## Sheet1
| Project | Baseline | Type | ID | Entry Type | Hazard Type | RequireKey | Requirement | Description | Status | Priority | Owner | Owner (First Name) | Owner (Last Name) | Created By | Responsible | Registered Interest | Created On | Last Modified On | Modified By | Source | Charm | Open Points | S1 Subsystem | Hazard Remark | Harm | AFAP justification | Kano Ranking | Test Level | Rearrange Operation | Rearrange Reason | Hazard | Initial Risk Assessment | Residual Risk Assessment | Validation Procedure State | New | Export | Level of Requirements | Level of Test | Usability | Domain | Applicable Generic Hazard | Hazard Effect To | System S1 | Target Release S1 | System IGS-Platform | Target Release IGS-Platform | System PreOPlan | Target Release PreOPlan | System ISS | IGS-Platform Subsystem | Target Release ISS | System S3 | Target Release S3 | S3 Subsystem | System NavSim | Target Release NavSim | Traced System | P1 before | P2 before | P1 after | Polarion Migration
0 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561938 | Headline | Chapter |  | Evaluation of Functions, Hazards, Risks and Measures | The evaluation of the risks is performed using the definitions and criteria in <XREF>561844</XREF>.The usage of the <XREF>561893</XREF> and the <XREF>568711</XREF> is explained in <XREF>561922</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563430 | Headline | Chapter |  | Move trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569143 | Headline | Function |  | Move trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 04.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
3 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561941 | Headline | Hazard |  | Crushing due to trolley tipping over |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
4 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561942 | Headline | Cause |  | Movement over inclined surface or monitor movement changes the center of gravity of the trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42#Haz01 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
5 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561943 | Requirement | Measure | design_of_system_angle_of_10_degree | Trolley design prevents overbalance | The trolley must be designed in a way that prevents tipping over when placed in any transport position on an inclined plane at an angle of 10 °. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.03.2022 11:21 | Rudolf Siebenhaar |  |  |  | TrolleyU | IEC 60601-1, cl. 9.4.2.1, 9.4.2.2 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
6 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561944 | Headline | Hazard |  | Crushing through tilting of the device |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
7 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561947 | Headline | Cause |  | Increase of tilt moment due to additional weight |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
8 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561949 | Requirement | Measure | design_of_system | Design of the trolley | The trolley must be designed in a way that prevents tilting of the trolley when driving it over an inclined plane (i.e. ramp). Remark:Compliance with DIN EN 60601-1 tilt test | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.03.2022 11:21 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
9 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561950 | Headline | Cause |  | Moving the trolley over a doorstep, against a doorframe, wall or other equipment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
10 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561951 | Requirement | Measure | rough_handling_test_trolley | Rough handling test trolley | The trolley shall be designed to comply with IEC 60601-1 cl. 15.3.5.Assumed speed of the trolley: 0,8 m/s +/- 0,1 m/s | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 25.03.2022 10:42 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
11 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563439 | Headline | Chapter |  | Startup main unit in standalone mode | Note, there are currently no risks related to "Startup main unit in standalone mode" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:11 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
12 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563431 | Headline | Chapter |  | Move main unit chassis in standalone mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
13 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564254 | Headline | Function |  | Move main unit chassis in standalone mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
14 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564255 | Headline | Hazard |  | Explosion, fire (HS_056, HS_030) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature,Release of stored energy,Flammability,loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
15 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564256 | Headline | Cause |  | Short circuit of the battery (external) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
16 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564257 | Requirement | Measure | coded_battery_connector | Coded battery connector | The battery connector for charging shall be coded in order to prevent connection with wrong polarity. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFSIEC 60601-1; 15.4.3.1 Housing, 15.4.3.2 Connection |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
17 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564258 | Requirement | Measure | battery_prevent_accidental_short_circuit | Battery housing preventing accidental short circuit | The housing of the battery shall prevent accidental short circuit. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 17.06.2022 08:43 | Johannes Schmidt |  |  |  | Batt |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
18 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564259 | Requirement | Measure | check_wiring_battery | Wiring check battery in manufacturing | The battery wiring shall be checked during manufacturing. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
19 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564267 | Headline | Cause |  | Overcharging of the battery (voltage, current) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
20 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564260 | Requirement | Measure | prevent_overcharging_battery | Prevent overcharging of battery | Overcharging of the battery shall be prevented regarding voltage and current. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
21 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571858 | Headline | Cause |  | Overcharging of the battery, sleeping error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
22 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571859 | Requirement | Measure | self-test_overcharging_battery | Self-test of battery overcharging protection | Self-test at startup of battery overcharging protection | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
23 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564268 | Headline | Cause |  | Internal short circuits and temperature increase of the battery packs |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
24 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564261 | Requirement | Measure | prevent_charging_high_temperature_battery | Prevent charging of battery with high temperature | The temperature of the battery unit shall be monitored. When a critical temperature is reached, the charging of the battery shall be blocked. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
25 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577242 | Requirement | Measure | ventilation_battery_housing | Ventilation battery housing | The battery pack housing shall be ventilated in such a way that in case of gas emission the accumulation of gases is prevented.Note: Gas emission only in error condition, not normal condition. Ventilation prevents accumulation of gases and possible ignition. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFSIEC 60601-1 cl. 15.4.3.1 |  |  | Basic | System Integration Test | empty | empty | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
26 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571860 | Headline | Cause |  | Internal short circuits and temperature increase of the battery packs, sleeping error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
27 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571861 | Requirement | Measure | self-test_prevent_charging_high_temperature_battery | Self-test of protection against charging of battery with high temperature | Self-test at startup of battery temperature protection | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.10.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
28 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564269 | Headline | Cause |  | Mechanical destruction |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
29 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564262 | Requirement | Measure | battery_housing_mechanical_strength | Battery housing mechanical strength | The battery shall be provided with a housing to protect the battery from external mechanical impact. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
30 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564263 | Requirement | Measure | battery_discharge_current_limitation | Battery discharge current limitation | The discharge current of the battery shall be limited (prevent sudden short circuit discharge, e. g. fuse at battery outlet). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
31 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564270 | Headline | Cause |  | Thermal runaway of battery pack (internal short circuit, overheating) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
32 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564265 | Requirement | Measure | fireproof_battery_housing | Battery housing fireproof and preventing flame break-out | The battery shall be provided with a fireproof housing (thermal runaway), preventing spread of flames. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
33 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564264 | Requirement | Measure | thermal_isolation_battery_cells | Thermal isolation of single battery cells | Thermal isolation of single battery cells against spread of thermal runaway chain reaction in other battery cells. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
34 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577220 | Headline | Hazard |  | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  |  | No gases are emitted during normal charging or discharging. |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature,Release of stored energy,Flammability,loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
35 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577236 | Headline | Cause |  | Increased temperature of the battery packs |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_056HS_030 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
36 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577237 | Headline | Measure | prevent_charging_high_temperature_battery | Reference - Prevent charging of battery with high temperature | <XREF>564261</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
37 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577241 | Headline | Measure | prevent_overcharging_battery | Reference - Prevent overcharging of battery | <XREF>564260</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
38 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577238 | Headline | Measure | battery_housing_mechanical_strength | Reference - Battery housing mechanical strength | <XREF>564262</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
39 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577239 | Headline | Measure | battery_discharge_current_limitation | Reference - Battery discharge current limitation | <XREF>564263</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
40 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577240 | Headline | Measure | thermal_isolation_battery_cells | Reference - Thermal isolation of single battery cells | <XREF>564264</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 25.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | Batt |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
41 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569265 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,slips, lapses and mistakes,Incorrect or inappropriate output,Software failure,Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
42 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569264 | Headline | Cause |  | Activation of control elements by non-authorized persons (e. g. passer-by in the corridor) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 16.08.2022 14:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
43 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564347 | Requirement | Measure | lock_movement_function | Lock movement initiation from chassis handle | The system shall provide the possibility to lock all motorized chassis movements activated via chassis handle by pressing a dedicated hardware button at the handle panel. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 13.08.2020 | 10.02.2022 10:23 | Johannes Schmidt |  |  |  | MCU,HICForceControl | Additionally, when the system is not operated for a certain time in standalone mode, it switches off (06S573383 in <XREF>561865</XREF>). |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
44 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569266 | Requirement | Measure | um_lock_motorized_movements | Caution - Lock chassis movement or switch off | CAUTION Cause: System is left unattended (e. g. in standalone mode).Possible consequences: Unauthorized operation of system movements.Remedy: Do not leave the system unattended when turned on. Use the lock movement function at the chassis handle panel or switch off the system. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 15.09.2020 | 07.10.2022 09:51 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
45 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569291 | Headline | Hazard |  | Severe limitation of movement functionality (HS_111) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
46 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569292 | Headline | Cause |  | Chassis battery runs empty, no more motorized movements possible |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 16.08.2022 14:34 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_111 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
47 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568840 | Requirement | Measure | indicate_battery_charge_standalone | Indicate battery charge for standalone mode | The system shall indicate the battery charge for standalone mode. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | Batt,MCU,HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
48 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564113 | Headline | Chapter |  | Move main unit chassis in transport mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
49 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564112 | Headline | Function |  | Move main unit chassis in transport mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
50 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564114 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
51 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564115 | Headline | Cause |  | User does not pay attention to obstacles in the movement path during transport movement (intentional) | User does not pay attention to obstacles in the movement path during movement (intended):User reacts too slowly to obstaclesSystem drives too fast for user reactions (by design, normal condition) or unpredictably (technical failures)Miscommunication in the ORSimultaneous movements of other equipment and the systemBad overview, structures and equipment covered other equipment | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
52 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564116 | Requirement | Measure | collision_avoidance_fast_drive_mode | Collision avoidance in fast drive mode | The system shall detect the approximation of solid obstacles at critical collision areas of the system during fast drive mode (e. g. towards moving direction), e.g. by means of a distance sensor. When an approximation is detected, the system shall reduce the speed to OR (operating room) mode speed before an imminent collision takes place. During normal operation (e. g. stopping without collision / approximation, emergency stop activation without collision / approximation), the system shall not create false positive collision / approximation detection information. The functionality of the sensor shall be tested during boot and monitored during operation (e. g. checking and calibrating by regularly scanning known fixed mounted objects with differing fixed distances to the sensor). When an error in the sensor system is detected, the system shall go to operating room speed (deactivate fast drive mode) and indicate the error to the operator. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.08.2022 13:56 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
53 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562463 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF> and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  |  | IEC 60601-1; 9.2.2.6 Speed of movement(s)IEC 60601-2-54; 201.9.2.2.6 Speed of movement(s) |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
54 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569215 | Headline | Measure | set_acceleration_values | Reference - Set acceleration values | The acceleration of motorized movements under normal conditions is defined in <XREF>569146</XREF>.For maneuverability of the system, the set speed values are deemed to be sufficient. The acceleration (and deceleration) is however monitored for mechanical stability <XREF>563680</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
55 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564118 | Requirement | Measure | restrictions_fast_drive_mode | Additional restrictions for fast drive mode | The max. fast drive speed (set speed value) shall only be allowed in forward and backward direction under consideration of the speed limits <XREF>563486</XREF> and permissible centrifugal acceleration <XREF>563681</XREF>, and when initiated from the chassis force handle. Furthermore, fast drive mode shall only be allowed if the footswitch is disconnected. When the system is in fast drive mode, and movements are initiated from other movement controls than from the chassis handle, the system shall apply max. OR mode speed (set speed). | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 22.09.2022 13:54 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
56 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564119 | Requirement | Measure | timeout_fast_drive_mode | Timeout for fast drive mode | After at most 300 seconds of chassis standstill in fast drive mode, there shall be a timeout deactivating the fast drive mode automatically. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 17.02.2022 10:33 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
57 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564120 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
58 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564121 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
59 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564122 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
60 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564123 | Headline | Measure | collision_detection_rear_side | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
61 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564124 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
62 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564125 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
63 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571907 | Headline | Cause |  | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
64 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571908 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
65 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578162 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
66 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571909 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF>The protection path knows when transport mode is activated. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
67 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571910 | Headline | Measure | restrictions_transport_mode | Reference - Additional restrictions for transport mode | <XREF>564118</XREF>Only chassis handle movements are allowed to initiate transport speed movements. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
68 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571911 | Headline | Cause |  | Unintentional activation of fast drive mode by the operator |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
69 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571914 | Requirement | Measure | protect_fast_drive_mode_button | Fast drive button protected against unintended activation | The button to activate the fast drive mode shall be protected against unintended activation, e. g. recessed button. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
70 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571912 | Headline | Measure |  | Reference - Timeout for transport mode | <XREF>564119</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
71 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571913 | Requirement | Measure | indication_fast_drive_mode | Indicate active fast drive mode | The system shall have an optical indication which shows that the fast drive mode is active. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU,HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
72 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578308 | Headline | Cause |  | Unintentional activation of fast drive mode by technical error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.02.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
73 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578309 | Requirement | Measure | safe_activation_fast_drive_mode | Safe activation fast drive mode | The fast drive mode shall be activated by two redundant paths. Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design contraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.02.2021 | 17.06.2022 08:53 | Johannes Schmidt |  |  |  | MCU,HICForceControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
74 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571915 | Requirement | Measure | check_fast_drive_mode_button | Check fast drive mode button state | The system shall check the status of both fast drive button signal paths during boot-up. When an active fast drive button signal path is detected, the fast drive mode shall be blocked and an error message shall be displayed until a deactivation and regular re-activation of both fast-drive release paths takes place. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 22.10.2020 | 29.09.2022 14:48 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
75 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563432 | Headline | Chapter |  | Drape the system | Note, drapes are handled at other positions of this document. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 27.06.2022 08:26 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
76 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563433 | Headline | Chapter |  | Connect trolley to line power | Note, there are currently no risks related to "Connect trolley to line power" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:19 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
77 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563434 | Headline | Chapter |  | Connect trolley and main unit | Note, there are currently no risks related to "Connect trolley and main unit" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:19 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
78 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563429 | Headline | Chapter |  | Startup the system via trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 20.06.2022 11:12 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
79 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563478 | Headline | Function |  | Startup the system via trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
80 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563479 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | HS_042Crushing by or collision with objects moving with highkinetic energy and / or driving force and contactsurfaces shaped in a way that contact force andpressure are high | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
81 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563480 | Headline | Cause |  | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | Unintentional activation of movement during startup:Unintended activation of control element by the operator, e. g. control element is clampedErroneous activation of movement due to technical error, e. g. movement key is stuck | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 16.08.2022 14:35 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
82 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562344 | Requirement | Measure | Check_AtStartUp_s1 | Check movement release signal status at startup | The system shall check the status of both movement release signal paths at startup.When an active movement release signal path is detected, movements from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both movement release paths takes place.? Note 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Note 2: The emergency stop of the affected control remains functional. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 08:58 | Johannes Schmidt |  |  |  | MCU | SFS (for all movement controls, and for all wired X-ray controls, but not for wireless footswitch no PM is involved) |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
83 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562378 | Requirement | Measure | flat_control_elements | Control elements protected against unintended movement or x-ray activation | The control elements shall be designed in such a manner that unintended movement or x-ray activation by the operator is prevented e.g.:- Joysticks are recessed and have an additional movement triggering switch- No locking of buttons in activated position possible- For the wireless and wired remote control, no function shall be activated when the component is dropped or put reversely on a flat surface (see also 60601-1 cl. 15.3.1).- At least control elements related to switching the system on/off, chassis movement initiation, brake control and x-ray release shall be recessed.- Handle for power assisted chassis movements shall have two switches / buttons / sensors which are physically separated.- The system shall only allow movement initiation from the handles if force is applied within 2 s timeout after the capacitive sensors are activated.Note: In case the timeout passes by without force being applied or when trying to reinitiate a movement after a temporary stop that exceeds the timeout, the operator needs to release the handles and start a new movement. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 23.08.2022 11:44 | Johannes Schmidt |  |  |  | HICForceControl,HICHandControl | SFSIEC 60601-1; 9.2.3.1 Unintended movementIEC 60601-2-54; 201.9.2.3.1 Unintended movement |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
84 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562377 | Requirement | Measure | dead_mans_push_buttons | Continuous activation of movements with redundant activation paths | Motorized movement shall be possible only when two redundant signals provided by two independent paths are present (continuous activation), preventing unintended movement activation. When only one of two paths is activated, all motorized movements shall be stopped and inhibited for the respective control.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. Further note: For network interfaces cyclic retriggering (cyclic messages within e.g. 150 ms) are seen as equivalent to continuous activation. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 27.09.2022 11:39 | Johannes Schmidt |  |  |  | HICSwitch,MCU,HICForceControl,HICHandControl | SFSIEC 60601-1; 9.2.3.1 Unintended movementIEC 60601-2-54; 201.9.2.3.1 Unintended movementIEC 60601-1; 9.2.2.5 Continuous activationIEC 60601-2-54; 201.9.2.2.5 Continuous activation |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
85 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562345 | Requirement | Measure | EmergencyStopMovement_s1 | Emergency stop at every point of motorized movement activation | Emergency stop shall be available at every point of motorized movement activation.Points of movement activation (control locations) are:- control panel at the FD cover- wireless hand control- wired hand control- Backside of Servo-Handle Chassis The emergency stop shall be designed according to IEC 60947-5-1 and IEC 60947-5-5. When an emergency stop is activated, all motorized movements of the system except the motorized monitor height adjustment shall be stopped (e. g. stop category 1, thereof e. g. safe stop 1 function according to 61800-5-2) and inhibited until it is deactivated again. ?Note 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. Note 2: For behaviour regarding radiation, see <XREF>562135</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 08:59 | Johannes Schmidt |  |  |  | MainU,MCU,HICForceControl,HICHandControl | SFSIEC 60601-1; 9.2.4 Emergency stopping devicesIEC 60601-1; 9.2.2.4.4 Other RISK CONTROL measuresIEC 60601-1; 9.2.2.5 Continuous activationIEC 60601-2-54; 201.9.2.3.1 Unintended movementIEC 60601-2-54; 201.9.2.4.101 Controls |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
86 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571863 | Headline | Cause |  | Unintentional activation of movement during startup, sleeping or unexpected error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
87 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571864 | Requirement | Measure | self-test_protection_hardware | Self-test of protection module hardware | The protection module hardware shall execute a self-test of its safety relevant hardware at startup.When the test fails, all motorized movements shall remain inhibited. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
88 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571866 | Headline | Measure | Check_AtStartUp_s1 | Reference - Check movement release signal status at startup | <XREF>562344</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
89 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578161 | Requirement | Measure | safe_power_supply_protection_hardware | Safe power supply of protection module hardware | The protection module hardware shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. When the startup test fails or one of the over-/undervoltage protection detects an error, all motorized movements and radiation shall be inhibited. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 28.06.2022 14:38 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
90 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578307 | Requirement | Measure | general_principles_protective_path | General principles protective path | The following general principles for designing the protective path shall be considered:- The software of the protective path shall be developed according to software safety class C (IEC 62304).- Common causes (errors or environmental conditions that could affect C and P simultaneously) shall be considered and excluded by design: ESD / transients (sufficient electrical isolation between P and C path signals), temperature out of specified limits, etc. For components which are connected to the protection module, especially the integration requirements of <XREF>580888</XREF> shall be considered. Note 1: Safe power supply of P and the mechanical mounting of the encoders have already separate requirements <XREF>578161</XREF>, <XREF>578163</XREF>,<XREF>578173</XREF>. Note 2: For design contraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 05.02.2021 | 22.08.2022 15:35 | Johannes Schmidt |  |  |  | HICSwitch,MainU,MCU,HICForceControl,HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
91 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569274 | Headline | Hazard |  | Medium-term delay of the clinical procedure (HS_091) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
92 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569275 | Headline | Cause |  | X-ray chain is not functional (e. g. communication error, hardware error) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_091 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
93 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564352 | Requirement | Measure | x-ray_chain_selftest_startup | X-ray chain selftest at startup | The system shall perform a cross-component self-test of the X-ray chain at startup of the system:HV_trigger, idexi line from FLC (IC) to Generator interface and HV_ON with the involved subsystems Xgen, GenIF, FLC, XCU.All involved subsystems / components shall report if the test was successful or not. If the test was not successful, a message shall be displayed to the operator (no X-ray blocking).Remark: HW error in HV_Trigger, HV_On then detectable before X-ray release, cross-component selftest (and that DEXI is not sent to generator due to network configuration issue) | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 13.08.2020 | 17.02.2022 13:18 | Johannes Schmidt |  |  |  | FLC,MCU,XCU,XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
94 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569297 | Headline | Cause |  | Delayed boot-up or restart time to recover system functions (e. g. after error, SW freeze) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_091 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
95 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568860 | Requirement | Measure | recovery_management | Recovery management | The following time restrictions shall be met for recovery management (this is also valid for EMC immunity testing): * Time from initiating a shut down via reset button, booting the system again, until X-ray ready: 120 s? * Time from initiating a shut down via reset button, until the previous settings are restored again (this includes registering a patient, set loading parameters (kV, mA, p/s), move the system from angular 0° to an exemplary previously stored position of angular 180°): 4 min * Time from initiating a regular shut down via off button, booting the system again, until X-ray ready: 4 min * Time from power supply interruption, booting the system again, until X-ray ready: 120 sNote: Due to the limited time to recover all system functions, no emergency radiography mode needs to be provided. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2020 | 28.02.2022 11:18 | Johannes Schmidt |  |  |  | FLC,MCU,PowerU,XCU,XGen | IEC 60601-2-43; cl. 201.4.102 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
96 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563435 | Headline | Chapter |  | Register patient | Note, there are currently no risks related to "Register patient" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:20 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
97 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563628 | Headline | Chapter |  | System operation - general |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
98 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563631 | Headline | Function |  | System operation - general |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
99 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563627 | Headline | Hazard |  | Crushing by or collision with falling or tipping objects with high potential energy and / or high contact pressure (HS_051) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Suspending parts,Instability,Vibration,Moving parts,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
100 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563629 | Headline | Cause |  | Mechanical support structures of the wheels break, the system suddenly tips to one side or tips over completely | Mechanical support structures of the wheels break, the system suddenly tips to one side or tips over completely (during movement or standstill possible):Design errorManufacturing errorAssembly error | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | High | A
101 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561959 | Requirement | Measure | design_of_wheels | Mechanical design drive and wheels | The gear, motor shaft and castors shall be designed appropriately for mechanical stress according to IEC 60601-1 cl. 9.8.2. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 12:51 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
102 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578748 | Headline | Hazard |  | Limited imaging functionality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Fatigue failure |  | S1M | VB10 | [12 empty] | NotDetermined
103 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578749 | Headline | Cause |  | Performance of the imaging system is degraded due to system mangement application |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_101 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | High | A
104 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578750 | Requirement | Measure | imaging_system_performance_test | Boundary conditions for imaging system performance test | During all phases of the imaging system performance test, the MNP has to be installed and be active all the time to verify that MNP does not disturb the overall system performance. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 15.06.2022 16:42 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
105 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597002 | Headline | Hazard |  | Effect on people with implants, e.g. pacemaker |  | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 27.06.2022 16:24 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
106 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597003 | Headline | Cause |  | Interference of electromagnetic fields into implant/s electronics |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz44 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Electro-magnetic influence (EMC) | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
107 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597004 | Requirement | Measure | Label_electromagnetic_fields | Label electromagnetic fields | A label shall be attached at the backside of the chassis, such that it is visible to an operator when moving the system. It shall contain the following safety signs: ISO 7010-P007, ISO 7010-W006, ISO 7010-W001, ISO 7010-M002 | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 07.10.2022 09:11 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
108 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597005 | Requirement | Measure | docUser_electromagnetic_fields | User Manual: Electromagnetic fields Chassis | CAUTION Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes, chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Possible consequences: Hazards on people with implants due to electromagnetic disturbances. Remedy: If you have active implanted cardiac devices keep a distance of >10 cm to the above mentioned components. | InWork | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 07.10.2022 09:53 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
109 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597007 | Requirement | Measure | docServ_electromagnetic_fields | Service Doc: Electromagnetic fields Chassis | CAUTION in Service Manual Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes, chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Possible consequences: Hazards on people with implants due to electromagnetic disturbances. Remedy: If you have active implanted cardiac devices keep a distance of >10 cm to the above mentioned components. | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 27.06.2022 13:56 | Johannes Schmidt | [7 empty] | Basic | Subsystem Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
110 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597011 | Requirement | Measure | docSCM_electromagnetic_fields | Factory Doc: Electromagnetic fields Chassis | CAUTION in Factory Documentation Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes, chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Possible consequences: Hazards on people with implants due to electromagnetic disturbances. Remedy: If you have active implanted cardiac devices keep a distance of >10 cm to the above mentioned components. | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 27.06.2022 13:57 | Johannes Schmidt | [7 empty] | Basic | Subsystem Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
111 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597013 | Headline | Hazard |  | User has skin irritation |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 27.06.2022 16:24 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | [16 empty] | NotDetermined
112 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597014 | Headline | Cause |  | System contains non-biocompatible material |  | InWork | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_085 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Bio-incompatibility | User,Service Personnel | [14 empty] | NotDetermined | C | Low | B
113 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597019 | Requirement | Measure | biocompatible_material | Biocompatible material | The clinical user and the service personnel get in regular hand skin contact with the following components of the system: AccesS (RCP cart, Keyboard, Mouse), DSmonS (Monitor), GridS, HICSwitch, HICTouchCon, MainU (Covers), TrolleyU (Covers, X10 cable and plug), HICForceControl, HICHandControl, These components shall be made out of biocompatible material or at least it shall be shown that the utilized material is in common use for regular consumer products or that there is an established history of safe use in a clinical environment, e.g. from previous CIOS systems. | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 09.08.2022 15:19 | Nils Naarmann |  |  |  | AccesS,DSMonS,GridS,HICSwitch,HICTouchCon,MainU,TrolleyU,HICForceControl,HICHandControl |  |  |  | Basic | Component Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
114 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563437 | Headline | Chapter |  | Position patient | Note, there are currently no risks related to "Position patient" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:20 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
115 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563438 | Headline | Chapter |  | Move main unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
116 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563475 | Headline | Function |  | Move main unit - motorized, general |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
117 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562459 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
118 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562462 | Headline | Cause |  | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | Unintentional activation of control elements by the operator:Operator leans over control elements or hits them by accident while doing something elseOperator hangs lead protection over control elements (handle for power assisted movement)Other objects fall onto control elementsRemote control falls on the floor and gets activated | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
119 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563485 | Headline | Measure | flat_control_elements | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
120 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578331 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 08.02.2021 | 17.06.2022 09:47 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
121 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578316 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
122 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578332 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 08.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
123 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563488 | Requirement | Measure | prerequisites_fast_drive | Prerequisites for fast drive mode | Prerequisites for fast drive mode: Vertical axle shall be in the max. downwards position and the C-arm in the upright position (orbital/angular = 0° +/- 5° tolerance). The correct position shall be detected single fault safe. The trolley-chassis cable shall be disconnected from the main unit.Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 05.08.2020 | 17.06.2022 09:00 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
124 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578291 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 04.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU | SFSIEC 60601-1; 9.2.2.4.4 Other RISK CONTROL measuresIEC 60601-2-54; 201.9.2.3.1 Unintended movementIEC 60601-2-43; 201.9.2.4 Emergency stopping devices + aa) |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
125 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578358 | Headline | Measure | collision_detection_rear_side | Reference - Collision detection at rear side of the chassis and FD | <XREF>563491</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 09.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
126 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571916 | Requirement | Measure | collision_detection_other_axles | Collision detection via position change of other axles | Position changes of motorized axles which are not used for motorized movements shall be detected (although the brakes of orb. and ang. axles e. g. are activated each during other motorized movements, they can be overridden manually).In case a non-expected position change of more than +/- 5° is detected, all motorized movements shall stop.Hint: The system may be configured to stop at even lower thresholds < 5°. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.10.2020 | 12.09.2022 11:34 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
127 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563493 | Requirement | Measure | um_risk_collision_general | Warning - General risk of collisions | WARNINGCause: System movements.Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment.Remedy: Ensure that movements are released only if neither the operator, the patient, other persons or other equipment cannot be endangered.Always pay attention to possible collisions during movements.Make sure that nobody is standing inside the danger area.Ensure that collision sensors of the system are not covered by any accessories e.g., drape fixtures, which would unintentionally activate or prevent activation of the collision detection system.Remove any objects from collision area, e.g. injector or infusion stand.Body penetrating objects introduced in the patient (e.g. catheter, biopsy / kyphoplasty needle) must not be actuated by system movements.Special care is required when the system has already detected a collision situation.Use the emergency stop if required. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 05.08.2020 | 11.10.2022 11:22 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
128 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563494 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
129 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563495 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
130 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571868 | Headline | Cause |  | Speed limits are exceeded due to sleeping or unexpected error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
131 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571869 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
132 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571870 | Requirement | Measure | Positioning_validity_of_motorized_axes | Positioning validity of motorized axes | During every movement in orbital or in angular direction, a comparison between their protective position sensors and their position sensors used by the control system shall be executed. When an error is detected, motorized movements of the affected axle shall be inhibited. Note1: This replaces a regular self-test at startup. This is also required since the protective position sensors do not provide safe positions according to SIL2 or higher. Note2:The movement stop affects all system drives. Movements from the not affected drive can be initiated directly again. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 20.10.2020 | 19.08.2022 10:59 | Rudolf Siebenhaar |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
133 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578163 | Requirement | Measure | independent_power_supply_protective_position_sensors | Independent power supply of protective position sensors | The protective position sensors (orbital, angular and for mecanum wheels) shall have an independent power supply from the other position sensors. Hint: The power supply for one position sensor shall be provided by the c-path, the power supply of the other position sensor shall be provided by p-path. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 02.02.2021 | 05.10.2022 13:38 | Peter Hefter |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | NEW | empty | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
134 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578164 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
135 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 595374 | Headline | Measure | prevent_incorrect_movement_mecanum | Reference - Prevent incorrect movement execution - mecanum | <XREF>563729</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 18.03.2022 | 18.03.2022 11:00 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,Service Personnel,Environment | S1M | VB10 | [12 empty] | NotDetermined
136 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571872 | Headline | Cause |  | Collision detection fails due to sleeping error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
137 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571874 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
138 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571873 | Requirement | Measure | self-test_collision_sensors | Self-test of the collision sensors | The functionality of the collision sensors shall be tested at startup and continuously during operation (no connection). When an error in the collision detection system is detected, the system shall stop all motorized movements and show an error message to the operator. Subsequent movements are allowed with collision mode speed. For detailed behaviour, see collision detection requirements <XREF>563490</XREF> and <XREF>563719</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 28.07.2022 09:21 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
139 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571876 | Requirement | Measure | design_collision_sensors | Reliable design of collision sensors | The collision sensors shall be designed reliably for an expected service life of 10 years (e. g. for an electro-mechanical switch e. g. 100.000 switching cycles). | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 20.10.2020 | 03.08.2022 13:17 | Johannes Schmidt |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
140 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563492 | Requirement | Measure | um_check_collision_sensors | Warning - Check collision sensors | WARNINGCause: Undetected failure of the collision sensors.Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment.Remedy: Perform a daily check of the collision sensors at the chassis and the C-arm according to the steps below. Check that the collision sensors are functioning and verify that if a collision has been detected all motorized movements are stopped immediately. Messages must be displayed. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 05.08.2020 | 09.02.2022 16:35 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
141 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563645 | Headline | Cause |  | Unintentional activation of movement due to technical failure (without the operator's request / approval) | Unintentional activation of movement due to technical failure (without the operator's request / approval):Technical failure, movement button failure, generatedMovement command gets generated by wrong CAN telegram via EMC influences | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
142 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563646 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 13.09.2022 17:18 | Johannes Schmidt |  |  |  | HICSwitch,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
143 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578359 | Headline | Measure | EMC_s1 | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 09.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,DapS,DSMonS,FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
144 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563622 | Headline | Cause |  | User does not pay attention to feet in the movement path during movement (intended): System drives over toes |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
145 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563625 | Requirement | Measure | mech_safety_gaps_mecanum | Mechanical safety - allowed gaps mecanum | The gaps between the floor and chassis cover around the mecanum wheels shall be < 35 mm according to IEC 60601-1 table 20 (adult foot). In addition to that, all areas of possible crushing between the covers and the floor shall be covered by <XREF>563490</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 06.08.2020 | 29.09.2022 15:09 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1 cl. 9.2.2.2, cl. 9.2.2.4.1, cl. 9.2.2.4.4 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
146 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599924 | Headline | Measure | hm_label_warning_crushing_foot | Reference - warning label risk of crushing foot | <XREF>562271</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 12.09.2022 | 12.09.2022 10:13 | Johannes Schmidt | [7 empty] | Basic | System Test | none | empty | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1D | VB10 | [12 empty] | NotDetermined
147 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563636 | Headline | Cause |  | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | User does not pay attention to obstacles in the movement path during movement (intended):User reacts too slowly to obstaclesSystem drives too fast for user reactions (by design, normal condition) or unpredictably (technical failures)Miscommunication in the ORSimultaneous movements of other equipment and the systemBad overview, structures and equipment covered by sterile drapes | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
148 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563638 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 06.08.2020 | 17.06.2022 09:48 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
149 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563639 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
150 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563640 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
151 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563641 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
152 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563642 | Headline | Measure | collision_detection_rear_side | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
153 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563643 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
154 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563644 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
155 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563674 | Headline | Cause |  | System does not move according to control parameters: Resulting system speed too high due to technical error | System does not move according to control parameters: Resulting system speed is too high due to technical error in the control system:SW error in control system (e. g. wrong speed is set by control system)Random HW error in control system | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
156 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563676 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
157 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563679 | Headline | Cause |  | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) and speed limit might be reached ""too late""SW error in control system (e. g. wrong acceleration is set by control system)Random HW error in control system | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
158 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563680 | Requirement | Measure | limit_acceleration_deceleration | System acceleration and deceleration limits | The system shall limit the acceleration and deceleration so that tipping over of the system is prevented. In case the acceleration / deceleration threshold is exceeded, the system shall stop all motorized movements.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:00 | Johannes Schmidt |  |  |  | MCU | SFS | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
159 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563681 | Requirement | Measure | limit_centrifugal_accelleration | Limit system centrifugal acceleration | The system shall limit the maximum permissible centrifugal acceleration so that tipping over of the system is prevented.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:01 | Johannes Schmidt |  |  |  | MCU | SFS | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
160 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563682 | Headline | Measure | check_plausibility_sensors | Reference - Plausibility check for sensors in movement control elements | <XREF>563673</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
161 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571877 | Headline | Cause |  | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
162 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571878 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
163 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578166 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
164 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571879 | Headline | Measure | self-test_protection_position_sensors | Reference - Self-test of the protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
165 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571880 | Headline | Measure | mounting_protection_position_sensors | Reference - Mounting of protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
166 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578165 | Headline | Measure | safe_power_supply_protective_position_sensors | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
167 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580816 | Headline | Cause |  | System does not move according to control parameters: During quickstop, the system violates trajectory trueness |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
168 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596883 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 02.06.2022 | 15.06.2022 14:56 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
169 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563720 | Headline | Cause |  | Collision sensor is active, operator moves further into collision by mistake |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
170 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563719 | Requirement | Measure | behaviour_collision | Behaviour in collision | When the operator wants to move the system out of a collision (all movements have been stopped), the system shall behave as follows:Movements shall only be possible with reduced speed, after new activation of movement. The reduced speed is maintained as long as the movement is activated (and not as long as collision sensor is activated, no "automatic" acceleration). When the movement is stopped, and re-initiated and the collision state has been removed, the subsequent movement shall be performed normally again. In case additional collision sensor are activated, the system shall also behave as described above (stop, reduced speed after new movement activation, check real system position changes). Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:02 | Johannes Schmidt |  |  |  | MCU | SFSFollow-up hazard: 06H563718 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
171 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563722 | Requirement | Measure | torque_limits_mecanum | Torque limits mecanum | The system shall restrict the torque of the mecanum drives when collision detection is active. If a torque exceeding the limit is detected, the mecanum drives shall stop. Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 18.06.2022 09:56 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
172 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563723 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
173 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576754 | Headline | Measure |  | Reference - Override angular / orbital brakes manually | The orbital and angular axles can be overriden manually, see <XREF>562449</XREF>. Thus, the system may be able to retract, depending on the initial axle positions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
174 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571893 | Headline | Cause |  | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
175 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571894 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
176 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578167 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
177 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571895 | Headline | Measure | self-test_protection_position_sensors | Reference - Self-test of the protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
178 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571896 | Headline | Measure | mounting_protection_position_sensors | Reference - Mounting of protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
179 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578168 | Headline | Measure | safe_power_supply_protective_position_sensors | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
180 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571881 | Headline | Cause |  | Collision sensor is active, operator moves further into collision by technical error (e. g. joystick deflection sensor) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
181 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571882 | Headline | Measure | behaviour_collision | Reference - Behaviour in collision | <XREF>563719</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
182 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571883 | Headline | Measure | torque_limits_mecanum | Reference - Torque limits mecanum | <XREF>563722</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 20.10.2020 | 17.06.2022 10:01 | Johannes Schmidt |  |  |  | MainU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
183 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563747 | Headline | Cause |  | System does not stop when user releases control element or stops too late | System does not stop when user releases control element or stops too lateSystem does not recognize deactivation of control elementElectro-magnetic disturbances prolongs movementControl system does not brake or brakes too slowly (does not initiate Quickstop or selects wrong braking ramp) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
184 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563748 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
185 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563749 | Requirement | Measure | safe_stop_motorized_movements | Safe stop motorized movements | The braking process (regular release of move enable or quickstop) shall be monitored. If the system detects an increased delay in reaction or a wrong braking ramp, the system shall stop, e. g. by implementing a safe stop 2 function according to IEC 60204-1 or equivalent.The fast braking process (quickstop) shall be implemented single fault safe. When the system detects an increased delay in reaction or a wrong braking ramp, the system shall stop, e. g. by implementing a safe stop 2 function according to IEC 60204-1 or equivalent. Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 29.09.2022 15:11 | Johannes Schmidt |  |  |  | MainU,MCU | SFS(The implementation is two stage: 1. PM tries to trigger QS at the motorcontroller, 2. PM triggers STO and brakes.) |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
186 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563750 | Headline | Measure | EMC_s1 | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,DapS,DSMonS,FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
187 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571884 | Headline | Cause |  | System does not stop when user releases control element or stops too late due to sleeping error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
188 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571885 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
189 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578169 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
190 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571886 | Headline | Measure | self-test_protection_position_sensors | Reference - Self-test of the protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
191 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578170 | Headline | Measure | safe_power_supply_protective_position_sensors | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
192 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571887 | Headline | Measure | mounting_protection_position_sensors | Reference - Mounting of protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
193 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576755 | Requirement | Measure | self-test_STO | Self-test STO functionality | A self-test of the ability to trigger STO via the protection module shall be performed at startup.In case the test is not passed, the motorized movements affected by the failed STO (e.g. mecanum wheel base) shall remain blocked. The motorcontrollers shall have a certified STO functionality according to SIL2. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.01.2021 | 27.09.2022 12:07 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
194 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576756 | Headline | Measure | automatic_brake_test | Reference - Self-test of the brakes | <XREF>564087</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFSFollow-up hazard: 06H571856 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
195 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563758 | Headline | Cause |  | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | System does not stop when user releases control element or system moves without activation:Erroneous prolongation of movement due stuck movement key | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
196 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563759 | Headline | Measure | flat_control_elements | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
197 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563760 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
198 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563761 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
199 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563762 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
200 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563763 | Headline | Cause |  | System does not stop when user releases control element: Overtravel is increased due to slip , e. g. fluid on the floor |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
201 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563765 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
202 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563725 | Headline | Hazard |  | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | Crushing by or collision with falling or tipping objectswith high potential energy and / or contact surfacesshaped in a way that contact force and pressure arehigh (HS_051) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Instability,Moving parts,Fatigue failure |  | S1M | VB10 | [12 empty] | NotDetermined
203 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563726 | Headline | Cause |  | System tips over during movement: One drive blocks, one brake gets activated, random HW error | One drive blocks in worst case speed and axle positionsBrake activatedRandom HW error in control path | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
204 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563727 | Headline | Measure | limit_acceleration_deceleration | Reference - System acceleration and deceleration limits | <XREF>563680</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 30.05.2022 09:45 | Johannes Schmidt |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
205 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563728 | Headline | Measure | limit_centrifugal_accelleration | Reference - Limit system centrifugal acceleration | <XREF>563681</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 30.05.2022 09:46 | Johannes Schmidt |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
206 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563729 | Requirement | Measure | prevent_incorrect_movement_mecanum | Prevent incorrect movement execution - mecanum | The system shall check if the current mecanum drive information is a valid mecanum movement (fulfill the mecanum equations). When an incorrect movement is detected, the system shall stop motorized movements.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:03 | Johannes Schmidt |  |  |  | MCU | SFS | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
207 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563730 | Requirement | Measure | prevent_tipover_bydesign | Prevent tip over by mechanical design | System shall not tip over by mechanical design in the following normal and fault conditions:- emergency stop during max. transport speed in transport position- emergency stop during curve movement in the worst case axle positions and movement speed- system is stopped with max. permissible external additional components mounted in the worst case axle positions and OR speed- one mecanum drive blocks completely during movement in the worst case axle position and movement speed- all mecanum drives block completely during movement (brakes activated) in the worst case axle position and movement speed- emergency stop at OR speed and inclined surface of 7° (no transport mode speed) Note:During normal stopping (releasing movement control element), the system shall not lift at one side at all.During other stopping scenarios, it is allowed that the rear mecanum wheels of the system lift temporarily. The short duration of the lifting combined with the expected position of the operator with respect to the possible trapping zone prevent a crushing of the operator foot with the system during this scenario. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 01.06.2022 09:23 | Johannes Schmidt |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E
208 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571888 | Headline | Cause |  | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
209 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571889 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
210 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578171 | Headline | Measure | safe_power_supply_protection_hardware | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
211 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571890 | Headline | Measure | self-test_protection_position_sensors | Reference - Self-test of the protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
212 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571891 | Headline | Measure | mounting_protection_position_sensors | Reference - Mounting of protective position sensors | <XREF>571870</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
213 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578172 | Headline | Measure | safe_power_supply_protective_position_sensors | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
214 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563731 | Headline | Cause |  | System tips over during movement or during braking: Safety mechanism is activated (e. g. quickstop is not initiated at all) | Safety mechanism is activated (e. g. quickstop is not initiated at all)Emergency stop is activated | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
215 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563732 | Headline | Measure | prevent_tipover_bydesign | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
216 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563733 | Headline | Cause |  | System tips over during movement or during braking: Additional weight is mounted at the C-arm |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
217 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563734 | Requirement | Measure | docNav_additional_weight | Weight restriction mounting external components | The weight of additional components mounted at the C-arm (like navigation components) shall be restricted to max. 2 kg. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
218 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563735 | Requirement | Measure | um_weight_limit_additional_components | Warning - Weight limit for additionally mounted components | WARNINGCause: Additional components are attached to the C-arm.Possible consequences: System movements behave differently, C-arm may be counterbalanced, risk of crushing or tipping over is increased.Remedy: Only mount approved components to the C-arm. The max. additional weight is max. 2 kg. Ensure that the floor is free of obstacles and is not slippery. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
219 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563770 | Headline | Measure | prevent_tipover_bydesign | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 01.06.2022 09:15 | Johannes Schmidt |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | empty | empty | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
220 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578343 | Requirement | Measure | um_steep_ramp | Warning - Inclined surface restriction | WARNINGCause: System is moved on a steep ramp, over a step or uneven ground. Possible consequences: System movements behave differently. The C-arm may be counterbalanced or collide with the floor. Risk of crushing or tipping over is increased.Remedy: The system is designed for steps of max. 2 cm and ramp pitches of max. 5° when in transport position. If moving on steeper ramps the lifting column of the C-arm needs to be raised enough to avoid collision with the floor and the system must be moved with care. Note, do not raise the lifting column unnecessarily high because the risk of tipping over is increased. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 08.02.2021 | 15.06.2022 18:00 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | empty | empty | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
221 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596712 | Requirement | Measure | serv_inclined_surfaces | Warning - Service instruction for inclined surfaces | Warning - Service Instruction for inclined surfaces WARNINGCause: ???????? System is moved on a steep ramp, over a step or uneven ground.Possible consequences: ??System movements behave differently. The C-arm may be counterbalanced or collide with the floor. Risk of crushing or tipping over is increased.Remedy: ????????The system is designed for steps of max. 2 cm and ramps of max. 7°. Note, when the system is in transport position ramp pitches of max. 5° are possible. If moving on steeper ramps the lifting column of the C-arm needs to be raised enough to avoid collision with the floor and the system must be moved with care. Do not raise the lifting column unnecessarily high because the risk of tipping over is increased. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.05.2022 | 29.09.2022 15:33 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
222 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563736 | Headline | Cause |  | System tips over during movement or during braking process: Curve radius too small in relation to speed | Curve radius too small in relation to speed, e. g. during braking:HW errorSW errorUser error (unintended steering command at the handles, e. g. due to tripping) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
223 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563737 | Headline | Measure | limit_centrifugal_accelleration | Reference - Limit system centrifugal acceleration | <XREF>563681</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
224 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563766 | Headline | Cause |  | System tips over during movement or during braking process: Normal braking ramp is too harsh | System is out of balance during stopping:Normal braking ramp is too harsh, e. g. due to design error (worst case axle positions not considered,...), or SW error, random HW error | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
225 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563767 | Headline | Measure | prevent_tipover_bydesign | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
226 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563768 | Headline | Measure | limit_acceleration_deceleration | Reference - System acceleration and deceleration limits | <XREF>563680</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
227 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580818 | Headline | Cause |  | System lifts at one side during braking process |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_051 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
228 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580819 | Headline | Measure | prevent_tipover_bydesign | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
229 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563632 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,slips, lapses and mistakes,Incorrect or inappropriate output,Software failure,Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
230 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563633 | Headline | Cause |  | Two operators attempt to initiate system movements at the same time (intentionally) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
231 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563634 | Requirement | Measure | only_one_active_control_element | Only one active movement control element at a time | The system shall be designed in such a way that if a control element (e.g., button or handle) is actively used for a specific movement, operation of any movement from a different control element is inhibited until the active control element is deactivated. Note:- Emergency stops and in particular collision sensors are not affected and always overrule the action of control elements if activated.- In case a movement has been terminated e.g. by means of a collision, a new movement can be initiated from a different control device even if the control element that triggered the initial movement is still activated, e.g. a still pressed button or touched handle.- Initiating more than one movement by pressing more than one button simultaneously at the same device shall also not be possible.- Other system functionalities apart from system movements such as activating the laser light shall not be affected by this requirement and can be performed simultaneously during movements. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 06.08.2020 | 23.09.2022 09:17 | Johannes Schmidt |  |  |  | MCU | Machinery directive (EHSR); 1.2.2 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
232 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563635 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
233 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569307 | Requirement | Measure | no_movement_takeover_across_controls | No movement take over across controls | When another control button is activated (not controlling the movement) while the movement controlled by another control button is stopped, the movement shall be stopped until a deactivation and reactivation of the movement command took place. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 17.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | Machinery directive (EHSR); 1.2.2 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
234 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569308 | Requirement | Measure | movement_signals_same_control_button | Redundant signals from same control button | Only movement release signal pairs from the same control button or from control buttons which share one enable line shall trigger movements.Note: Chassis handle panel transport position button and chassis handles share the enable line. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 17.09.2020 | 23.12.2021 14:27 | Johannes Schmidt |  |  |  | MCU | Machinery directive (EHSR); 1.2.2 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
235 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563647 | Headline | Cause |  | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | System moves in wrong direction due to operating error (intentional movement):User mixes up directions of a joystickUser mixes up buttonsFD force handle: Operator has unintentionally locked an electro-magnetic brake or has locked the wrong brake (this locks the movement direction when movements are initiated via FD handle) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
236 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563648 | Requirement | Measure | movement_preview | Movement preview for position movements | For all position movements, the system shall have an optical indication of the planned system movement (movement preview), as soon as the position movement is selected (not initiated) (e.g. pictogram, color marking displaying the difference between the current pose and the target pose with respect to the system orientation).?Exceptions: During image acquisition in any mode no preview is required. The movement preview shall be shown after the image acquisition has finished. For movement to the transport and 0/0 position, no preview is required.?Note 1: The operator has the possibility to check the subsequent movement.?Note 2: Position movements are e. g. position storage, park & return, default positions.?Note 3: Only a preview is required. The preview is required before start of movement. During movement, no display is required (functionally it might be reasonable though in order to see whether the desired position has been reached, at least towards the end of the movement sequence), since the operator observes the mechanical movement of the system. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 06.08.2020 | 18.06.2022 09:57 | Johannes Schmidt |  |  |  | FLC,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
237 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563650 | Requirement | Measure | um_collisions_position_movements | Warning - Collisions during position movements without test run | WARNINGCause: Movement to a predefined position.Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment.Remedy: Prior to initiating a movement to a predefined position, check the planned movement path via the movement preview and ensure that the path is free of any obstacles.Ensure that the patient is fixated at the table and that the table is positioned correctly (tilt etc.). Observe the movement continuously. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 06.08.2020 | 14.01.2022 09:09 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
238 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563651 | Requirement | Measure | orientation_indication_remote_controls | Orientation indication for remote movement controls | For control modules where the operator can switch his / her orientation towards the system, the affected control shall indicate the currently selected location of the operator in relation to the system movements.For controls or buttons which have several functions, it shall be clearly indicated to the operator, which function is currently selected. Note: The operator has the possibility to check the subsequent movement. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 03.09.2021 09:50 | Judith Welzel |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
239 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564191 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
240 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563652 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 17.06.2022 09:49 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
241 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563655 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
242 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563654 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
243 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563656 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
244 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563657 | Headline | Measure | collision_detection_rear_side | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
245 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563658 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
246 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563659 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
247 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563671 | Headline | Cause |  | System moves in wrong direction due to technical error (intentional movement): force handles | System moves in wrong direction due to technical error (intentional movement):FD force handle technical errorChassis force handle error (one side or both sides) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
248 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563672 | Requirement | Measure | deactivate_chassis_handle | Deactivate whole chassis handle when error in movement activation is detected | When an error in the movement activation circuit in one side of the chassis handle is detected, movement activation from the whole chassis handle shall be inhibited. Note: This is checked during startup. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 08:21 | Johannes Schmidt |  |  |  | MCU,HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
249 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563673 | Requirement | Measure | check_plausibility_sensors | Plausibility check for sensors in movement control elements | The system shall permanently check the plausibility of analogue sensor outputs for movement control of the force handle sensors and the hand control joystick during movement activation. When a significant implausible deviation is detected, movement initiated from this control location shall be stopped and inhibited and an error message shall be shown as long as the malfunction is present. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 02.03.2022 10:38 | Ingo Hollenborg |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
250 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563683 | Headline | Cause |  | System does not move according to control parameters, does not follow the planned movement path due to external factors | System does not move according to control parameters, does not follow the planned movement path due to external factors (while maintaining speed limits etc.):Grip is suddenly removed, fluids on the floor, partiallyObstacles on the floor: cables,...Uneven floorDoorsteps in the waySystem moves onto an inclined planeSlip of one or more wheels (correct speeds and directions of all drives, contamination of the wheels…), resulting movement direction of system is not correctAdditional components / accessories are mounted on the system | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
251 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563684 | Requirement | Measure | main_unit_movement_shock_detection | Main Unit movement shock detection | The system shall check the movement for significant shocks.If a significant shock is detected, motorized movements shall be stopped (but not inhibited) and an appropriate error message shall be displayed?to the user, so that the user can decide to keep or delete stored positions.?This is relevant in cases when the collision detection of the system has not been activated. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 24.08.2022 10:59 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
252 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563685 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:50 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
253 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564164 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
254 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563687 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
255 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563688 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
256 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563689 | Requirement | Measure | cable_deflectors | Cable deflectors and wheel protection | The wheels shall be equipped with cable deflectors and cover at the upper side to prevent entanglement of external objects in the wheels and prevent external mechanical influence (cables, other disturbing parts which could get caught in the wheels). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
257 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563018 | Headline | Measure | docNav_additional_weight | Reference - Weight restriction mounting external components | <XREF>563734</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
258 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561948 | Headline | Measure | um_weight_limit_additional_components | Reference - Warning - Weight limit for additionally mounted components | <XREF>563735</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
259 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563690 | Headline | Cause |  | System does not move according to control parameters, does not follow the planned movement path due to technical errors | System does not move according to control parameters due to failure of the control system (while maintaining speed limits etc.):Wrong single wheel speed and direction incorrect (e. g. SW error, random HW error…)Wrong path is calculated or incorrectly translated into system movements (SW..)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | B
260 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563691 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 24.08.2022 11:30 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
261 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599809 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:32 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
262 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563692 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:51 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
263 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563693 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
264 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563695 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
265 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563696 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
266 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563697 | Headline | Measure | design_of_wheels | Reference - Mechanical design drive and wheels | <XREF>561959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
267 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568836 | Requirement | Measure | class_shock_vibration_operation | Classification for mechanical shock and vibration - operation | The system shall be designed according to class 3M12 according to IEC 60721-3-3 during operation.Note: Only components located in the main unit are relevant, not in the trolley. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 21.09.2022 12:32 | Johannes Schmidt |  |  |  | Batt,ColS,FDmediumH,GridS,HICTouchCon,MainU,MCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
268 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568848 | Requirement | Measure | docSCM_check_movement_manufacturing | Movement check in manufacturing | There shall be a dedicated check of the movement functionality in manufacturing. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 01.09.2020 | 21.09.2022 12:27 | Johannes Schmidt | [7 empty] | empty | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
269 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568847 | Requirement | Measure | serv_check_movement_service | Movement check after service / maintenance | There shall be a dedicated check of the movement functionality after service / maintenance activity. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 01.09.2020 | 21.09.2022 12:29 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
270 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563698 | Headline | Cause |  | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | System does not move according to control parameters due to failure of the control system (while maintaining speed limits etc.):One drive blocks during movement (e. g. power supply failure and brake gets activated, random HW error, mechanical error of gear, brake is activated,…) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | B
271 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563699 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 24.08.2022 11:34 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
272 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599810 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:35 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
273 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563700 | Headline | Measure | design_of_wheels | Reference - Mechanical design drive and wheels | <XREF>561959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
274 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563702 | Headline | Measure | cable_deflectors | Reference - Cable deflectors and wheel protection | <XREF>563689</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
275 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563751 | Headline | Cause |  | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | System does not move according to control parameters during stopping:One wheel does not follow QuickstopMechanical failure of the brake (when Quickstop is not initiated correctly)One wheel slips resulting in non-path movement until standstillAdditional components / accessories are mounted on the system | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | A
276 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563755 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
277 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563756 | Headline | Measure | docNav_additional_weight | Reference - Weight restriction mounting external components | <XREF>563734</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
278 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563757 | Headline | Measure | um_weight_limit_additional_components | Reference - Warning - Weight limit for additionally mounted components | <XREF>563735</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | empty | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
279 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563752 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 24.08.2022 11:36 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
280 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571853 | Headline | Cause |  | Brakes have been opened manually, subsequent motorized movement is performed with brakes open, arbitrary movement of axle |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | A
281 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571854 | Requirement | Measure | activate_other_brakes | Activate other brakes for motorized movement | If a motorized movement is requested, the movement shall only be started after the brakes of all non-relevant axes are activated. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 16.10.2020 | 19.07.2022 17:39 | Johannes Schmidt |  |  |  | MCU | Machinery directive (EHSR); 1.2.2 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
282 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580821 | Headline | Cause |  | During braking process, orbital and angular axle are moved unintentionally due to inertia (overridden brakes) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | A
283 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580822 | Headline | Measure | braking_force | Reference - Braking force | <XREF>564086</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU | SFSIEC 60601-2-54; 201.9.2.3.1.101 Unintended movement during PATIENT loading/unloading |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
284 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563704 | Headline | Hazard |  | Complete loss of movement functionality (e. g. chassis) (HS_110) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
285 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563706 | Headline | Cause |  | Chassis movement is blocked: E. g. Brakes cannot be unlocked due to power supply failure, motor not functional |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_110 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
286 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562205 | Requirement | Measure | docUser_backup_system | Warning - Establish emergency procedures | WARNING:Cause: System functions including X-ray imaging and system movements get disturbed or get unavailable due to component failures or failures of the external power supply. Possible consequences: Disturbance or unavailability of system functions. Interventions cannot be started or continued.Remedy: Consider establishing emergency procedures to manage situations related with the failures of the system, e. g. keeping a reserve system available. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
287 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563718 | Headline | Cause |  | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | Chassis movement is blocked:Collision sensor remains activated in collisionMore than one collision sensor is activated at the same time, system is blocked | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_110 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
288 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563721 | Headline | Measure | behaviour_collision | Reference - Behaviour in collision | <XREF>563719</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
289 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571892 | Headline | Measure | behaviour_collision | No additional measures are required | In case the collision state cannot be removed by one control (e. g. error in joystick deflection), an alternative control can be used to maneuver the system out of collision zone. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
290 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580797 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
291 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580798 | Headline | Cause |  | Unintended movement activation with activated e-stop: Operator releases manual brakes during active e-stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
292 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580799 | Headline | Measure |  | Reference - Behaviour in collision | No risk identified, no measure required, since with activated e-stop, the STO is still active, so that a restart of the drives is prohibited safely. The operator is allowed to open the brakes and perform manual movements. During motorized movements, it is not possible to open the brakes manually. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
293 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563705 | Headline | Hazard |  | Nonrecoverable or long-term loss of imaging functionality (HS_090) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
294 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563707 | Headline | Cause |  | System gets damaged due to a previous collision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_090 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
295 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563708 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
296 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563741 | Headline | Cause |  | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_090 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined
297 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563742 | Headline | Measure | prevent_tension_tc_cable | Reference - Prevent tension at trolley-chassis cable | <XREF>563740</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
298 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563709 | Headline | Hazard |  | Critical influence on external electric devices (HS_011) | Critical influence on external electric devices (e.g.magnetostatic fields generated by the system exceedthe defined limits for the safe operation ofPacemakers/ICDs) (HS_011) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC),Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
299 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563710 | Headline | Cause |  | Electro-magnetic emissions of the drives in combination with the other emissions of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_011 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
300 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563712 | Headline | Measure | EMC_s1 | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,DapS,DSMonS,FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
301 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562061 | Requirement | Measure | docUser_note_uncommon_system_behavior | Warning - EMC disturbances of other equipment | Note: Warning that, even when the system complies with 60601-1-2, other interferences could result in functional disturbances:WARNINGCause: The electromagnetic radiation of the system may cause interference.Risk: This interference may disturb or disrupt the operation of nearby equipment.Remedy:- Examinations involving external equipment require that there are healthcare professionals present, who are trained for managing the malfunctions of these external equipment and associated emergency situations.- In case of disturbances in external equipment operation, try to re-orient or re-locate these.- To prevent such disturbances, it may be necessary to shield the external equipment/location. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
302 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563714 | Headline | Hazard |  | Non-critical influence on external electric devices (HS_010) | Non-critical influence on external electric devices'function (e.g. the electromagnetic emission of thesystem exceeds the defined limits for disturbance-freeoperation of recording systems) (HS_010) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC),Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
303 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563715 | Headline | Cause |  | Electro-magnetic emissions of the drives in combination with the other emissions of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_010 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
304 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563716 | Headline | Measure | EMC_s1 | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,DapS,DSMonS,FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
305 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563717 | Headline | Measure | docUser_note_uncommon_system_behavior | Reference - Warning - EMC disturbances of other equipment | <XREF>562061</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
306 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563738 | Headline | Hazard |  | Contact with high level electrical energy (HS_002) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
307 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563739 | Headline | Cause |  | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_002 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
308 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563740 | Requirement | Measure | prevent_tension_tc_cable | Prevent tension at trolley-chassis cable | The system shall detect any tensile load on the trolley-chassis cable. When a hazardous tensile load is detected, motorized chassis movements shall be stopped before any damage to the system or any influence on the movement trajectory occurs. Subsequent movements may only be initiated with collision speed <XREF>576777</XREF>. At the same time, the torque supervision <XREF>563722</XREF> shall be applied.An appropriate message or indication shall be displayed to the operator.?Note 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. Note 2: When the trolley is moved (manually always), there is direct feedback to the operator who pushes or pulls the trolley, so that it is not likely that any damage occurs due to that, also based on the field information based on previous systems. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.08.2020 | 17.06.2022 09:04 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
309 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576777 | Requirement | Measure | behaviour_tension_detection_tc_cable | Behaviour with activated trolley-chassis cable tension detection | With activated trolley-chassis cable tension detection (that means after stopping and new movement initiation), the system shall only allow motorized chassis movements with reduced speed. The behaviour shall be the same as in a collision detection situation.<XREF>563490</XREF> Note: In case the operator cannot relieve the tension with motorized movements, he / she can unplug the cable, move the cable, untangle the cable or the setup, or move the trolley. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 18.01.2021 | 09.02.2022 13:43 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
310 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576778 | Headline | Cause |  | System is moved away from the trolley, trolley-chassis cable tension detection is not working |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 18.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_002 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
311 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576779 | Requirement | Measure | self-test_tension_detection_tc_cable | Self-test trolley-chassis cable tension detection | The system shall perform a self-test of the trolley-chassis cable tension detection as long as the cable is plugged in.If an error is detected the behavior shall be the same as if cable tension is detected (see <XREF>576777</XREF> and <XREF>563740</XREF>). | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 18.01.2021 | 24.02.2022 13:18 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
312 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563623 | Headline | Hazard |  | Short-term delay of the clinical procedure (HS_092) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
313 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563624 | Headline | Cause |  | Collision is detected, the system is stopped, but the operator is not aware of the situation and is confused about the stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
314 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563626 | Requirement | Measure | collision_state_indication | Collision state indication | When a collision is detected, the system shall indicate the state to the operator. Note: This refers to collisions detected by collision detection sensors. For Chassis & FD handle see <XREF>563491</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.08.2020 | 26.09.2022 08:28 | Johannes Schmidt |  |  |  | MCU,HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
315 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563743 | Headline | Hazard |  | Severe limitation of movement functionality (HS_111) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
316 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563744 | Headline | Cause |  | Collision sensor fails during any other operation (direction movement, standstill,…): permanently activated collision sensor | Collision sensor fails during any other operation (direction movement, standstill,…)cable breakage, mechanical defect, that results in permanently activated collision sensor | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
317 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563745 | Headline | Measure |  | No additional measure required |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
318 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576792 | Headline | Hazard |  | Unnecessary exposure to low levels of ionizing radiation (HS_060) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
319 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576793 | Headline | Cause |  | Regulatory source to skin distance is decreased when collision detection mechanism of X-ray source cover is activated |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_060 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | A | Low | A
320 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576794 | Headline | Measure |  | No measure required | IEC 60601-2-54, clause 203.9 requires a minimum focal spot to skin distance (aka next surface which mechanically prevents that the operator cannot place any anatomy closer to the X-ray source, e. g. X-ray window) of 200 mm.Theoretically, it is possible that the anatomy is positioned directly on the X-ray window so that the focal spot to skin distance would be violated. However, this is not considered as reasonably foreseeable sequence of events. Therefore, it is concluded, that no significant risk is related to the design. No additional measures are required and the design is acceptable. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 20.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | IEC 60601-2-54 cl. 203.9 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
321 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564220 | Headline | Function |  | Perform movements via wireless hand control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
322 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564221 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
323 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564223 | Headline | Cause |  | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | Operator activates system movement via wireless remote control not pertaining to the system near the other system it pertains to:Interchanged controls controls due to cleaning or moving the systems between different rooms | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
324 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564224 | Requirement | Measure | pairing_wireless_hand_control | Pairing of hand control and system | The hand control and the system shall be paired. The pairing shall be performed by an interaction on the control and on the system side within a defined time to ensure the paring of specified devices only. Only pairs shall operate together. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 01.08.2022 15:51 | Ingo Hollenborg |  |  |  | HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
325 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564225 | Requirement | Measure | optical_indication_wireless_pair | Optical indication of system and hand control pairs | There shall be an optical indication on the system and the hand control, e. g. colored symbols, indicating that they belong together. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 15.06.2022 14:18 | Johannes Schmidt |  |  |  | MainU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
326 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596993 | Requirement | Measure | docSCM_label_wireless_hand_control | Label wireless hand control | Factory Documentation: Attach a label on the wireless hand control to which system it belongs to. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 23.06.2022 | 23.06.2022 13:53 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
327 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600715 | Requirement | Measure | docServ_optical_indication_wireless_collision | Optical indication of system and hand controls - collision | In Service Documentation?WARNINGCause: Movement released from a remote hand control in another roomPossible Consequences: Collision of system with persons or objectRemedy: Before performing any service activities with the remote hand control, check that the remote hand control belongs to the system and check that the indicating label is in place and properly visible. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 20.09.2022 11:12 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
328 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600716 | Requirement | Measure | docUser_labeling_wireless_hand_control_collision | Labeling wireless hand control - collision | WARNINGCause: Movement released from a remote hand control in another roomPossible Consequences: Collision of system with persons or objectRemedy: As part of the system installation, the wireless hand control is labeled according to which system it belongs to.When utilizing the wireless hand control, check that it belongs to the system according to the indicating labels and ensure there is visual contact with the system. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 21.09.2022 08:21 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
329 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564234 | Headline | Cause |  | Control is dropped and function gets activated |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
330 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564235 | Headline | Measure | flat_control_elements | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
331 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564237 | Headline | Cause |  | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | (when this emergency stop is activated by the operator) due to: Wireless hand control loses RF link, e. g. person takes hand control and walks away.Wired hand control non-functional or unplugged | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
332 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564238 | Requirement | Measure | deactivate_movement_radiation_wireless | Deactivate movement and radiation at wireless hand control receiver | The movement and radiation request shall be deactivated at the wireless hand control receiver, when the receiver looses its connection to the sender or when the communication is disturbed (independent of a potential running movement or radiation initiated from this control).If the wireless hand control receiver becomes non-functional, movement and radiation shall be deactivated from this control.In case an emergency stop has been activated before connection loss, it shall remain activated by the receiver.?Note: Safety architecture "mixed", see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 05.10.2022 17:01 | Peter Hefter |  |  |  | MCU,HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
333 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 595749 | Requirement | Measure | deactivate_movement_radiation_wired_HC | Deactivate movement and radiation at wired hand control | If the wired hand control is disconnected or becomes non-functional, movement and radiation shall be deactivated from this control.In case an emergency stop has been activated before unplugging or non-functionality of the wired hand control, it shall remain activated. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 31.03.2022 | 06.04.2022 09:25 | Ingo Hollenborg |  |  |  | MCU,HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
334 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564239 | Requirement | Measure | indicate_e-stop_functionality_wireless | Indication of e-stop functionality | The system shall indicate on the wireless hand control and on the user interface (if available, not available in battery mode) if the emergency stop of the wireless hand control is functional.It is acceptable if the wireless hand control uses e.g., a blinking LED for displaying the functionality. Note, there are two independent time outs, one for the hand control (for detecting the connection status and indicating it) and one at the receiver (for detecting the connection status and indicating it on system side). | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 11:51 | Johannes Schmidt |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
335 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571957 | Requirement | Measure | message_wireless_control_drawer | User message about non-functional wireless hand control and instruction to put it in the drawer | A message shall be displayed when the wireless hand control becomes non-functional or connection gets lost which shall instruct the operator to put the control into the drawer, and which shall be confirmed by the operator. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.11.2020 | 06.04.2022 09:25 | Ingo Hollenborg |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
336 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 595744 | Requirement | Measure | message_wired_control_removal | User message if wired hand control is non-functional | A message shall be displayed when the wired hand control is non-functional and shall instruct the operator to put the wired hand control away. The message shall be confirmed by the operator. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Johannes Schmidt |  |  | 31.03.2022 | 25.04.2022 14:52 | Peter Hefter |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
337 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564240 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 31.03.2022 11:21 | Johannes Schmidt |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
338 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564241 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
339 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564242 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
340 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578630 | Headline | Cause |  | Emergency stop on the wireless hand control is not functional when control goes to sleep mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.02.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
341 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578632 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>Emergency stop on wireless hand control is functional when control is in sleep mode (with only slightly increased delay time). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
342 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564246 | Headline | Cause |  | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged, e-stop is deactivated automatically and system starts moving (again) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
343 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564247 | Headline | Measure | deactivate_movement_radiation_wireless | Reference - Deactivate movement and radiation at wireless hand control | <XREF>564238</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
344 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564248 | Headline | Measure |  | Reference - Alternative deactivation of emergency stop - hand control | <XREF>564251</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
345 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568773 | Headline | Cause |  | Error in wireless data transmission: Delay, loss of data packet(s) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
346 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568776 | Headline | Measure | deactivate_movement_radiation_wireless | Reference - Deactivate movement and radiation at wireless hand control receiver | <XREF>564238</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
347 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568774 | Headline | Cause |  | Error in wireless data transmission: Wrong sequence of data packets, repetition of invalid packets |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
348 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568779 | Headline | Measure | safe_wireless_transfer | Reference - Safe wireless transfer of safety relevant information | <XREF>564349</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICHandControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
349 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568775 | Headline | Cause |  | Error in wireless data transmission: Wrong information transmitted |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
350 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564349 | Requirement | Measure | safe_wireless_transfer | Safe wireless transfer of safety relevant information | The transfer of safety relevant movement and radiation signals / information from wireless remote controls shall be safe and protected against transmission errors (ensure integrity of packets including sequence).When an integrity error is detected, the movement and radiation enable shall be deactivated at the wireless hand control receiver (independent of a potential running movement or radiation initiated from this control).In case an emergency stop has been activated when the error is detected, it shall remain activated by the receiver. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 13.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
351 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576770 | Headline | Cause |  | Error in wireless data transmission: Sleeping error in protection mechanisms of wireless transfer |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
352 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576771 | Requirement | Measure | self-test_safe_wireless_transfer | Self-test safe wireless transfer components | A self-test of the components required for protection mechanisms for safe wireless transfer shall be performed at system startup. This includes the protection mechanism of the receiver as well. In case components cannot be self-tested, the component shall be implemented redundantly. Note: For system reaction see <XREF>564351</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.01.2021 | 29.06.2022 14:53 | Johannes Schmidt |  |  |  | HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
353 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578173 | Requirement | Measure | safe_power_supply_wireless_hand_control | Safe power supply for safe wireless transfer components | The components required for protection mechanisms for safe wireless transfer shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. Note: For system reaction see <XREF>564351</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.02.2021 | 29.06.2022 14:58 | Johannes Schmidt |  |  |  | HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
354 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568777 | Headline | Cause |  | Error in wireless data transmission: Internal error in receiver |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
355 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564351 | Requirement | Measure | safe_state_wireless_receiver | Safe state wireless hand control receiver | When the wireless hand control receiver detects an internal problem (self-test, watchdog), the receiver shall go to a safe state: X-ray and movement enable deactivated and emergency stop activated. ?Note: Safety architecture "mixed", see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 13.08.2020 | 17.06.2022 09:05 | Johannes Schmidt |  |  |  | HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
356 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568778 | Headline | Cause |  | Error in movement release circuit of the wireless control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
357 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578324 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 08.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
358 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578325 | Headline | Measure | check_plausibility_sensors | Reference - Plausibility check for sensors in movement control elements | <XREF>563673</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 08.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
359 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568780 | Headline | Cause |  | Wireless control comes into reach or connect. reestablished and mov. button is activ. or tech. error in mov. release circuit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
360 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568781 | Requirement | Measure | check_movement_release_AtHotPlug_s1 | Block movement when controls connecting | The system shall check the status of both movement release signal paths when controls are connected to the system.If an active movement release signal path is detected at hot plug, movements from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both movement release paths take place. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 24.02.2022 13:23 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
361 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591782 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 02.09.2021 15:45 | Judith Welzel | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,slips, lapses and mistakes,Incorrect or inappropriate output,Software failure,Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
362 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591783 | Headline | Cause |  | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates non-correctly calibrated control | Operator activates wireless remote control pertaining to another system, but joystick is calibrated for the original system it belongs to. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
363 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591784 | Headline | Measure | pairing_wireless_hand_control | Reference - Pairing of hand control and system | <XREF>564224</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 02.09.2021 15:50 | Judith Welzel |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
364 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591785 | Headline | Measure | optical_indication_wireless_pair | Reference - Optical indication of system and hand control pairs | <XREF>564225</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 02.09.2021 15:51 | Judith Welzel |  |  |  | MainU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
365 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564226 | Headline | Hazard |  | Short-term delay of the clinical procedure (HS_092) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined
366 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564227 | Headline | Cause |  | Wireless remote control is not functional: Battery runs empty |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
367 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564228 | Requirement | Measure | check_battery_charge_wireless | Check battery charge of wireless remote control | The system shall check the battery charge of the wireless remote control at start-up shall warn the operator when the remaining battery charge would not be sufficient for one day of system operation with the expected use profile. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.02.2022 16:50 | Miriam Moreno Pradillo |  |  |  | MCU,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
368 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564229 | Requirement | Measure | wired_control_mandatory | Wired remote control mandatory | The system shall always be equipped with a wired remote control when it is equipped with a wireless remote control (option). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Configuration Management | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
369 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564230 | Headline | Cause |  | Wireless remote control is not functional: Liquid intrudes into the electronics |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
370 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564231 | Requirement | Measure | protection_class_remote_controls | IP protection class remote controls | The wireless and the wired remote control shall have protection class IPx5 according to EN 60529 with closed battery compartment. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICHandControl |  |  |  | Basic | System Integration Test | empty | empty | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
371 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568818 | Requirement | Measure | correct_exchange_batteries_wireless | Caution - Correct change of batteries wireless remote control | CAUTIONCause: Incorrect exchange of batteries of the wireless remote control.Possible consequences: Wireless remote control not functional.Remedy: Correct exchange of batteries with dry hands and dry control. Do not let the battery compartment open, close it as soon as you are finished, e. g. also for storing the control. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.10.2022 09:54 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
372 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564232 | Headline | Cause |  | Wireless remote control is not functional: Control is dropped and gets damaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | C
373 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564233 | Headline | Measure | flat_control_elements | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
374 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564243 | Headline | Cause |  | Emergency stop on the wired / wireless remote control is activated unintentionally |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
375 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564244 | Requirement | Measure | protection_unintended_activation_e-stop | Prevent unintended activation of e-stop | The emergency stops shall be protected against unintended activation (for the wired and wireless remote control e. g. mounting within a recessed surface of the surrounding control, but no closed mechanical shroud; for other emergency stops e. g. mounted away from areas which are foreseeably unintentionally touched etc.). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
376 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564245 | Requirement | Measure | indication_activated_e-stop | Indication of e-stop activation and location | When an emergency stop is activated, the system shall indicate that an emergency stop is activated and at which location.In battery mode it is acceptable, that only the activation state of emergency stop is indicated, and not the location. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 27.07.2022 09:10 | Johannes Schmidt |  |  |  | MCU,HICForceControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
377 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591780 | Headline | Cause |  | Wireless remote control comes into wireless range or becomes functional again with a pressed emergency stop | and interrupts any movement | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_092 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
378 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591781 | Headline | Measure |  | No additional measure required | When the wireless hand control comes into wireless range again or becomes functional again after being out of the range or non-functional, the wireless hand control is not accepted as a valid control when the emergency stop is activated at that time due to regulatory reasons.This prevents that an ongoing movement is interrupted by an unintentional emergency stop activation while the wireless hand control is out of reach or not functional (emergency stop pressed when control is carried in the pocket of a coat, or when the batteries are exchanged etc.). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 02.09.2021 14:53 | Judith Welzel |  |  |  | MainU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
379 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564249 | Headline | Hazard |  | Severe limitation of movement functionality (HS_111) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | Patient | S1M | VB10 | [12 empty] | NotDetermined
380 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564250 | Headline | Cause |  | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged, e-stop remains active, cannot be deactivated and blocks further motorized movements | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
381 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564252 | Headline | Measure | deactivate_movement_radiation_wireless | Reference - Deactivate movement and radiation at wireless hand control receiver | <XREF>564238</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICTouchCon,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
382 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564251 | Requirement | Measure | alternative_deactivation_e-stop | Alternative deactivation of emergency stop - hand control | There shall be an alternative safe way to deactivate an emergency stop of the wireless and the wired hand control with a deliberate action (two subsequent actions) by the operator.For the wireless hand control: Only if the hand control connection is lost with an activated e-stop by this control: After the connection loss, an activation and deactivation must take place at another emergency stop.For the wired hand control: Only if the wired hand control functionality is lost with an activated e-stop by this control: Unplugging the wired hand control and thereafter activating and deactivating another emergency stop.Note 1: The single actions of unplugging the wired hand control or putting the wireless hand control in the drawer shall not deactivate the emergency stop.If a connection is re-established and the e-stop has been released in the meanwhile at another device, the e-stop shall not be present. If the e-stop has not been released in the meanwhile at another device, then the e-stop shall remain present until resolved. Note 2: This is deemed to sufficiently prevent the scenario, that an e-stop is deactivated unintentionally when e. g. two operators activate their e-stops nearly simultaneously (followed by a deactivation of the second one), since it is deemed very unlikely that the sequence connection loss, and e-stop already is activated, another e-stop is activated for another purpose or unintentionally than to deactivate the e-stop in general. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 14.06.2022 17:10 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
383 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569231 | Headline | Function |  | Perform movements via wired hand control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
384 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569232 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
385 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562466 | Headline | Cause |  | Disconnect wired hand control with activated emergency stop, system starts moving |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Test | NEW | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
386 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571978 | Headline | Measure | alternative_deactivation_e-stop | Reference - Alternative deactivation of emergency stop - hand control | <XREF>564251</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 09.11.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
387 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564077 | Headline | Function |  | Activate emergency stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
388 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564078 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
389 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564079 | Headline | Cause |  | Control system does not brake / brakes too slowy (Quickstop not initiated or wrong braking ramp) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
390 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564080 | Headline | Measure | safe_stop_motorized_movements | Reference - Safe stop motorized movements | <XREF>563749</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 24.03.2022 14:39 | Nils Naarmann |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
391 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564081 | Headline | Cause |  | After activating the emergency stop, movements stop, operator deactivates emergency stop, movement is initiated again | After activating the emergency stop (e. g. stuck movement key, not noticed), movements stop, operator deactivates emergency stop again, movement is initiated again | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
392 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564082 | Requirement | Measure | movement_after_estop_deactivation | No automatic movement release after e-stop deactivation | If the move request signal is still active after the emergency stop is deactivated again, the system shall block movements until a normal deactivation of the enable signal and re-activation took place. See <XREF>564245</XREF> for user messaging for emergency stop. Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.´ For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 29.09.2022 14:15 | Johannes Schmidt |  |  |  | MCU | SFSIEC 60601-1; 9.2.4 Emergency stopping devices | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
393 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562460 | Headline | Cause |  | Emergency stop not functional |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
394 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563483 | Requirement | Measure | SelfTestEmergencyStop | Emergency stop self-test and failure exclusions | The emergency stop circuit shall be checked during start-up. In case the test is not passed, motorized movements shall remain blocked. Failure modes which cannot be self-tested or prevented shall be covered by failure exclusions.?Note 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. Note 2: This mechanism does not cover mechanical failures of the emergency stop switches, therefore see <XREF>563484</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 17.06.2022 09:06 | Johannes Schmidt |  |  |  | MCU,HICHandControl | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
395 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563484 | Requirement | Measure | um_check_e-stop | Warning - Check emergency stop | WARNINGCause: Emergency Stop button not functioning.Risk: Risks of injury by crushing or collision, system could stop too late in a critical situation.Remedy: Perform the check of the emergency stop buttons on a daily basis according to the steps below. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
396 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571900 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 21.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
397 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571898 | Headline | Cause |  | After activating the emergency stop, movement is initiated again, sleeping error in P path |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 21.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
398 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571899 | Headline | Measure | self-test_protection_hardware | Reference - Self-test of protection module hardware | <XREF>571864</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 21.10.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
399 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564083 | Headline | Function |  | Maintain stable system position |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
400 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564085 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Fatigue failure |  | S1M | VB10 | [12 empty] | NotDetermined
401 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564084 | Headline | Cause |  | System moves unintentionally during standstill |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
402 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564086 | Requirement | Measure | braking_force | Braking force | The braking force needs to be sufficient to hold the system in standstill on an inclined plane of 5° in non-transport position (worst case position of axles), and of 10° in transport position.The braking force needs to be sufficient to hold the system in standstill during 3D. During any stopping scenario of the mecanum platform, the orbital and angular axle shall not move more than +/- 5° (brakes overridden due to inertia). Hint: Definition of ramp inclinations (5° and 10°) were in alignment with IEC 60601-1, clause 9.4.2.2 | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 12.05.2022 18:03 | Johannes Schmidt |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
403 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564087 | Requirement | Measure | self-test_brakes | Self-test of the brakes | An automatic brake test for mecanum wheels shall be performed at first chassis movement in order to detect if a brake does not work properly. This shall include the ability of the protection module to activate the brakes. The test shall be performed before a motorized movement takes place. If the test is not passed, the chassis movements shall remain blocked. ?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 17.06.2022 09:07 | Johannes Schmidt |  |  |  | MCU | SFSFollow-up hazard: 06H571856 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
404 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571856 | Headline | Cause |  | System moves unintentionally during brake test |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 19.10.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
405 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594854 | Headline | Measure | protect_unintentional_activation_rescue_user | Reference protect unintentional activation rescue user | <XREF>564146</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 18.02.2022 | 18.02.2022 11:14 | Johannes Schmidt |  |  |  | MainU |  |  |  | Basic | empty | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
406 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594855 | Headline | Measure | inhibit_motorized_movements_rescue | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 18.02.2022 | 18.02.2022 11:13 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VxFUTURE | [12 empty] | NotDetermined
407 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564152 | Headline | Function |  | Perform position movement: Park and Return |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 02.02.2022 16:46 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
408 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564153 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
409 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564154 | Headline | Cause |  | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
410 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564155 | Requirement | Measure | positioning_accuracy | Positioning accuracy | The absolute precision for chassis positioning over a travel range of 1 m shall be at least 2 cm. The repetitive accuracy for chassis repositioning over a travel range of 1 m (back- and forth movement) shall be at least 2 cm. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 29.09.2022 15:17 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
411 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564156 | Headline | Cause |  | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | Technical error in control systemWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
412 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564158 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 24.08.2022 11:38 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
413 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599811 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:39 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
414 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564159 | Headline | Cause |  | Deviation of target position too large: External influences (slip, obstacles,…) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
415 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564160 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 24.08.2022 11:40 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
416 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564161 | Headline | Cause |  | Operator has selected an incorrect target position from position storage or is not aware of selection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
417 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564162 | Headline | Measure | movement_preview | Reference - Movement preview for position movements | <XREF>563648</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
418 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564193 | Requirement | Measure | timeout_position_movement | Timeout for position movement selection | After at most 60 seconds of motorized standstill after selection of a target position for any position movement, there shall be a timeout deselecting the selected target position automatically. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
419 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569216 | Requirement | Measure | movement_range_limit_park_return | Movement range limit park and return | The distance between starting and end point of the wheel platform for park and return movements shall be limited to 4 m. The maximum wheel platform rotation shall be limited to 135°. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.09.2020 | 09.06.2022 16:32 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
420 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564170 | Headline | Measure | um_collisions_position_movements | Reference - Warning - Collisions during position movements without test run | <XREF>563650</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
421 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564192 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
422 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564163 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 17.06.2022 09:51 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
423 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564165 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
424 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564166 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
425 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564167 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
426 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564168 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
427 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564171 | Headline | Cause |  | Operator initiates park and return on outdated / invalid position data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
428 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564172 | Requirement | Measure | prerequisites_park_and_return_position_memory | Prerequisites park and return and position memory | A park and return or position memory movement shall only be possible when the system can reconstruct the movement path (e. g. when the system is rebooted, when the rescue function has been activated, the function is not available). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
429 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564198 | Headline | Function |  | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
430 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564199 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
431 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564200 | Headline | Cause |  | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
432 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564201 | Headline | Measure | repetitive_accuracy_positioning | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
433 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564202 | Headline | Cause |  | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | Technical error in control systemWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
434 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564203 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:41 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
435 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599812 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:41 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
436 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564205 | Headline | Cause |  | Deviation of target position too large: External influences (slip, obstacles,…) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
437 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564206 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:43 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
438 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564207 | Headline | Cause |  | Operator has selected an incorrect target position from position storage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
439 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564208 | Headline | Measure | movement_preview | Reference - Movement preview for automatic movements | <XREF>563648</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
440 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564209 | Headline | Measure | timeout_automatic_movement | Reference - Timeout for automatic movement selection | <XREF>564193</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | empty | empty | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
441 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564219 | Requirement | Measure | movement_range_limit_position_memory | Movement range limit position memory | The distance between starting and end point for position memory movements of the wheel platform shall be limited to 4 m. The maximum wheel platform rotation shall be limited to 135°. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 09.06.2022 16:42 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
442 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564210 | Headline | Measure | um_collisions_automatic_movements | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
443 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564211 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
444 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564212 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 17.06.2022 09:52 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
445 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564213 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
446 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564214 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
447 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564215 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
448 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564216 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
449 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564217 | Headline | Cause |  | Operator initiates movement to stored position on outdated / invalid position data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
450 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564218 | Headline | Measure | prerequisites_park_and_return_position_memory | Reference - Prerequisites park and return and position memory | <XREF>564172</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
451 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564173 | Headline | Function |  | Move along center beam (chassis and lift simultaniously) |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 10.02.2022 16:20 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
452 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564174 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
453 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564175 | Headline | Cause |  | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
454 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564176 | Headline | Measure | repetitive_accuracy_positioning | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
455 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564177 | Headline | Cause |  | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | Technical error in control systemWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
456 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564178 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 24.08.2022 11:44 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
457 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599813 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:44 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
458 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594826 | Headline | Measure |  | Reference - no movement takeover across controls | <XREF>569307</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 10.02.2022 | 08.03.2022 14:57 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
459 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594827 | Headline | Measure |  | Reference - only one active control element | <XREF>563634</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 10.02.2022 | 08.03.2022 14:58 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
460 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564179 | Headline | Cause |  | Deviation of target position too large: External influences (slip, obstacles,…) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
461 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564180 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 24.08.2022 11:46 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
462 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564181 | Headline | Cause |  | Operator has selected an incorrect movement function | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
463 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564182 | Headline | Measure | movement_preview | Reference - Orientation indication for remote movement controls | <XREF>563651</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 03.09.2021 09:44 | Judith Welzel |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
464 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564183 | Headline | Measure | um_collisions_automatic_movements | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
465 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564195 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
466 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564184 | Headline | Measure | set_speed_values | Reference - Set speed values | <XREF>562463</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
467 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564185 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
468 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564186 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
469 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564187 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
470 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564188 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
471 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564310 | Headline | Function |  | Perform position movement: Click and Go (chassis and lift simultaneously) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
472 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564311 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
473 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564312 | Headline | Cause |  | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
474 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564313 | Headline | Measure | repetitive_accuracy_positioning | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
475 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564314 | Headline | Cause |  | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | Technical error in control systemWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
476 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564315 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:47 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
477 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599814 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:48 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
478 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564316 | Headline | Measure | lock_other_movements_applications | Reference - no movement takeover across controls | <XREF>569307</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 10.02.2022 16:58 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
479 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594828 | Headline | Measure |  | Reference - only one active control element | <XREF>563634</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 10.02.2022 | 08.03.2022 14:59 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
480 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564317 | Headline | Cause |  | Deviation of target position too large: External influences (slip, obstacles,…) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
481 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564318 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:49 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
482 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564319 | Headline | Cause |  | Operator has selected an incorrect target position on the image |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
483 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564320 | Headline | Measure | movement_preview | Reference - Movement preview for automatic movements | <XREF>563648</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
484 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564321 | Headline | Measure | timeout_position_movement | Reference - Timeout for position movement selection | <XREF>564193</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | empty | empty | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
485 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569237 | Requirement | Measure | movement_range_limit_click_go | Movement range limit click and go | The travel range for click and go movements shall be limited to 1 m. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 14.09.2020 | 18.02.2022 12:10 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
486 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564322 | Headline | Measure | um_collisions_automatic_movements | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
487 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564323 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
488 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564324 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 17.06.2022 09:53 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
489 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564325 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
490 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564326 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
491 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564327 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
492 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564328 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
493 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564329 | Headline | Cause |  | Operator initiates click and go on outdated / invalid position / image data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
494 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564190 | Requirement | Measure | prerequisites_click_and_go | Prerequisites click and go | Click and go shall only be possible when a patient is registered, an image has been acquired, and when the system can reconstruct the movement path. The function shall consider system movements between image acquisition and function activation, and shall consider image rotation and flip.(I. e. when a previous patient is loaded and an image is loaded again, the feature is not available; after a system reboot or activation of rescue function, the feature is not available.) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 11.02.2022 14:51 | Johannes Schmidt |  |  |  | FLC,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
495 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563440 | Headline | Chapter |  | Move main unit axles - manual | Note, there are currently no risks related to "Move main unit axles - manual" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:21 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
496 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563441 | Headline | Chapter |  | Store position | Note, there are currently no risks related to "Store position" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:21 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
497 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563467 | Headline | Chapter |  | Rescue or transfer patient (e. g. CPR, relocate) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
498 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564126 | Headline | Function |  | Rescue or transfer patient (e. g. CPR, relocate) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
499 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564127 | Headline | Hazard |  | Patient rescue hindered (HS_140) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
500 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564128 | Headline | Cause |  | Chassis is not movable: Power supply, random hardware error (brakes are activated) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_140 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
501 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564129 | Requirement | Measure | provide_rescue_button | Provide rescue button | The system chassis shall be able to be movable manually after an intentional action by the operator (e. g. releasing the brakes via a separate button which is secured against unintended activation) in case the chassis is not movable. The movement shall be possible with reasonable force (max. 400 N (200 N is threshold for mobile equipment propulsion under normal conditions)). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU | IEC 60601-1; 9.2.5 Release of PATIENTIEC 60601-2-43; 201.7.9.2.102 Provisions for CPRIEC 60601-2-43; 201.15.101 Configuration for cardiopulmonary resuscitation (CPR) | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | empty | empty | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
502 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562449 | Requirement | Measure | c-arm_manually_movement | Override angular / orbital brakes manually | The orbital and angular axle shall be designed in a way that the axle is movable manually e. g. in case of a power failure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.04.2022 11:12 | Rudolf Siebenhaar |  |  |  | MainU | When orb. / ang. is functional, the brakes can be unlocked even with activated e-stop (no gravity loaded axles due to weight balance). |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
503 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564130 | Requirement | Measure | use_recue_function | Caution - Use rescue function | CAUTIONRisk: The system cannot be removed from the patient table in case of an emergency situation, e. g. cardiopulmonary resuscitation.Cause: Motorized system movements are blocked.Remedy: Check if any emergency stop has been activated and deactivate it. Use the rescue button in order to move the system away from the patient. Note that the orbital and angular axles still can be moved manually. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.10.2022 09:55 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
504 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564131 | Headline | Cause |  | Chassis is not movable: Mechanical failure of gears or wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_140 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
505 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564132 | Headline | Measure | design_of_wheels | Reference - Mechanical design drive and wheels | <XREF>561959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
506 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564133 | Requirement | Measure | rotate_chassis_blocked_mecanum_wheel | Rotate chassis with blocked mecanum wheel | In case one of the mecanum wheels is permanently blocked, the system shall still be able to be moved / rotated 45 ° with reasonable force (max. 400 N) after activating the rescue function. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
507 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564140 | Headline | Measure | provide_rescue_function | Reference - Provide rescue function | <XREF>564129</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [5 empty] | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
508 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564134 | Headline | Measure | c-arm_manually_movement | Reference - Override angular / orbital brakes manually | <XREF>562449</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
509 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564135 | Headline | Measure | c-arm_manually_movement | Reference - Caution - Use rescue function | <XREF>564130</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
510 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564136 | Headline | Cause |  | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | Chassis movement is blocked:Collision sensor remains activated in collisionMore than one collision sensor is activated at the same time, system is blocked | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_140 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
511 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564137 | Headline | Measure | behaviour_collision | Reference - Behaviour in collision | <XREF>563719</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
512 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564139 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
513 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564141 | Headline | Measure | provide_rescue_button | Reference - Provide rescue button | <XREF>564129</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
514 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564142 | Headline | Measure | c-arm_manually_movement | Reference - Override angular / orbital brakes manually | <XREF>562449</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
515 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564143 | Headline | Measure | c-arm_manually_movement | Reference - Caution - Use rescue function | <XREF>564130</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
516 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564149 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
517 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564150 | Headline | Cause |  | Technical failure in the rescue function activates the function and releases the mecanum brakes, system rolls |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
518 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564151 | Requirement | Measure | protect_unintentional_activation_rescue | Prevent unintentional activation of rescue function by design | The unintentional activation of the rescue function due to technical errors shall be prevented by redundant activation channels (e. g. two switches and safety relays).Only when both channels are activated, the function shall be activated.Inconsistent activation of one channel or unintended permanent activation of both channels shall be checked during startup. In these cases movement shall be inhibited. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 20.09.2022 15:51 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
519 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576703 | Requirement | Measure | serv_check_rescue_function | Check rescue functionality during planned maintenance | Work step during planned maintenance to ensure the correct functionality of the rescue function including correct deactivation of the function after releasing the button. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 21.12.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
520 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564145 | Headline | Cause |  | Operator misuses the rescue function intentionally or unintentionally | e. g. being not aware of consequences, or unintentionally activates motorized movement | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 16:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
521 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564146 | Requirement | Measure | protect_unintentional_activation_rescue_user | Prevent unintentional activation of rescue function by user error | The button for initiating the rescue function shall be protected against unintended activation. The function shall require continuous activation by the operator and the activation signal shall be implemented redundantly.A consistency check of the rescue signals shall be performed at startup in order to detect a first failure in the rescue circuit. If the test is not passed the system shall be blocked. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 28.09.2022 15:32 | Johannes Schmidt |  |  |  | MainU,MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
522 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564147 | Requirement | Measure | inhibit_motorized_movements_rescue | Inhibit motorized movements of chassis - rescue function | If the rescue button is activated during an ongoing movement, the movement shall be stopped by braking (regular braking ramp or quick-stop) in order to prevent spinning out of the drives. Once the movement is stopped the STO shall be activated and movement enable blocked for all motorized axes.Once the rescue button is disengaged the STO is deactivated and movements can be initiated again.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.The SFS functionality is applicable for the breaking functionality, not for recognition and display of of first errors. For design constraints for single fault safety, see <XREF>578311</XREF>.´ | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 11.08.2020 | 26.09.2022 11:21 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
523 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564148 | Requirement | Measure | docUser_correct_use_rescue | Warning - Correct use of rescue function | WARNINGCause: Rescue button is operated incorrectly.Risk: Increased risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment.Remedy: The rescue button is only to be used in emergency situations when immediate access to the patient is required, e. g. for cardiopulmonary resuscitation. Do not operate other motorized movements when rescue button is activated. Do not operate the function during movements.Note that the collision detection and emergency stop will have no effect when this function is activated. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.08.2020 | 18.06.2022 10:52 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
524 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576759 | Headline | Cause |  | Rescue function is activated during activated emergency stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
525 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576760 | Headline | Measure | inhibit_motorized_movements_rescue | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
526 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580834 | Headline | Cause |  | Operator misuses the rescue function intentionally or unintentionally | e. g. activating it during an ongoing motorized movement. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
527 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580835 | Headline | Measure | inhibit_motorized_movements_rescue | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 19.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
528 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576757 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
529 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576758 | Headline | Cause |  | Unexpected arbitrary movement |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
530 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576762 | Headline | Measure | protect_unintentional_activation_rescue_user | Reference - Prevent unintentional activation of rescue function by user error | Continuous activation <XREF>564146</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
531 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576761 | Headline | Cause |  | Rescue function is activated during collision - collision detection ""non-functional"" |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
532 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576763 | Headline | Measure | protect_unintentional_activation_rescue_user | Reference - Prevent unintentional activation of rescue function by user error | Continuous activation <XREF>564146</XREF>, as soon as the button is released, the brakes are activated again (other than non-motorized design with manual brake activation). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
533 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576764 | Headline | Measure | docUser_correct_use_rescue | Reference - Warning - Correct use of rescue function | <XREF>564148</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
534 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563442 | Headline | Chapter |  | Adjust the field of view | Note, there are currently no risks related to "Adjust the field of view" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:22 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
535 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563443 | Headline | Chapter |  | Select and adjust examination settings and imaging parameters | Note, there are currently no risks related to "Select and adjust examination settings and imaging parameters" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:22 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
536 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569262 | Headline | Chapter |  | Acquire images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
537 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563444 | Headline | Function |  | Acquire images - general |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
538 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580656 | Headline | Hazard |  | Unnecessary radiation, dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1I,S1F | VA10,VA20,VA30 | [12 empty] | NotDetermined
539 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580654 | Headline | Cause |  | Radiation is switched off later than required in regulatory limits (IEC 60601-2-54:2018), normal conditions |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1I,S1F | VA10,VA20,VA30 | [12 empty] | NotDetermined | F | Low | F
540 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580655 | Headline | Measure |  | No additional measures required | See feature 505763. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 02.09.2021 15:04 | Judith Welzel | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1I,S1E,S1F,S1P | VA10,VA20,VA30 | [12 empty] | NotDetermined
541 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569260 | Headline | Function |  | Acquire images via wireless hand control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
542 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568782 | Headline | Hazard |  | Unnecessary exposure to low levels of ionizing radiation (HS_060) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
543 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568783 | Headline | Cause |  | Wireless control comes into reach and radiation button is activated or technical error in radiation release circuit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_060 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
544 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562292 | Requirement | Measure | Check_X-RaySwitchAtHotPlug_s1 | Block radiation when controls connecting | The system shall check the status of both radiation release signal paths when controls are connected to the system.When an active radiation release signal path is detected, radiation from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths takes place. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
545 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600717 | Headline | Cause |  | User mistakenly uses a wireless hand control from another system |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_060 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
546 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596992 | Requirement | Measure | docUser_labeling_wireless_hand_control_radiation | Labeling wireless hand control radiation | CAUTIONCause: Radiation released from a wireless hand control in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy: As part of the system installation, the wireless hand control is labeled according to which system it belongs to.When utilizing the wireless hand control, check that it belongs to the system according to the indicating labels. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 23.06.2022 | 07.10.2022 09:56 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
547 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596634 | Requirement | Measure | docServ_optical_indication_wireless_radiation | Optical indication of system and hand controls radiation | In Service Documentation?CAUTIONCause: Radiation released from a remote hand control in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy: Before performing any service activities with the remote hand control, check that the remote hand control belongs to the system and check that the indicating label is in place and properly visible. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 18.05.2022 | 20.09.2022 11:34 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
548 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600718 | Headline | Measure | pairing_wireless_hand_control | Reference pairing wireless hand control | <XREF>564224</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 20.09.2022 11:46 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
549 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600719 | Headline | Measure | optical_indication_wireless_pair | Reference Optical indication of system and hand control pairs | <XREF>564225</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 20.09.2022 11:46 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
550 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 600720 | Headline | Measure | docSCM_label_wireless_hand_control | Reference Label wireless hand control | <XREF>596993</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.09.2022 | 20.09.2022 11:48 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
551 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569279 | Headline | Function |  | Acquire images via wired footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
552 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569277 | Headline | Hazard |  | Medium-term delay of the clinical procedure (HS_091) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
553 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569278 | Headline | Cause |  | Footswitch gets damaged due to ingress of fluid |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_091 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
554 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568819 | Requirement | Measure | protection_class_footswitch | IP protection class footswitch | The wireless and the wired footswitch shall have protection class IPx8 according to EN 60529 with closed battery compartment. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
555 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563447 | Headline | Function |  | Acquire Roadmap images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
556 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591776 | Headline | Hazard |  | Unnecessary exposure to limited levels of ionizing radiation (HS_061) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 20.06.2022 13:54 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
557 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591777 | Headline | Cause |  | Position mismatch between subtraction and roadmap phases | e. g. due to system movements between the different phases, or selecting a peak opacification image from a different position | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_061 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,use by unskilled/untrained personnel,slips, lapses and mistakes,Foreseeable misuse,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
558 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591778 | Requirement | Measure | _sub_road_no_system_movement | No system movement during or between subtraction and roadmap phases | Case 1: Movement during radiation in subtraction or roadmap mode:When a movement of any system axle is detected during any phase of acquiring an image in subtraction or roadmap mode, the workflow shall be aborted, radiation shall be stopped. Case 2: Movement between acquired peak opacification image within a subtraction workflow and roadmap acquisition:When a movement of any system axle is detected between an acquired peak opacification image within a subtraction and a planned roadmap acquisition, the system shall discard this outdated peak opacification image (not allow roadmap), but instead, shall start the subtraction workflow automatically. | ToBeReviewed | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 02.09.2021 | 20.06.2022 14:10 | Johannes Schmidt |  |  |  | FLC,MCU,XCU | IEC 60601-2-43, cl. 201.12.4.105 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
559 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591779 | Requirement | Measure | docUser_prevent_movement_roadmap | Warning - Prevent movements during roadmap workflow | WARNING Cause: Movements of the patient, of the system, or deformations of the vessel tree between acquiring the mask image and using it as mask in the Roadmap workflow. Possible consequences: Treatment at the wrong location.Remedy:- Prevent the movement of the patient’s relevant body region between acquiring the mask and using it as a mask in the Roadmap workflows.- Check whether the device position in the image is matching with the vessel tree.- In case of doubt, visualize the actual vessel position by means of a short contrast injection and check it versus the vessel tree in the mask. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 02.09.2021 | 29.06.2022 15:00 | Johannes Schmidt |  |  |  |  | IEC 60601-2-43, cl. 201.12.4.105 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
560 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563448 | Headline | Function |  | Acquire 3D images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
561 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564089 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Foreseeable misuse,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
562 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564090 | Headline | Cause |  | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | Brakes are too weak, e. g. when 3D is performed (inertial force), e. g. one brake not working (e. g. mechanical failure) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
563 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564091 | Headline | Measure | braking_force | Reference - Braking force | <XREF>564086</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
564 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564092 | Headline | Measure | automatic_brake_test | Reference - Automatic brake test (selftest) | <XREF>564087</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
565 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564102 | Headline | Cause |  | Operator moves chassis (motorized) or other motorized axle during collision check or shortly after collision check |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
566 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564103 | Requirement | Measure | repeat_collision_check3D | Repeat collision check | Motorized movements other than orbital axle of the system during collision check or shortly after completion of the collision check shall be blocked.Note: If the user wants to perform any other motorized movement, the user has to actively abort the collision check / 3D scan workflow. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 22.06.2022 10:59 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
567 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564105 | Headline | Cause |  | Other motorized movement is triggered by technical fault during 3D scan |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_042 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
568 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564106 | Requirement | Measure | block_other_movements3D | Block other axles during 3D scan | During 3D scans, the system shall block all other motorized movements. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
569 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564097 | Headline | Hazard |  | Severe limitation of the imaging functionality (HS_100) | Severe limitation of the imaging functionality (e. g. subtraction, roadmap, 3D not available, medium-level arcing leading to the interruption of scenes), basic imaging still available | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function |  | S1M | VB10 | [12 empty] | NotDetermined
570 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564098 | Headline | Cause |  | Collision sensor fails or is activated during 3D scan, scan is aborted, restart possible | Brakes are too weak, e. g. when 3D is performed (inertial force), e. g. one brake not working (e. g. mechanical failure) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_100 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
571 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564101 | Headline | Measure | collision_detection | Collision detection during 3D | Hint: Collision handling during 3D is equivalent to normal collision detection. An unwanted interruption of a 3D scan via wheel-based sensors is assessed to be highly improbable because during 3D the wheel-base is blocked and no operator staff is in proximity of the C-Arm. Therefore, it does not need to be addressed with further measures. <XREF>563490</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 10.08.2020 | 10.02.2022 10:13 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
572 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564093 | Headline | Hazard |  | Severely reduced image quality (HS_120) | Severely reduced image quality (e. g. artifacts hinder identification of pathological structures or result in misidentification, …), e. g. due to exposure parameter or postprocessing | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function |  | S1M | VB10 | [12 empty] | NotDetermined
573 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564094 | Headline | Cause |  | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | Brakes are too weak, e. g. when 3D is performed (inertial force), e. g. one brake not working (e. g. mechanical failure) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
574 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564095 | Headline | Measure | braking_force | Reference - Braking force | <XREF>564086</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
575 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564096 | Headline | Measure | automatic_brake_test | Reference - Automatic brake test (selftest) | <XREF>564087</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 10.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
576 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578312 | Headline | Hazard |  | Unnecessary exposure to low levels of ionizing radiation (HS_060) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.02.2021 | 15.06.2022 11:52 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Deterioration in function,Potential for use errors triggered by design flaws,ambiguous or unclear device state,user interface,Untrained users,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
577 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578313 | Headline | Cause |  | Movement stops during acquisition |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.02.2021 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
578 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578315 | Requirement | Measure | no_movement_stop_radiation | Stop radiation when movement is stopped | When the system movement stops for any reason (e.g. collision) during a 3D or a panorama scan, radiation shall be stopped. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 06.02.2021 | 15.06.2022 17:03 | Johannes Schmidt |  |  |  | MCU,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
579 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564271 | Headline | Function |  | Acquire Panorama images |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:49 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
580 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564272 | Headline | Hazard |  | Severely reduced image quality (HS_120) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:49 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function |  | S1M | VxFUTURE | [12 empty] | NotDetermined
581 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564273 | Headline | Cause |  | Wrong stiching of images: Increased mechanical tolerances, compromised position resolution, vibrations (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
582 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564274 | Requirement | Measure | trajectory_accuracy | Trajectory accuracy | The required trajectory accuracy shall be reached, which is necessary to provide panorama images with sufficient quality and geometric fidelity.The travel range for this application shall be restricted so that this value is reached. Note: It must be ensured, that the trajectory accuracy is sufficient to be able to perform image based corrections. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 15.06.2022 11:26 | Rudolf Siebenhaar |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
583 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564291 | Requirement | Measure | reduce_vibrations_panorama | Reduce vibrations during panorama | The system shall perform the movement for panorama image acquisition with a vibration free drive profile. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 15.06.2022 11:25 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
584 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564275 | Headline | Cause |  | Wrong stiching of images: Source to object information is not known |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | C | Low | C
585 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564276 | Headline | Measure |  | No additional measure required | The operator has a correction factor to change the source to object information before and after scan (reconstruction can be repeated). Default is isocenter. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:50 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
586 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564285 | Headline | Cause |  | Wrong synchronization at panorama imaging: Technical error | Wrong synchronization at panorama imaging:Technical error in control system, movement and X-ray generation are not synchronized correctlyWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
587 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564286 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:49 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
588 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564288 | Headline | Cause |  | Wrong stiching of images: Software error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | C | Low | C
589 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564289 | Requirement | Measure | save_all_images_panorama | Save all images panorama | All single 2D images of a panorama scan shall be stored. Note: Multiple reconstructions can be performed, as well as review of single slices. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 11.02.2022 14:52 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
590 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564277 | Headline | Hazard |  | Severe limitation of the imaging functionality (HS_100) | Severe limitation of the imaging functionality (e. g. subtraction, roadmap, 3D not available, medium-level arcing leading to the interruption of scenes), basic imaging still available | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:55 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function |  | S1M | VxFUTURE | [12 empty] | NotDetermined
591 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564278 | Headline | Cause |  | Operator initates panorama imaging on any (outdated) image data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_100 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
592 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564280 | Requirement | Measure | prerequisites_panorama_imaging | Prerequisites panorama imaging | The acquisition of a panorama image shall only be possible when- a patient is registered, an image has been acquired on which the planning took place (path planning must have taken place and been completed in standstill)- when the system has not been moved and no image rotation or flip took place between path planning / confirmation and movement activation- no collimation error (collimator must be able to reach the correct slot positions for the panorama slices)- enough storage capacity must be available for the maximum scan length- tube temperature must be sufficiently low to be able to complete the maximum scan length | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 11.02.2022 14:52 | Johannes Schmidt |  |  |  | FLC,MCU,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
593 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569290 | Headline | Cause |  | Vibrations during scan loosens connectors |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) |  |  |  |  | HS_100 |  | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
594 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596956 | Headline | Measure | class_shock_vibration_operation | Reference - class_shock_vibration_operation | <XREF>568836</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 15.06.2022 | 15.06.2022 11:22 | Johannes Schmidt |  |  |  | Batt,ColS,FDmediumH,FDmediumC,GridS,HICTouchCon,MainU,MCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
595 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564281 | Headline | Cause |  | Wrong synchronization at panorama imaging: Technical error | Wrong synchronization at panorama imaging:Technical error in control system, movement and X-ray generation are not synchronized correctlyWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_100 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
596 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564283 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:52 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
597 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564282 | Headline | Cause |  | Wrong synchronization at panorama imaging: External influences (slip, obstacles,...) | Wrong synchronization at panorama imaging:Technical error in control system, movement and X-ray generation are not synchronized correctlyWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) |  |  |  |  | HS_100 |  | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
598 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564290 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:53 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
599 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569287 | Headline | Cause |  | Undetected error in mecanum mechanics ( e. g. one castor is damaged) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) |  |  |  |  | HS_100 |  | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
600 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596954 | Headline | Measure | serv_check_movement_service | Reference - movement check after service/maintenance | <XREF>568847</XREF> | InWork | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 15.06.2022 | 24.08.2022 11:54 | Johannes Schmidt | [7 empty] | Basic | System Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VxFUTURE | [12 empty] | NotDetermined
601 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596955 | Headline | Measure | docSCM_check_movement_manufacturing | Reference - check_movement_manufacturing | <XREF>558848</XREF> | InWork | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 15.06.2022 | 15.06.2022 10:34 | Johannes Schmidt | [7 empty] | empty | System Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
602 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569288 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 24.08.2022 11:54 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
603 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569289 | Headline | Measure | design_of_wheels | Reference - Mechanical design drive and wheels | <XREF>561959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 15.06.2022 11:32 | Johannes Schmidt |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VxFUTURE | [12 empty] | NotDetermined
604 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599815 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:55 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
605 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564292 | Headline | Hazard |  | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:57 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Incorrect or inappropriate output,Foreseeable misuse,Untrained users |  | S1M | VxFUTURE | [12 empty] | NotDetermined
606 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564293 | Headline | Cause |  | User error, planning of movement path is not done correctly (mix-up of directions…) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
607 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564294 | Headline | Measure | movement_preview | Reference - Movement preview for automatic movements | <XREF>563648</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:57 | Johannes Schmidt |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
608 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568854 | Requirement | Measure | restricted_movement_range_panorama | Restricted movement range for panorama imaging | The movement range for panorama imaging shall be restricted to 110 cm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 01.09.2020 | 10.06.2022 08:57 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
609 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564301 | Headline | Measure | timeout_automatic_movement | Reference - Timeout for automatic movement selection | <XREF>564193</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:57 | Johannes Schmidt |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
610 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564295 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 17.06.2022 09:54 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
611 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564296 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:57 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
612 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564297 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:57 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
613 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564298 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:58 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
614 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564299 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:58 | Johannes Schmidt |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
615 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564300 | Headline | Measure | um_collisions_automatic_movements | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:58 | Johannes Schmidt |  |  |  | FLC,MCU |  |  |  | Basic | n.a. | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
616 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564302 | Headline | Cause |  | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
617 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564303 | Headline | Measure | repetitive_accuracy_positioning | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:58 | Johannes Schmidt |  |  |  | MainU,MCU,HICForceControl |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
618 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564304 | Headline | Cause |  | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | Technical error in control systemWrong geometric input (axle positions)Technical error, e. g. wrong movement direction of one axle / drive, resulting in wrong trajectory (e. g. one castor of one wheel is blocked, but the motor control can still manage to bring the correct speed and positions, nevertheless this results in a incorrect movement path due to the resulting forces) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
619 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564305 | Headline | Measure | main_unit_movement_shock_detection | Reference - Main Unit movement shock detection | <XREF>563684</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 12.08.2020 | 24.08.2022 11:56 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
620 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599816 | Headline | Measure | positioning_accuracy | Reference - Positioning accuracy | <XREF>564155</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 24.08.2022 | 24.08.2022 11:57 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
621 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 564306 | Headline | Measure | lock_other_movements_applications | Reference - no movement takeover across controls | <XREF>569307</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 12.08.2020 | 10.06.2022 08:58 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VxFUTURE | [12 empty] | NotDetermined
622 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594829 | Headline | Measure |  | Reference - only one active control element | <XREF>569307</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 10.02.2022 | 10.06.2022 08:59 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
623 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569228 | Headline | Cause |  | Emergency stop too far away when panorama image is acquired via footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VxFUTURE | [12 empty] | NotDetermined | D | Low | C
624 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577223 | Headline | Measure | set_speed_values | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>?and to not exceed the speed limits <XREF>563486</XREF> in normal conditions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.01.2021 | 17.06.2022 09:54 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
625 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577225 | Headline | Measure | collision_detection | Reference - Collision detection | <XREF>563490</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.01.2021 | 10.06.2022 08:59 | Johannes Schmidt |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
626 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577226 | Headline | Measure | speed_limits | Reference - System speed limits | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.01.2021 | 10.06.2022 08:59 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
627 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 577224 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | Multiple emergency stops are available at the system <XREF>562345</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.01.2021 | 10.06.2022 08:59 | Johannes Schmidt |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
628 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563461 | Headline | Chapter |  | Display acquired image(s) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
629 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569299 | Headline | Function |  | Display acquired image(s) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
630 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569300 | Headline | Hazard |  | Slightly reduced image quality (HS_121) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | Patient | S1M | VB10 | [12 empty] | NotDetermined
631 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569301 | Headline | Cause |  | Impaired displayed image quality due to dirt on the touch monitor |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_120 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
632 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569148 | Headline | Measure |  | Reference - Cleaning of the touch monitor | Cleaning of the monitor is covered in cleaning guide.Reference to Caution to follow the cleaning instruction <XREF>562240</XREF>.The monitor must withstand the procedure.Regarding DIN 6868-157: 5 c) Before starting these tests, the viewing surface of the IMAGE DISPLAY DEVICE is to be cleaned in accordance with operating instructions, if needed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC,MainU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
633 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563460 | Headline | Chapter |  | Display dose related information during radiation | Note, there are currently no risks related to "Display dose related information during radiation" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:09 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
634 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563454 | Headline | Chapter |  | Postprocessing - 2D images | Note, there are currently no risks related to "Postprocessing - 2 D images" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:09 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
635 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563455 | Headline | Chapter |  | Postprocessing - Subtraction | Note, there are currently no risks related to "Postprocessing - Subtraction" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:10 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
636 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563456 | Headline | Chapter |  | Postprocessing - 3D | Note, there are currently no risks related to "Postprocessing - 2 D" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:09 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
637 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563457 | Headline | Chapter |  | Postprocessing - Panorama | Note, there are currently no risks related to "Panorama" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:10 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
638 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563436 | Headline | Chapter |  | Save images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
639 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569280 | Headline | Function |  | Save images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
640 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569282 | Headline | Hazard |  | Image data is unavailable or lost (HS_130) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
641 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569283 | Headline | Cause |  | Storage rate of images during acquisition is reduced |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_130 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
642 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594856 | Headline | Measure | full_storage_rate_radiography | Reference - full storage rate radiography | <XREF>578317</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 18.02.2022 | 18.02.2022 12:07 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
643 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563462 | Headline | Chapter |  | Perform 2D navigation | Note, there are currently no risks related to "Perform 2D navigation" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:26 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
644 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563463 | Headline | Chapter |  | Perform 3D navigation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
645 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569284 | Headline | Function |  | Perform 3D navigation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
646 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569285 | Headline | Hazard |  | Contact with high level electrical energy (HS_002) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
647 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569286 | Headline | Cause |  | Damaged 3rd party devices are connected to the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_002 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,Moving parts | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | B | Med | B
648 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568831 | Requirement | Measure | electrical_safety_ext_interfaces | Electrical safety external interfaces | Galvanic separation (voltage withstand 4 kV) shall be provided for the following external interfaces: * HDMI * Network interface * Injector interface | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | SysCabl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
649 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568833 | Requirement | Measure | electrical_safety_ext_usb | Warning - Electrical safety USB | WARNINGCause: Connect damaged component with external power supply to USB interface.Possible consequences: Electric shock, increased leakage current.Remedy: Only connect USB sticks, external hard drives without external power supply or other approved components to the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 14.06.2022 10:16 | Rudolf Siebenhaar | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
650 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563464 | Headline | Chapter |  | Operate application on application host | Note, there are currently no risks related to "Operate application on application host" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:27 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
651 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563458 | Headline | Chapter |  | Protocol examination data | Note, there are currently no risks related to "Protocol examination data" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:27 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
652 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563453 | Headline | Chapter |  | Administer patient data | Note, there are currently no risks related to "Administer patient data" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:27 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
653 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563466 | Headline | Chapter |  | Close patient | Note, there are currently no risks related to "Close patient" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:28 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
654 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563465 | Headline | Chapter |  | Delete images | Note, there are currently no risks related to "Delete images" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:28 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
655 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563459 | Headline | Chapter |  | Shutdown system | Note, there are currently no risks related to "Shutdown system" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:28 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
656 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563469 | Headline | Chapter |  | Disconnect trolley and main unit | Note, there are currently no risks related to "Disconnect trolley and main unit" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:29 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
657 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563470 | Headline | Chapter |  | Disconnect trolley from line power | Note, there are currently no risks related to "Disconnect trolley from line power" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:29 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
658 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563468 | Headline | Chapter |  | Clean the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
659 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568820 | Headline | Function |  | Clean the mecanum wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
660 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568821 | Headline | Hazard |  | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
661 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568822 | Headline | Cause |  | Cleaning cover of mecanum wheels is not closed after cleaning |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
662 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568824 | Requirement | Measure | docUser_cleaning_mecanum | Caution - Close mecanum cover after cleaning | CAUTIONCause: Incorrect cleaning of wheels.Risk: Increased risk of crushing e.g., of fingers, risk of damage to the system.Remedy: Clean the wheels according to the instructions. Disconnect the C-arm chassis from the trolley. Do not move the system while a person is cleaning wheels. Before opening the wheel covers, activate the lock movement function at the chassis handle panel. Store wireless hand control in drawer and wired hand control in holder. Close the wheel covers after cleaning activities are finished. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.10.2022 09:57 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
663 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 571852 | Headline | Measure |  | Reference - Lock all motorized movements | <XREF>564347</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.10.2020 | 31.05.2022 17:30 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
664 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580832 | Requirement | Measure | label_hand_injury_mecanum | Label - Hand injury mecanum | Labels for hand crushing shall be placed near each mecanum wheel. They shall be visible when the covers for cleaning are open. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 16.07.2021 | 07.10.2022 09:23 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
665 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568823 | Headline | Cause |  | Unintended movement activation during cleaning activities of the mecanum wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_041 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
666 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578106 | Headline | Measure | docUser_cleaning_mecanum | Reference - Warning - Close mecanum cover after cleaning | <XREF>568824</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 01.02.2021 | 31.05.2022 17:32 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
667 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578107 | Headline | Measure |  | Reference - Lock all motorized movements | <XREF>564347</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 01.02.2021 | 31.05.2022 17:32 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
668 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568826 | Headline | Hazard |  | Contact with high level electrical energy (HS_002) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 31.05.2022 17:35 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | empty | empty | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
669 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568827 | Headline | Cause |  | Contact with electrical energy during mecanum wheel cleaning (e. g. ingress of cleaning agent) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_002 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | empty | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
670 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568828 | Requirement | Measure | ingress_fluids_mecanum | Protection against ingress of fluids - mecanum | The mecanum wheel electronics shall be protected against ingress of fluids with IPX4 according to IEC 60529. Hint: This applies to the electric drives. The mechanical rollers on the wheels themselves do not have to fulfill this standard. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 21.02.2022 11:58 | Johannes Schmidt |  |  |  | MainU,SysCabl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
671 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578318 | Headline | Measure | IEC60601_1_cl8-8_dielectric_strength | Reference - IEC 60601-1 Clause 8.7 and 8.8 | <XREF>562365</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | AccesS,DSMonS,HICSwitch,HICTouchCon,MainU,MCU,PowerU,TrolleyU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
672 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569267 | Headline | Chapter |  | Transport the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
673 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569269 | Headline | Function |  | Transport the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
674 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569268 | Headline | Hazard |  | Complete loss of movement functionality (e. g. chassis) (HS_110) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
675 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569270 | Headline | Cause |  | System gets damaged due to transport (e. g. vibrations, rough handling) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_110 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
676 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568837 | Headline | Measure |  | No additional measure required | The classification for mechanical shock and vibration during transport according to IEC 60721-3-2 is sufficient (2M4, except for free fall 2M5), see <XREF>561868</XREF>.No additional risk control measures are required. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 31.08.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,FDmediumH,FDmediumC,FLC,GridS,HICSwitch,HICTouchCon,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
677 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569271 | Headline | Hazard |  | Nonrecoverable or long-term loss of imaging functionality (HS_090) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions |  | S1M | VB10 | [12 empty] | NotDetermined
678 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569272 | Headline | Cause |  | System gets damaged due to transport (e. g. vibrations, rough handling) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_090 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
679 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569273 | Headline | Measure | class_shock_vibration_transport | Reference - Classification for mechanical shock and vibration - transport | The classification for mechanical shock and vibration during transport according to IEC 60721-3-2 is sufficient (2M4, except for free fall 2M5), see <XREF>561868</XREF>.No additional risk control measures are required. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 17.06.2022 10:09 | Johannes Schmidt |  |  |  | AppHost,Batt,ColS,FDmediumH,FDmediumC,FLC,GridS,HICSwitch,HICTouchCon,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
680 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563471 | Headline | Chapter |  | Install the system | Note, there are currently no risks related to "Install the system" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:29 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
681 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563472 | Headline | Chapter |  | Service the system | Note, there are currently no risks related to "Service the system" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:30 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
682 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563473 | Headline | Chapter |  | Refurbish the system | Note, there are currently no risks related to "Refurbish the system" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:30 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
683 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563474 | Headline | Chapter |  | Dispose the system | Note, there are currently no risks related to "Dispose the system" identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.08.2020 | 15.06.2022 12:30 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
684 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561939 | Headline | Chapter |  | Setup Trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
685 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561940 | Headline | Function |  | Move Trolley into OR |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
686 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561954 | Headline | Hazard |  | Pinching through breakage of supporting parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Instability,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
687 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561955 | Headline | Cause |  | Breakage of welded joints or wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 15:17 | Johannes Schmidt | [5 empty] | #Haz09#Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
688 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561956 | Requirement | Measure | design_of_the_welded_joints | Design of the welded joints | The welded joints or wheels must be designed in a way so that an overload can be carried in compliance with DIN EN 60601-1 (2006), chap. 9.8., table 21 | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 15:17 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
689 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561960 | Headline | Hazard |  | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
690 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561961 | Headline | Cause |  | Brake failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause036 | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
691 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561962 | Requirement | Measure | docUser_note_autonomous_movements_break_failure | User Manual - Brake failure | CAUTIONCause: Brake failurePossible consequences: Risk of crushingRemedy: Before beginning the examination, perform the daily function and safety checks. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:58 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
692 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561963 | Headline | Cause |  | Braking power not sufficient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause036 | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
693 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561964 | Headline | Measure |  | Reference - Design of the brakes | <XREF>562436</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.03.2022 11:21 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
694 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561965 | Requirement | Measure | docUser_parking_trolley_incline | User Manual - Parking Trolley on inclined Plane (i.e. Ramp) | CAUTIONCause: Insufficient braking force.Possible consequences: Risk of crushing!Remedy:- When parking the monitor trolley on an inclined surface, face the front toward the incline and lock the brakes.- Make sure the monitor trolley is not parked on a wet floor or other places where the adhesion of the wheels is greatly reduced. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:42 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
695 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561966 | Requirement | Measure | label_parking_position_trolley | Warning Label - Parking Position Trolley | Place warning label (3868681) next to foot brake. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:25 | Johannes Schmidt |  |  |  | TrolleyU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
696 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561967 | Headline | Hazard |  | Risk of injury through damaged components | Translation key:Risk of injury through damaged components | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
697 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561968 | Headline | Cause |  | Crossing a step |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause036 | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
698 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561969 | Headline | Measure |  | Reference - Rough handling test trolley | <XREF>561951</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
699 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561970 | Headline | Hazard |  | Crushing when adjusting heights of monitors |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
700 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561971 | Headline | Cause |  | Motor-driven movement |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:47 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause036 | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
701 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561972 | Requirement | Measure | distance_for_fault | Distance in case of fault according to IEC 60601-1 | Compliance with DIN EN 60601-1 (2006) Kap. 9.2; Table 20, Acceptable gaps for finger > 25mmNote: the monitor movements are not considered by the emergency stop | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 25.03.2022 10:43 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
702 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561973 | Requirement | Measure | docUser_note_risk_of_pinching | Attention note in user manual | CAUTIONCause: Motorized movements.Possible consequences: Risk of crushingRemedy: Do not stand in the travel range of the monitor or place objects on the trolley housing. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:43 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
703 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561974 | Headline | Hazard |  | Electric shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
704 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561975 | Headline | Cause |  | Protective earth broken |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause072 | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
705 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597040 | Headline | Measure | cable_protection | Reference Cable protection | <XREF>562104</XREF> and <XREF>562103</XREF> | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 29.06.2022 | 27.07.2022 09:24 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
706 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561980 | Headline | Hazard |  | Injury through dropping parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
707 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561981 | Headline | Cause |  | During transportation not secured parts may drop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause125 | #Haz55 #Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
708 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561982 | Requirement | Measure | storage_have_to_be_secured | Built-in parts have to be secured because of the risk for dropping | Built-in parts have to be secured because of the risk for dropping | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 25.03.2022 10:43 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
709 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561983 | Headline | Function |  | Startup Trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
710 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561984 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Untrained users,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
711 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561985 | Headline | Cause |  | Untrained users / Missing knowledge of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | See <XREF>273686</XREF>. | #Haz42 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
712 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561986 | Requirement | Measure | um_TrainedUsersOnly_s1 | User Manual - Trained Users | WARNINGCause: Operation of the system by non-trained usersPossible consequences: Misinterpretation of image information or of indications of the system, leading to inappropriate decisions in the clinical workflow, incorrect diagnosis or treatment, delay of the clinical procedure, or exposure to radiation without clinical benefit.Remedy: The system must only be used by persons with the necessary knowledge, after an appropriate application training. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
713 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561987 | Headline | Cause |  | Not intended use of system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause007 | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse,Patients with contraindications | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
714 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561988 | Headline | Measure |  | Reference - Operator Manual - Not intended use | <XREF>562190</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
715 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561989 | Headline | Hazard |  | Loss of functionality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Availability in critical functions,Foreseeable misuse,Untrained users,Inappropriate maintenance |  | S1M | VB10 | [12 empty] | NotDetermined
716 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561990 | Headline | Cause |  | Complexity of the system (component failures), mains supply failure or voltage out of specification |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz42 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
717 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561991 | Headline | Measure |  | Reference to User Manual - Establish emergency procedures | <XREF>562205</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
718 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561992 | Headline | Cause |  | (Cooling) Liquid is stored in the trolley and leaks, damaging the electronics |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz42 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
719 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561993 | Requirement | Measure | docServ_storage_cooling_liquid | No fluid storage in trolley | Service DocumentsWARNINGCause: Fluids are stored (e. g. cooling liquid) in the monitor trolley and leaks, destroying electronics.Possible Consequences: Damage of components in the trolley, possibly delay or interruption of the procedure.Remedy: Do not store fluids like the Service Pack with cooling liquid in the monitor trolley. Note:This is relevant to systems with cooling option. Nevertheless, storing of liquids inside of trolley is prohibited for all systems. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
720 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561994 | Headline | Cause |  | Foreseeable misuse |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause009 | #Haz20#Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
721 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561995 | Headline | Measure |  | Reference - OperatorManual-Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
722 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561996 | Headline | Cause |  | Some parts in system could lead to a hazard by wear |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz22 #Haz40 | See general justifications in the introduction which apply here. The electrical and mechanical components are designed according toapplicable regulations and technical standards. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
723 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561997 | Headline | Measure |  | Reference - OperatorManual-Abrasion of wear | <XREF>562201</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
724 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561998 | Headline | Hazard |  | Hazards from fire or smoke |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Flammability |  | S1M | VB10 | [12 empty] | NotDetermined
725 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 561999 | Headline | Cause |  | Overheating and Flammability Materials |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 15:50 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause019 | Haz36 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | D
726 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562000 | Headline | Measure | EN60601-1_s1 | Reference - IEC 60601-1 Standards | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 11.10.2022 17:42 | Bence David Fodor (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | empty | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
727 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562001 | Headline | Measure | IT_equipment_safety | Reference - IT equipment safety | Compliance with IEC 62368-1, 04S452203 and 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 11.10.2022 17:42 | Bence David Fodor (EXT) |  |  |  | AccesS,AppHost,FLC,PowerU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
728 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562002 | Headline | Measure |  | Reference - User Manual - Switch off the system in the event of fire User Manual | <XREF>562218</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
729 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562003 | Headline | Cause |  | Auto-ignition or Melting of Material |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
730 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562004 | Headline | Measure |  | No measure required | Result of risk evaluation: Temperature inside of trolley will never reach such threshold to enable auto-ignition or any melting of material. Therefore, ho risk control measures are necessary.Note:System specific documents (e.g. protocols etc., i.e. blue "Service-Folder") may be stored inside of trolley. This wish come up as system specific documents needs to be on-site during product life time. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
731 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562005 | Headline | Cause |  | Explosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz37 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | A | High | A
732 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562006 | Requirement | Measure | um_NoAPClassification | User Manual - Danger of explosion | WARNINGCause: The system is not designed for operation in explosion-endangered areas. It does not fulfill the requirements for AP/APG classification.Possible consequences: Danger of explosionRemedy: The system may not be used in locations where explosive atmospheres may occur. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
733 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562007 | Headline | Cause |  | Too high temperature of the detector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 #Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
734 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562008 | Headline | Measure |  | Reference - System Warning - Temperature of the Detector | See <XREF>562226</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | ReadyForTCReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
735 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562009 | Headline | Hazard |  | Chemical hazards |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Toxicity of chemical constituents |  | S1M | VB10 | [12 empty] | NotDetermined
736 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562010 | Headline | Cause |  | Toxic materials and fluids |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz38#Haz59 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
737 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562011 | Headline | Measure |  | Reference - Technical instruction for the disposal procedure | <XREF>562246</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
738 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562012 | Requirement | Measure | docServ_harmful_cooling_liquid | Service Instruction - Toxic Fluid | CAUTIONCause: Cooling liquid (such as Antifrogen N) is not handled correctly.Possible consequences: Skin or eye irritation, irritation of the breathing passages, toxic or allergic reactions caused by contact with hazardous substances.Remedy: Avoid skin contact, inhalation and swallowing, use personal protective equipment such as gloves and goggles. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
739 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562013 | Headline | Hazard |  | electromagnetic disturbances |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC),loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
740 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562014 | Headline | Cause |  | Delayed or disturbed transmission of images, or other functional disturbance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51 #Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
741 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562016 | Requirement | Measure | docUser_degradation_of_essential_performance | User Manual - Degradation of essential performance by electromagnetic interferences | CAUTIONCause: Electromagnetic interference of the system.Possible consequences: Degradation of the image quality or accuracy of the radiation parameters!Remedy: - Remove the source of interference (for instance, a mobile phone) from the area.- Shut down the system and turn it on again.- Pay attention to error messages during restart. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:45 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
742 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562017 | Headline | Measure |  | Reference to generator measure for IEC 60601-1-2 | See HM_S1_mec_F1G1U1M1 and HM_S1_emv_F1G1U2M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | AppHost,ColS,DapS,DSMonS,FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
743 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562018 | Headline | Hazard |  | Hearing impairment of patients and/or staff |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Acoustic energy |  | S1M | VB10 | [12 empty] | NotDetermined
744 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562019 | Headline | Cause |  | Excessive audible noise |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
745 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562020 | Headline | Measure |  | No measure required | The acoustic energy / audible noise is limited via IEC60601-1:2005, chapter 9.6.2.1. Therefore, no additional risk control measures are necessary. has to be smaller/equal 80 dB(A). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | AccesS,FLC,HICTouchCon,MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
746 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562021 | Headline | Hazard |  | Injuries due to monitor dropping |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
747 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562022 | Headline | Cause |  | Breakage of the monitor support |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
748 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562023 | Requirement | Measure | compliance_DIN_EN_60601_1 | Compliance with DIN EN 60601-1 | The monitor support must be designed in a way that an overload can be carried. Remark:Compliance with DIN EN 60601-1 (2006), Kap. 9.8 Table 21 | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.03.2022 13:01 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
749 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562024 | Headline | Hazard |  | Risk of injury through impaired parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
750 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562025 | Headline | Cause |  | Transport damages e.g. through collision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
751 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562026 | Requirement | Measure | safe_transport_position_for_monitors | Protection of monitor during transport | The monitor shall be robust against mechanical impacts from all sides:A cover shall protect the monitor on all sides, except the front side from impact of collisions in order to meet at least IEC 60601-1 cl. 15.3.3. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 12.07.2022 11:04 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
752 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596969 | Requirement | Measure | robust_touch_glass_of_monitor | Robust touch glass of monitor | The touch glass of the monitor shall be robust against mechanical impacts (e. g. at least meeting IEC 60601-1 cl. 15.3.3). If the touch glass is broken due to an impact, it shall not splinter or expell shards. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 20.06.2022 | 21.09.2022 10:54 | Johannes Schmidt |  |  |  | DSMonS |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
753 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562027 | Headline | Hazard |  | Electric shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
754 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562028 | Headline | Cause |  | High bleeder resistor |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
755 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562029 | Requirement | Measure | docSCM_leakage_current_measurement | Leakage current measurement | Leakage current measurement during final inspection | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
756 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562030 | Requirement | Measure | docCS_note_leakage_current_measurement_after_maintenance | Work step in maintenance documents | Work step in maintenance documents: operate leakage current measurement after maintenance work | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
757 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562031 | Requirement | Measure | docCS_leakage_current_measurement_after_service | Attention note in service manual | Attention note in the relevant service manual: Operate leakage current measurement after service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
758 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562032 | Headline | Cause |  | Interruption of the protective earth |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
759 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562033 | Requirement | Measure | docSCM_protective_earth_measurement | Protective earth measurement | Protective earth measurement during final inspection | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
760 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562034 | Requirement | Measure | docCS_protective_earth_measurement_after_service | Attention note in service manual | Attention note in service manual: Perform protective earth measurement after service work. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
761 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562035 | Requirement | Measure | docCS_protective_earth_measurement_after_maintenance | Work step in service instructions | Work step in maintenance documents: Perform protective earth measurement after maintenance work. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
762 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562036 | Headline | Cause |  | Mechanically impaired cables/wiring |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
763 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562037 | Requirement | Measure | cable_conduit_for_movable_cables | Cable conduit for movable cables | Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition.This measure is in accordance to IEC60601-1:2005 chap. 8.10. Note:Movable components are affected by e.g. adjustment of monitor height or horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 12:54 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
764 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562038 | Headline | Cause |  | Static charge |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz11 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | D
765 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562039 | Requirement | Measure | create_conductance | Create conductance on the trolley and main unit | In order to prevent electrostatic charging (e. g. when the system is moved over the floor), the system shall provide a conductive connection to the ground, e.g. conductance band, conductive wheels. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 01.04.2022 10:20 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
766 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562040 | Requirement | Measure | docCS_note_scanning_for_damage_and_fouling | Work step in maintenance documents | Work step in maintenance documents: Scanning of the conducting elements for damage and fouling | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
767 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562041 | Headline | Cause |  | Residual charge in components after power off or impaired electrical integrity of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | HS_003 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | B
768 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562042 | Headline | Measure |  | Reference to compliance to IEC 60601-1 |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
769 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562043 | Headline | Measure |  | Reference to Service Documentation - Trained service personnel |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
770 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562044 | Requirement | Measure | serv_discharge_capacitor | Caution in Service Documentation - Discharge capacitor | Service DocumentationCAUTIONCause: Replacing the fuses F1 or F2 on D819 without discharging the capacitor C107Possible Consequences: Electrical shock due to capacitor dischargeRemedy: Discharge capacitor C107 on D819 by connecting a multimeter for 10 minutes to X34 and X35. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Line voltage | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
771 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562045 | Headline | Hazard |  | Risk of injury due to bad surface finish |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Sharp edges and points |  | S1M | VB10 | [12 empty] | NotDetermined
772 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562046 | Headline | Cause |  | Sharp edges |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
773 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562047 | Requirement | Measure | design_without_sharp_edges | Develop design without sharp edges | Develop design without sharp edges.(referring to IEC60601-1:2005 chapter 9.3.) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 29.06.2022 12:57 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
774 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562048 | Headline | Measure |  | Reference to measures | HM_S1_mec_F4G2U1M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,IOC,MainU,PowerU,TrolleyU,XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
775 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562049 | Headline | Cause |  | Rough surface finish |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
776 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562050 | Requirement | Measure | smooth_surfaces | Specify smooth surface finishes | Specify smooth surface finishes(referring to IEC60601-1:2005 chapter 9.3) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 09:38 | Johannes Schmidt |  |  |  | HICTouchCon,MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
777 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562052 | Headline | Hazard |  | Effect on people, e.g. tissue injury, reddened skin |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC) |  | S1M | VB10 | [12 empty] | NotDetermined
778 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562053 | Headline | Cause |  | Electromagnetic radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC) | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
779 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562054 | Headline | Measure |  | No measure required | The permissible electromagnetic emissions of the system and its components (e. g. WLAN) fulfill the limits specified in IEC 60601-1-2. Therefore, no further risk control measures are necessary. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 23.09.2022 14:12 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
780 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562055 | Headline | Hazard |  | Effect on people with implants e.g. pacemaker |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC) |  | S1M | VB10 | [12 empty] | NotDetermined
781 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562056 | Headline | Cause |  | Interference into the implant\s electronics |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz44 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
782 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562058 | Requirement | Measure | docUser_note_minimum_distance | User Manual - Interference into Implants | CAUTIONCause: InterferencePossible consequences: Effect on people with implantsRemedy: Maintain a minimum distance in accordance with IEC 60601-1-2. If this is not possible, switch off the WLAN client. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:48 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
783 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 576774 | Headline | Measure | EMC_s1 | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost,Batt,ColS,DapS,DSMonS,FDmediumH,FDmediumC,FLC,HICTouchCon,IOC,MainU,MCU,PowerU,TrolleyU,XCU,XGen,SysCabl,HICForceControl,HICHandControl | IEC 60601-1; 17 Electromagnetic compatibility of ME EQUIPMENT and ME SYSTEMSIEC 60601-1-2 + HF surgical equipment |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
784 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562063 | Headline | Chapter |  | Setup Main Unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
785 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562064 | Headline | Function |  | Move Main Unit in OR |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
786 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562065 | Headline | Hazard |  | Crushing due to system tipping over |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
787 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562066 | Headline | Cause |  | C-arm movement shifts over-all center of gravity |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
788 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562067 | Requirement | Measure | design_of_system_angle_of_5_degree | Design of system | The system must be designed in a way that prevents tipping over on inclined plane (i.e. ramp) at an angle of 5°. Remark:Compliance with DIN EN 60601-1 tilt test | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 15:18 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
789 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562068 | Headline | Cause |  | Weight put on in addition changes over-all center of gravity |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
790 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562069 | Requirement | Measure | docUser_note_weight_put_on | User Manual - Weight Balancing | WARNING:Cause: Change of overall center of gravityPossible consequences: Risk of crushingRemedy: Only products/components approved by Siemens Healthineers may be installed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:15 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
791 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562070 | Headline | Cause |  | Transport of the device over an inclined plane (i.e. ramp) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
792 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562071 | Requirement | Measure | design_of_system_ramps | Design of system | The system are designed in a way that prevents tipping over on inclined plane (i.e. ramp) at an angle of 10°. Remark:Compliance with DIN EN 60601-1 tilt test for transportation | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:16 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
793 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562072 | Headline | Cause |  | Crossing a step |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
794 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562073 | Headline | Measure |  | Reference - Rough handling test (main unit) | <XREF>562087</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:32 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
795 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562074 | Headline | Hazard |  | Crushing through breakage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Instability,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
796 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562075 | Headline | Cause |  | Breakage of welded joints |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
797 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562076 | Headline | Measure |  | Reference - Design of the welded joints | <XREF>561956</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
798 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562077 | Headline | Hazard |  | Crushing through breakage of the wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
799 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562078 | Headline | Cause |  | Carrying capacity of the wheels too small |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
800 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562079 | Headline | Measure |  | Reference - Design of the wheels | <XREF>561959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
801 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562080 | Headline | Hazard |  | Crushing due to autonomous movements on inclined plane (i.e. ramp) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
802 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562081 | Headline | Cause |  | Brake failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
803 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562082 | Headline | Measure | ref_self_test_brakes | Reference - self test of the brakes | <XREF>564087</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 27.07.2022 14:12 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
804 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562083 | Headline | Cause |  | Braking power not sufficient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 #Haz41 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
805 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562084 | Headline | Measure |  | Reference - Specify sufficient braking power | <XREF>562436</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:42 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
806 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562085 | Headline | Hazard |  | Risk of injury through damaged components |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
807 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562086 | Headline | Cause |  | Moving the system over a doorstep, against a doorframe, wall or other equipment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:24 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 #Haz42#Haz12 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
808 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562087 | Requirement | Measure | rough_handling_test | Rough handling test (main unit) | The chassis shall be designed to comply with IEC 60601-1 cl. 15.3.5. Scenarios to be tested: - Transport position of the system reached and fast drive move activated for a) Ascending step shock, b) Descending step shock, c) Door frame shock and forward direction (Note: During a) and c) the collision avoidance is activated and slows down the system to operating room speed. For b) fast drive mode speed is applied.) - Transport position of the system and operating room speed for a), b), c) in forward and lateral direction (Note: During a) and c), this is expected to repeat the first test with a) and c) with collision avoidance activated. b) is expected to be less critical than the first test with b) with transport speed). | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:17 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
809 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 594861 | Headline | Measure | um_steep_ramp | Reference - steep ramp | <XREF>578343</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 21.02.2022 | 08.03.2022 17:33 | Johannes Schmidt | [7 empty] | Basic | System Test | none | n.a. | Not Determined | empty | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
810 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562090 | Headline | Hazard |  | Injury through damaged control element |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
811 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562091 | Headline | Cause |  | Control element falls on the floor during transportation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz18 | See general justifications in the introduction which apply here. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
812 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562092 | Requirement | Measure | drop_test | Drop test | The portable control elements except Hand-Helds with masses <= 10 kg have to be designed so that they can be dropped from a height of 5 cm without having damages that cause unacceptable risks.Hand-Helds have to be designed to withstand drop from heights at which the ME EQUIPMENT is used (as specified in the ACCOMPANYING DOCUMENTS), or from a height of 1 m, whichever is greater, onto a 50 mm +/-5 mm thick hardwood board (hardwood > 600 kg/m3) lying flat on a concrete or a similar rigid base.Test criterion:Execute drop tests according to standard IEC 60601-1 at heights of 5 cm for portable mobile control panels or heights of >= 1m for Hand-Helds. Possible emerged damages have to be examined for hazards (e.g. access to electrical and moveable parts has to be prevented). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 03.02.2022 14:37 | Ingo Hollenborg |  |  |  | HICSwitch,HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
813 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562093 | Headline | Hazard |  | Biological and Chemical Hazards | The devices must be designed and manufactured in such a way as to reduce to a minimum the risks posed by substances leaking from the device. Special attention shall be given to substances which are carcinogenic, mutagenic or toxic to reproduction (MDD / MDR) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature,Flammability |  | S1M | VB10 | [12 empty] | NotDetermined
814 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562094 | Headline | Cause |  | Detector is damaged, service personnel comes into contact with CsI |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz59 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Toxicity of chemical constituents,Leaking subst. that are carcinogenic, mutagenic, toxic | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
815 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562096 | Requirement | Measure | manu_detector_toxic_material_s1 | Supply Chain Documentation - Toxic Materials | Make note in the relevant factory documents. How these parts are disposed in an environmentally friendly manner have to be documented in the disposal instructions.Cause: Toxic materialsPossible consequences: Toxic contamination in case of detector panel is broken by mechanical shockRemedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens Healthineers. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:16 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Leaking subst. that are carcinogenic, mutagenic, toxic |  | S1M | VB10 | [12 empty] | NotDetermined
816 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562097 | Requirement | Measure | docServ_detector_toxic_material_s1 | Service Instruction - Toxic Materials | Make note in the relevant service documents. How these parts are disposed in an environmentally friendly manner have to be documented in the disposal instructions. CAUTION Cause: Toxic material is accessible when the image intensifier / flat detector is damaged due to mechanical impact.Possible consequences: Skin or eye irritation, irritation of the breathing passages, toxic or allergic reactions caused by contact with hazardous substances.Remedy: Handle the image intensifier / flat detector with care. If it is damaged, do not use it. Inspect if any particles are spilled. Avoid skin contact and inhalation of dust, use personal protective equipment if required. Spilled particles have to be collected and kept in a sealed container and returned to Siemens Healthineers for disposal. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:16 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Leaking subst. that are carcinogenic, mutagenic, toxic |  | S1M | VB10 | [12 empty] | NotDetermined
817 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562098 | Headline | Cause |  | Detector not disposed correctly |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz38 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Toxicity of chemical constituents,Leaking subst. that are carcinogenic, mutagenic, toxic | Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
818 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562099 | Headline | Measure | docServ_detector_toxic_material_s1 | Service Instruction - Toxic Materials - Disposal | See <XREF>407319</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Leaking subst. that are carcinogenic, mutagenic, toxic |  | S1M | VB10 | [12 empty] | NotDetermined
819 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569205 | Headline | Cause |  | Detector is damaged, operator comes into contact with CsI |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 09.09.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz59 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Toxicity of chemical constituents,Leaking subst. that are carcinogenic, mutagenic, toxic | User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
820 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562095 | Requirement | Measure | um_detector_toxic_material_s1 | User Manual - Detector Cover Broken | CAUTIONCause: Flat detector cover brokenPossible consequences: Skin or eye irritation, irritation of the breathing passages, toxic or allergic reactions caused by contact with hazardous substances.Remedy: If the flat detector cover is broken, do not touch the detector, do not use the system at all. Please call Siemens Healthineers service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:49 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Leaking subst. that are carcinogenic, mutagenic, toxic |  | S1M | VB10 | [12 empty] | NotDetermined
821 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569206 | Headline | Measure |  | No additional measure required | 2 covers around hazardous substance provided. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 09.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Leaking subst. that are carcinogenic, mutagenic, toxic |  | S1M | VB10 | [12 empty] | NotDetermined
822 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562100 | Headline | Function |  | Connect Trolley and Main Unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
823 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562101 | Headline | Hazard |  | Electric shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
824 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562102 | Headline | Cause |  | Damage of the cable |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
825 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562103 | Requirement | Measure | certified_power_cable | Certified mains power cable | The mains power supply cable shall be in accordance with IEC 60601-1 cl. 8.11.3.Country specific certificates may be used when equivalent. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 01.04.2022 10:22 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
826 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569245 | Requirement | Measure | robust_tc_cable | Robust trolley-chassis cable | Robust trolley-chassis cable The system shall prevent damage to trolley-chassis cable from extensive tension caused by movements of the Main Unit.Note, a collision detection system is implemented which detects tension at the TC cable. The trolley chassis cable shall tolerate forces and bending up to the limit when tension detection is activated.<XREF>563740</XREF> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 14.09.2020 | 04.03.2022 14:25 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
827 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562104 | Requirement | Measure | cord_anchorage | Cable anchorage | All power cables have to be protected against cable damage to prevent electrical shock. In addition, it has to ensured, that system's connection to power line remains active as long as possible (a connection via socket outlet and plug is not sufficient). All power cables need to be secured with cable anchorage / strain-relief, e.g. by means of a strain-relief sleeve.Pull forces for cable retention and cable anchorage shall be in accordance to IEC62444-2010 (Table 2). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 09:36 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
828 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562105 | Requirement | Measure | docUser_note_check_power_cable | User Manual - Damage of Cable | WARNING:Cause: Cable damagePossible consequences: Risk of electrical shockRemedy: Check the power cable. Do not use the device if the power cable is damaged. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
829 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562106 | Headline | Cause |  | Interruption of the protective earth |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
830 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562107 | Headline | Measure |  | Reference - Protective earth measurement | <XREF>562033</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
831 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562108 | Headline | Measure |  | Reference - Attention note in service instructions | <XREF>562034</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
832 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562109 | Headline | Cause |  | Excess current or short-circuit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
833 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562110 | Requirement | Measure | fuse_for_the_supply_line | Fuse for the supply line | Fuses shall be provided in each supply lead | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:19 | Rudolf Siebenhaar |  |  |  | MainU,PowerU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
834 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562111 | Headline | Measure |  | Reference - Warning label | <XREF>563226</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
835 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562112 | Headline | Cause |  | Static charge |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz11 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | D
836 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562113 | Headline | Measure |  | Reference - Create conductance on the trolley and main unit | <XREF>562039</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
837 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562114 | Headline | Measure |  | Reference - Work step in maintenance documents | <XREF>562040</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
838 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562115 | Headline | Hazard |  | Risk to fall |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
839 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562116 | Headline | Cause |  | Tripping over cables |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz19 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
840 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562117 | Requirement | Measure | docUser_note_Installing_cables | User Manual - Installing Cables | CAUTIONCause: Tripping over cablePossible consequences: Risk of fallingRemedy: When routing cables, make sure they run on the floor properly and without loops. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 10:50 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Moving parts | User | S1M | VB10 | [12 empty] | NotDetermined
841 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562118 | Headline | Hazard |  | Excessive temperature, fire, explosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature,Flammability,loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
842 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562119 | Headline | Cause |  | Overcharge of rechargeable batteries |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz35 #Haz15 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
843 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562120 | Requirement | Measure | provide_overcharge_protection | Provide protection from overcharge HICTouchCon | Provide protection from overcharge | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VxFUTURE | [12 empty] | NotDetermined
844 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562121 | Requirement | Measure | provide_overcharge_protection_2 | Provide protection from overcharge HICSwitch | Provide protection from overcharge | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
845 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562122 | Requirement | Measure | use_approved_batteries | Use approved batteries | Use approved batteries | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
846 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562123 | Headline | Cause |  | Battery Gas Emission during Charge |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 #Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
847 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562124 | Requirement | Measure | aerated_battery_cases | Batteries with Temperature Sensors | Use batteries with temperature sensors to enable interrupt of charging device. Interruption of charging will prevent (toxic) gas emission. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 16:40 | Johannes Schmidt |  |  |  | Batt |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
848 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562125 | Headline | Cause |  | Overcharge of the capacitors |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:27 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz35 #Haz15 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
849 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562126 | Requirement | Measure | mounted_under_cover | Mounted under covers | Capacitors are mounted under a dedicated cover. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:38 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
850 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562127 | Headline | Measure |  | Reference - Protection of the energy storage | <XREF>562523</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
851 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562128 | Headline | Measure |  | Reference - Rupture joint by Overcharge and pressure relief valve | <XREF>562525</XREF><XREF>562524</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
852 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562129 | Headline | Hazard |  | Foreseeable missuse |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
853 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562130 | Headline | Cause |  | Misplugging of cables |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. Usability is considered within the development process. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
854 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562131 | Requirement | Measure | misplugging_cables | Design considering misplugging | The design of the external accessible connectors must consider misplugging of cables. Compare also chapter 16.9 of IEC60601-1:2005. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:20 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts | User | S1M | VB10 | [12 empty] | NotDetermined
855 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562132 | Headline | Hazard |  | Unintentional Radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
856 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562133 | Headline | Cause |  | Exposure release switch is pressed by mistake during start up |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
857 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562134 | Requirement | Measure | Check_X-RaySwitchAtStartUp_s1 | Check radiation release signal status at startup | The system shall check the status of both radiation release signal paths at startup of every control.When an active radiation release signal path is detected, radiation from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths take place. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 08:54 | Johannes Schmidt |  |  |  | MCU | SFS |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
858 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562135 | Requirement | Measure | emergency_stop_radiation | Emergency stop radiation | If an emergency stop <XREF>562345</XREF> is activated, radiation shall be stopped, but not inhibited.When the operator activates radiation again, the system shall release radiation again. Note: The SysCB monitors the status of the emergency stop and provides this information to the XCU. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 27.07.2022 14:18 | Johannes Schmidt |  |  |  | MCU,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
859 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562136 | Requirement | Measure | um_EmergencyStopInterruptsXray | User Manual - Emergency Stop Button | CAUTIONCause: Because of a technical error the acquisition is not stopped or there is an unintended release of radiation.Possible consequences: Risk of unnecessary radiation exposureRemedy: In the event of unwanted radiation, press the nearest EMERGENCY STOP button. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 12:52 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
860 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562137 | Requirement | Measure | docUser_ShutDownSwitchInterruptsXray | User Manual - Shutdown switch | CAUTIONCause: Because of a technical error the acquisition is not stopped or there is an unintended release of radiation.Possible consequences: Risk of unnecessary radiation exposureRemedy: Press the OFF button on the monitor trolley, even if the EMERGENCY STOP button has already been pressed. Radiation terminates immediately. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 12:54 | Johannes Schmidt |  |  |  |  | IEC 60601-1: cl. 7.2.13 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
861 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562138 | Headline | Measure |  | Reference to radiation indication signal | See HM_S1_hw_F1G1U1M3 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
862 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562139 | Headline | Cause |  | Radiation indicator fails |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. In this specific case X-Ray imaging is not prevented in case of a failing x-ray indicator. The probability of two or more failing x-ray indicators failing at the same time is rated as very low. The benefit for system availability and to continue x-ray imaging is rated as high. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
863 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568841 | Requirement | Measure | monitor_radiation_indicator | Monitoring of radiation indicators | The radiation indicator circuit shall be monitored (e. g. current in indicator circuit when radiation is emitted). When an error is detected, the system shall show an appropriate error message (no blocking of radiation). Note, radiation indicators are provided on top of the trolley, on top of the CCP and on top of the RCP, if connected to the system. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 22.03.2022 16:28 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Environment | S1M | VB10 | [12 empty] | NotDetermined
864 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562140 | Requirement | Measure | plug_stablized_connection | Plug with stablized connection | Utilize plug with stabilized mechanical connection against disconnection. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 25.03.2022 10:45 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Environment | S1M | VB10 | [12 empty] | NotDetermined
865 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562141 | Requirement | Measure | docUser_note_radiation_indicator_failed | User Manual - Radiation Indicator | CAUTIONCause: Failure of radiation indicator.Possible consequences: Danger of unintended radiation!Remedy: If a radiation indicator fails contact Siemens Healthineers Customer Service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:18 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Environment | S1M | VB10 | [12 empty] | NotDetermined
866 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562142 | Headline | Cause |  | Radiation indicator LED fails |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
867 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562143 | Headline | Measure |  | No measure required | Note:Lifetime of LED (approx. 100,000 hours) will exceed lifetime of C-arm system. In addition, radiation indicator is equipped with multiple LEDs, so indication of X-ray is possible while failure of a single LED. Therefore, no measure is required. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
868 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562144 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Lack of integrity of data,Loss of data,unintended interactions within and among SW sys,unknown aspects or quality of third-party software,lack of data sec, particularly vulnera. to tampering,Malicious software,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
869 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562145 | Headline | Cause |  | Prevent startup |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51 #Haz31 | See general justifications in the introduction which apply here. The electrical and mechanical components are designed according toapplicable regulations and technical standards. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
870 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562146 | Headline | Measure |  | Reference - User Manual - Wear and tear | <XREF>562201</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
871 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562147 | Headline | Cause |  | Data loss because the System is not safely turned off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz40 #Haz30 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
872 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562148 | Requirement | Measure | docUser_shutdown_xray | User Manual - ShutDown during X-Ray | CAUTIONCause: Loss of image dataPossible consequences: Risk of unnecessary radiation exposure!Remedy: Do not shut down the system while radiation is being released. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
873 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562149 | Headline | Cause |  | System not in a proper condition e.g. error in a component |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51 #Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
874 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562150 | Requirement | Measure | DisplaySystemStatus | Display system status | System status is visible to indicate the presence of an error. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
875 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562151 | Headline | Cause |  | User loads non-approved SW or add non approved HW to the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 #Haz42 #Haz52 #Haz01 | See general justifications in the introduction which apply here. The operator manual warns to not perform non-approved modifications of the system, incl. combining with non-approved accessories or other devices. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
876 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562152 | Requirement | Measure | um_NonapprovedSwHw | User Manual - Use of non-approved parts | WARNING:Cause: Use of unreleased or altered software or hardware components.Possible consequences: Risk of malfunction that could endanger the patient or product.Remedy:- Only use Siemens Healthineers-authorized software and hardware components.- Only perform repairs, or have them performed by an authorized third party, after having received our express written approval. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:19 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
877 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562153 | Headline | Cause |  | Impermissible or faulty manipulation of SW from network. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz23#Haz41 #Haz51#Haz31#Haz24 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Malicious software | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
878 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562154 | Requirement | Measure | um_ImpermissibleSW | User Manual - Impermissible or faulty manipulation from network | CAUTIONCause: Unauthorized or incorrect alterations/changes of the software or connection of the system to a network.Possible consequences: Risk of unauthorized access.Remedy: Make sure that all necessary precautions are taken with regard to the applicable security level if a functionality is added or the configuration of the condition on delivery is changed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
879 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562155 | Headline | Measure |  | Reference to measures against unauthorized installation of non-approved software | <XREF>563266</XREF><XREF>563267</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
880 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569133 | Requirement | Measure | FLC_VirusProtection_OS_Access | Imaging system access restrictions | The system user does not have access to the system level.This includes the access to the underlaying operating system and the belonging functions like file menu, system properties or start menu. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
881 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569135 | Requirement | Measure | FLC_VirusProtection_Autorun | Disable autorun for removable media | The autorun feature for removable media must be disabled on the FLC imaging system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
882 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569136 | Requirement | Measure | FLC_Internet_and_BIOS_Access | Imaging system access restrictions for internet and BIOS | Define a testcase to verify that the access to the internet is not possible for the clinical user. The service user shall be able to access internet via a special service environment.The access to BIOS is password-protected.. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 04.09.2020 | 29.09.2022 15:33 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
883 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562156 | Headline | Cause |  | Incorrect treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface,Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
884 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562157 | Headline | Measure |  | Reference - User Manual - Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
885 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562158 | Headline | Cause |  | Harddisk or file system defect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21#Haz50 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss of data | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
886 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562159 | Requirement | Measure | docUser_no_archive_medium | User manual - No Archive Medium | CAUTIONCause: Loss of image dataPossible Consequences: Risk of unnecessary X-ray exposureRemedy:- The hard disk of the imaging system is not suitable for long-term archiving of patient and image data. There are statutory requirements governing the archiving period, data availability, and data security (data integrity), as well as recommendations concerning fire protection or water damage for the archiving of image data. The operator of the archive is responsible for observing these regulations.- Because of advances in technical development and statutory requirements, the storage of image data and access to it cannot be realized with a single storage and media technology alone. Therefore, the migration of data may be necessary under the responsibility of the operator of the digital archive. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
887 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562160 | Headline | Cause |  | Electric power failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz40 #Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss of data | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
888 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562161 | Headline | Measure |  | Reference - The imaging system performs a controlled shutdown | <XREF>563249</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
889 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562162 | Headline | Cause |  | Insufficient charge (power) of UPS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
890 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562163 | Headline | Measure |  | No measure required | Utilization of UPS has changed: intended use is not keeping system operational but ensuring a controlled shut down.Hazard as well as hazardous situation/Cause is not applicable. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
891 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562164 | Headline | Cause |  | Charging capacity of the Uninterruptible Power Supply worn out |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
892 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562165 | Headline | Measure |  | No measure required | Utilization of UPS has changed: intended use is not keeping system operational but ensuring a controlled shut down.Hazard as well as hazardous situation/Cause is not applicable. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
893 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562166 | Headline | Function |  | Performing system relevant check |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
894 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562167 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
895 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562168 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
896 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562169 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
897 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562170 | Headline | Function |  | Performing quality control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
898 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562171 | Headline | Hazard |  | No hazard on system level | (image quality) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
899 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562172 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
900 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562173 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
901 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562174 | Headline | Function |  | Calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
902 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562175 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
903 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562176 | Headline | Cause |  | Incorrect parameters |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
904 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562177 | Requirement | Measure | FD_Calibration_s1 | Automatic procedure to perform FD calibration correctly | Automatic FD calibration procedure has to be initiated manually. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
905 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562178 | Headline | Cause |  | Operating error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
906 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562179 | Requirement | Measure | um_Calibration | Admin Guide - Calibration Procedure | CAUTIONCause: Operator error.Possible Consequences: Risk of unnecessary radiation exposure! Remedy: Follow the calibration procedure instructions. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 11.10.2022 13:35 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Admin Guide | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
907 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562180 | Headline | Hazard |  | Image artifacts - Wrong interventional decision possible |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
908 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562181 | Headline | Cause |  | Bad calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
909 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562182 | Headline | Measure |  | Reference - Automatic procedure to perform calibration correctly | <XREF>562177</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
910 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562183 | Headline | Measure |  | Reference - User Manual - Calibration Procedure | <XREF>562179</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
911 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562184 | Headline | Hazard |  | Unwanted radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
912 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562185 | Headline | Cause |  | Person in room during calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:33 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
913 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562186 | Requirement | Measure | serv_calibrating_s1 | Service training | Product specific service training (for service calibration) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
914 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562187 | Headline | Function |  | System Operation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
915 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562188 | Headline | Hazard |  | Danger for patient or user |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Assignment of image, data, objects,Workflow and configuration,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
916 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562189 | Headline | Cause |  | Wrong Intended Use |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Assignment of image, data, objects,Workflow and configuration,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
917 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562190 | Requirement | Measure | docUser_intended_use | User Manual - Intended use only | CAUTIONCause: Improper use of the system beyond its intended use.Possible consequences: Risk to the patient and user!Remedy: The system may only be used as defined by its intended use for clinical applications. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
918 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562191 | Headline | Cause |  | Crash of subsystems |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. In case of a recoverable malfunction, the operator manual advises the to restart the system. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
919 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562192 | Headline | Measure |  | Architectural Design of Subsystems and Components | see following sections in System Architecture (<XREF>561860</XREF>): * <xref>286376</xref> , see spacing * <xref>286377</xref> * <xref>353392</xref> * <xref>353414</xref> | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 09:27 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
920 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562193 | Requirement | Measure | docUser_reset_procedure | User Manual - Reset procedure for a freezed imaging SW | CAUTIONCause: Imaging system is frozen or restart is not possible.Possible consequences: Risk of delay in the clinical procedureRemedy: Restart the system as presented in the following reset procedure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
921 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562194 | Requirement | Measure | FLC_Radiation_Inputs_blocked | Block keyboard / mouse inputs during radiation | Functionality that would cause dialogs or other UI elements to cover the live image or that would cause performance bottlenecks for live image processing and display are disabled during radiation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
922 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562195 | Headline | Hazard |  | Loss of functionality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Assignment of image, data, objects,Workflow and configuration,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
923 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562196 | Headline | Cause |  | Complexity of the system or line voltage below the specified threshold |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz42 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | loss of basic safety or essential performance,Deterioration in function,use by unskilled/untrained personnel,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
924 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562197 | Headline | Measure |  | Reference to User Manual - Establish emergency procedures | <XREF>562205</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
925 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562198 | Headline | Cause |  | Foreseeable misuse |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz51#Haz41#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
926 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562199 | Headline | Measure |  | Reference - User Manual - Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
927 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562200 | Headline | Cause |  | Some parts in system could lead to a hazard by wear |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 #Haz41 | See general justifications in the introduction which apply here. The electrical and mechanical components are designed according toapplicable regulations and technical standards. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
928 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562201 | Requirement | Measure | um_RegularMaintenance | User Manual - Wear and tear | CAUTIONCause: Wear and tear.Possible consequences: Risk of injury to the patient, operating personnel, and other persons. Permanent damage to the system.Remedy: Follow the maintenance guidelines to maintain the safety and functionality of the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
929 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562202 | Headline | Cause |  | Operational hazard due to missing knowledge of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz42 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | use by unskilled/untrained personnel,slips, lapses and mistakes | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
930 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562203 | Requirement | Measure | serv_Trained_Technician_only_s1 | Provide training for technician | Provide training for technician.Service Documentation:WARNING:Cause: Missing training or not following service instructions can result in failure.Risk: Incorrect service can lead to damage for patient, personnel and system.Remedy: Service activities are only permitted by trained personnel and according to the appropriate service instructions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
931 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562204 | Headline | Cause |  | Component failure lead to system failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Due to the fact that no occurrence of harm is known for the lifetime of the whole platform (as of 2014), and that the calculated P1 is at the lower boarder of the range D (same order of magnitude), P harm is reduced to A. | #Haz42 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Assignment of image, data, objects,Workflow and configuration,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
932 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562206 | Headline | Hazard |  | Danger to electric high voltage or line voltage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage,Line voltage,Leakage current,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
933 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562207 | Headline | Cause |  | Problems with protective ground |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
934 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562208 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:30 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
935 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562209 | Requirement | Measure | ProtectiveEarth_s1 | Start up instruction | Start up instructionProtective earth measurements are described in start up instructions and done during system installation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 29.09.2022 15:34 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
936 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562210 | Headline | Cause |  | Contact with Parts Carrying Voltage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
937 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562211 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:32 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
938 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562212 | Requirement | Measure | label_Voltage_is_present | Warning Label - Voltage is present | Attach label at the Main Unit with safety sign: Label 3 of IEC 60601-1 Table D.2, dangerous voltage present Additionally it shall contian the following inscription:WARNING: Voltage is present after shutdown. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:26 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
939 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562213 | Headline | Cause |  | Isolation defect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
940 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562214 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:32 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
941 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562215 | Headline | Hazard |  | Hazards from fire or smoke |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Flammability,loss of basic safety or essential performance |  | S1M | VB10 | [12 empty] | NotDetermined
942 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562216 | Headline | Cause |  | Fire inside the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz35 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
943 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562217 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:32 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
944 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562218 | Requirement | Measure | um_HintsForFire | User Manual - Switch off the system in the event of fire | WARNING:Cause: Fire in or near the system.Possible consequences: Injury to patient and personnel, and damage to device. Risk of gas poisoning due to burning plastic.Remedy:- In case of fire switch off the system.- Be aware of and inform patients of escape routes.- Be aware of where fire extinguishers are located and know how to use them.- In case of fire switch off the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
945 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562219 | Headline | Hazard |  | Thermal energy and high temperature |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature,Flammability |  | S1M | VB10 | [12 empty] | NotDetermined
946 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562220 | Headline | Cause |  | Parts too warm |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
947 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562221 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:33 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
948 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562222 | Requirement | Measure | surface_temperature_applied_parts | Surface temperature applied parts | For all surfaces which can be touched by the patient (applied parts according to IEC 60601-1 cl. 4.6, <XREF>561933</XREF>), the surface temperatures under normal and single fault conditions according to IEC 60601-1, table 24 (see <XREF>434357</XREF>) for the following contact duration shall be applied:Affected components and areas:- Cover of the detector incl. FD handle and panel: 10 min <= t- Cover of the monoblock, excluded monoblock backside: 1 min <= t <10 min- Wired and wireless hand control: 10 min <= t- Remote control panel on the rear: t < 1 min (relatively heavy, therefore it is not expected to be placed on the patient directly, besides shortly putting it down, when mounted at the table, the patient cannot come into contact with the rear side)- Remote control panel on the top and the sides: 1 min <= t < 10 min- C-shaped mechanics of the system: 1 min <= t < 10 min | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 15.07.2022 14:43 | Nils Naarmann |  |  |  | HICTouchCon,MainU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
949 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562225 | Requirement | Measure | TemperatureSensor_SelfTest | Temperature sensor self test | During operating of the system there will be a self test of the temperature sensing circuit. In the case of a defect x-ray radiation will be prevented. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
950 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562226 | Requirement | Measure | FDTemperature_Switchoff_s1 | System Warning - Temperature of the Detector | The temperature inside the detector is checked quasi-continously (i.e. short interval) by FLC. According to the detector type an appropiate warning for the user is displayed, when first warning temperature is reached. When the second warning temperature is reached, an additional warning message is displayed to inform the user.Temperature for first warning message:- 52°C for FDmediumHTemperature for 2nd warning message:- 54°C for FDmediumHThe detector will not be switched off at any time, during an intervention. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FDmediumH,FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
951 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562228 | Requirement | Measure | um_tube_hot | User Manual - Hot tube | CAUTIONCause: Under extreme operating conditions (continuous operation close to the maximum tube voltage), the X-ray tube can become very hot. This can shorten the service life of the system components. Possible consequences: Risk of burns and prolongation of the operation! Remedy:- Avoid touching the X-ray tube housing.- Change your parameter settings, e.g. by reducing the dose level.- We also recommend keeping a second system ready for continuous radiation lasting longer than 30 minutes.- Please make sure that the system has cooled down before performing critical interventional applications. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
952 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562229 | Requirement | Measure | serv_um_tube_hot | ServiceDoc - Hot tube housing | Service DocumentsCAUTIONCause: In extreme operating mode (continuous operation), the temperature of the X-ray tube assembly can heat up.Possible consequences: Danger of burns.Remedy: Avoid contact with the tube housing | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
953 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562230 | Headline | Cause |  | Leakage of hot oil from tube assembly |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32#Haz16 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Leaking subst. that are carcinogenic, mutagenic, toxic | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
954 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562231 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:33 | Johannes Schmidt |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
955 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562232 | Requirement | Measure | PressureSwitch_s1 | Pressure switch in tube | Pressure switch in tube is monitored during exposure. If excess pressure is detected, the switch is opened and the radiation is switched off. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
956 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562233 | Headline | Hazard |  | Injuries due to laser radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
957 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562237 | Headline | Hazard |  | Patient is infected by a virus due to a bad disinfection, system is damaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection |  | S1M | VB10 | [12 empty] | NotDetermined
958 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562238 | Headline | Cause |  | Incorrect cleaning of system after each patient, liquids or blood enter the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33#Haz42 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
959 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562239 | Headline | Measure |  | Reference - Spillage of Liquids on System | <XREF>562238</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
960 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562240 | Requirement | Measure | um_CleaningOfMaterial | User Manual - Cleaning and draping | WARNINGCause: Fluid or blood seeps into the system, use of inappropriate cleaning agents.Possible consequences: Risk of damage to the system, risk of infection or electrical hazard.Remedy: Use sterile or sterilizable covers at least in C-arm area (C-arm, cover of the flat detector and X-ray monobloc). Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. Only use the recommended substances for cleaning and disinfection. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 14:04 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
961 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562241 | Headline | Hazard |  | Environmental hazard |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Toxicity of chemical constituents |  | S1M | VB10 | [12 empty] | NotDetermined
962 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562242 | Headline | Cause |  | Explosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | D
963 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562243 | Headline | Measure |  | Reference - User Manual - Danger of explosion | <XREF>562006</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
964 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562244 | Headline | Hazard |  | Chemical hazards |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Toxicity of chemical constituents,Inadequate Disposal and scrapping |  | S1M | VB10 | [12 empty] | NotDetermined
965 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562245 | Headline | Cause |  | Toxic materials and fluids |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz38 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
966 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562246 | Requirement | Measure | serv_Scrapping_s1 | Technical instruction for the disposal procedure | A technical instruction for the disposal procedure shall exist, which describes which parts must be removed prior to scrapping and how these parts are disposed in an environment friendly manner. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 10:07 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
967 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562247 | Headline | Hazard |  | Electro magnetic disturbances |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC) |  | S1M | VB10 | [12 empty] | NotDetermined
968 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562248 | Headline | Cause |  | Delayed or disturbed transmission of images, or other functional disturbance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51 #Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
969 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562249 | Headline | Measure |  | Reference - EMC compliance - emissions, immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.03.2022 14:00 | Rudolf Siebenhaar | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
970 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562250 | Headline | Hazard |  | Danger by releasing of stored mechanical energy |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy |  | S1M | VB10 | [12 empty] | NotDetermined
971 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562251 | Headline | Cause |  | Rotation of c-arm when power fails |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
972 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562252 | Requirement | Measure | um_Tube_Rotation | User Manual - Rotation of c-arm | CAUTIONCause: C-arm rotation.Possible consequences: Risk of crushingRemedy: Pay particular attention to crushing risks between moving system parts and the corresponding guide openings. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
973 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562253 | Headline | Cause |  | Cable c-arm falling down |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
974 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562254 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:49 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
975 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562255 | Headline | Cause |  | Collimator comes loose |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
976 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562256 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:50 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
977 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562259 | Headline | Hazard |  | Collision with operator, patient or equipment, crushing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
978 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562260 | Headline | Cause |  | Unintentional activation of control elements |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
979 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578758 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
980 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562262 | Headline | Cause |  | Manual or motorized movements of any parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
981 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562263 | Requirement | Measure | um_Risk_of_Crushing_s1 | User Manual - Risk of Crushing | CAUTIONCause: Manual or motorized movementsPossible consequences: Risk of crushing and collision for persons and objects in the vicinity of moving parts.Remedy:- During manual and motorized movements of the device, make sure you take into account any persons or objects in the way of movement, and control the movement to avoid collisions.- Make sure that everyone is outside the hazardous zone. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
982 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562264 | Requirement | Measure | um_danger_zones_s1 | User Manual - Danger zones | CAUTIONCause: The points marked on the figure indicate hazardous locations around the system.Possible consequences: Risk of injury to the patient and personnel due to crushing or collision.Remedy: Be careful around the hazardous locations indicated. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
983 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562265 | Requirement | Measure | motorized_orbital_movement_limit | Limit motorized orbital operating radius | After saving the software end positions for motorized orbital movement, it shall be checked that the orbital movement stops so that the remaining gap to the end stop is >= 25 mm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
984 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562266 | Requirement | Measure | label_warning_crushing_hand2 | Warning Label - Risk of crushing hand | Warning labels "risk of crushing - hand" shall be placed on the angulation cover (upper and lower side), for orbital movement. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:27 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
985 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562267 | Headline | Cause |  | Motorized movements of lifting column |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
986 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562268 | Headline | Measure |  | Reference - User Manual - Risk of crushing | <XREF>562263</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
987 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562269 | Headline | Measure |  | Reference - User Manual - Danger Zones | <XREF>562264</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
988 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562270 | Headline | Measure |  | Reference - User Manual - Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
989 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562271 | Requirement | Measure | label_warning_crushing_foot | Warning Label - Risk of crushing foot | Warning labels "risk of crushing - foot" shall be placed on the cover of the transport skid on the front side, as well as on and the left and right side. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:32 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
990 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562272 | Requirement | Measure | label_warning_crushing_hand | Warning Label - Risk of crushing hand | A warning label ""risk of crushing - hand"" shall be placed on the upper main unit cover to indicate crushing between the vertical movement unit. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:31 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; Cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
991 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562273 | Headline | Hazard |  | Injuries because of falling parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Suspending parts,Instability,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
992 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562274 | Headline | Cause |  | Wrong calculated max load, unforeseeable load. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
993 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562275 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:50 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
994 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562276 | Headline | Hazard |  | Injury or crushing due to emergency stop out of order |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
995 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562277 | Headline | Cause |  | Defect in button or cable harness |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
996 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562279 | Requirement | Measure | serv_CheckEmergencyStop | Check emergency stop | Check emergency stop during periodic maintenance. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
997 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562280 | Headline | Hazard |  | Operator fails to actuate emergency stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Workflow and configuration,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
998 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562281 | Headline | Cause |  | Emergency stop is not readily accessible and easily identifiable |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:42 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
999 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578759 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1000 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578760 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1001 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562284 | Requirement | Measure | um_LocationEmergencyStop | User Manual - Emergency stop location | CAUTIONCaution: Position of EMERGENCY STOP button not known.Possible consequences: In an emergency, the EMERGENCY STOP button will not be actuated in a timely manner.Remedy: Make sure you know where the EMERGENCY STOP buttons are located. EMERGENCY STOP buttons are in the following locations (see the figure). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1002 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562285 | Headline | Hazard |  | Malfunction of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1003 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562286 | Headline | Cause |  | Inappropriate accessory |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. The operator manual warns to not perform non-approved modifications of the system, incl. combining with non-approved accessories or other devices. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1004 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562287 | Requirement | Measure | um_Appropiate_Accessory | User Manual - Inappropriate accessory | CAUTIONCause: Inappropriate accessories.Possible consequences: The use of accessories that do not comply with the safety requirements of this equipment can reduce the safety of the entire system.Remedy: Use only original Siemens Healthineers accessories or accessories approved by Siemens Healthineers. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:20 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1005 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562288 | Headline | Cause |  | Any defect parts in the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1006 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562289 | Requirement | Measure | um_system_defect | User Manual - Defect of system components | CAUTIONCause: System component error.Possible consequences: An error on one system component can reduce the safety of the entire system.Remedy:- Terminate system operation.- Contact Siemens Healthineers Customer Service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:20 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1007 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562290 | Headline | Hazard |  | Unintentional Radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1008 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562291 | Headline | Cause |  | Exposure release switch is activated by mistake when wired control is connected to the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1009 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562295 | Headline | Function |  | Stop button |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1010 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562296 | Headline | Hazard |  | Emergency stop out of order |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface |  | S1M | VB10 | [12 empty] | NotDetermined
1011 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562297 | Headline | Cause |  | Defect in button or cable harness |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1012 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562298 | Headline | Measure |  | Reference - Check emergency stop - maintenance | <XREF>562279</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1013 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562299 | Headline | Function |  | Clean system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1014 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562300 | Headline | Hazard |  | Patient collision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1015 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562301 | Headline | Cause |  | Swapped wireless footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1016 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562302 | Headline | Measure |  | Reference to measures regarding footswitch | See <XREF>563088</XREF>, <XREF>563089</XREF>, <XREF>563090</XREF>, <XREF>563091</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:58 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1017 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562303 | Headline | Hazard |  | Unwanted radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1018 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562304 | Headline | Cause |  | Swapped wireless footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1019 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562305 | Headline | Measure |  | Reference to measures regarding footswitch | See <XREF>563088</XREF>, <XREF>563089</XREF>, <XREF>563090</XREF>, <XREF>563091</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 10:58 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1020 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562306 | Headline | Function |  | Store patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1021 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562307 | Headline | Hazard |  | Unnecessary minor exposition to radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing,data transfer,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
1022 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562308 | Headline | Cause |  | Image store does not keep pace with acquisition under error conditions |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing,Loss of data | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1023 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562309 | Requirement | Measure | warning_message_no_saving_possible | Warning Message | When the image storage does not keep pace with the acquisition (e.g., due to buffer issues), radiography imaging shall be stopped and a message shall be displayed to the operator. Fluoroscopy imaging shall not be stopped or inhibited. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 11.10.2022 10:09 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | empty | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1024 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562310 | Headline | Cause |  | Storage Limit Reached (License) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing,Loss of data | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1025 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591790 | Headline | Measure | NotReady_s1 | Reference - Storage capacity exhausted | <XREF>562612</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2021 | 03.09.2021 13:12 | Judith Welzel |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1026 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 591791 | Headline | Measure | DiskCapacity1_s1 | Reference - Display storage space | <XREF>562613</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2021 | 03.09.2021 13:12 | Judith Welzel |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1027 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562312 | Headline | Hazard |  | Wrong interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing,data transfer |  | S1M | VB10 | [12 empty] | NotDetermined
1028 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562313 | Headline | Cause |  | Image quality cannot be improved after acquisition | Image quality in post processing cannot be improved after acquisition was performed with unsuited window values. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:45 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz23 #Haz41 #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1029 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562314 | Requirement | Measure | store_raw_data | Store raw data | The imaging system shall store acquired 2D images immediately to the hard disk, before any user manipulation is done. The operator shall always be able to restore the original image visualization. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1030 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562315 | Headline | Chapter |  | Connect to Network |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1031 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562316 | Headline | Function |  | HIS/RIS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1032 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562317 | Headline | Hazard |  | Non-harmonized configuration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1033 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562318 | Headline | Cause |  | Incorrect Configuration of HIS/RIS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz22 #Haz40 | See general justifications in the introduction which apply here. The operator manual advises how to check the correct configuration. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | System | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1034 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562319 | Requirement | Measure | docUser_HISRIS_configuration | User Manual - HIS/RIS Configuration | CAUTIONCause: Non-harmonized configuration of HIS/RISPossible consequences: Impossibility of CommunicationRemedy: Do not use configuration "patient ID only" within HIS/RIS as Imaging System (FLC) identifies a patient by quadruple (patient name, ID, sex, date-of-birth). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1035 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562320 | Headline | Function |  | Transfer exam |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1036 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562321 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Lack of integrity of data,user interface |  | S1M | VB10 | [12 empty] | NotDetermined
1037 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562322 | Headline | Cause |  | Network disturbance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1038 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562323 | Requirement | Measure | Network_Problems | Use of reliable data transfer standards | For transfer of dicom images (e.g. to a pacs or a share) reliable data transfer standards and protocols shall be used. Error messages in case of network problems in particular by transferring DICOM images to specific destinations(e.g. Pacs, Share,..). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1039 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562324 | Requirement | Measure | Local_Deleting_Data | Prompt by local deleting of data | Confirmation by user when deleting studies where nothing has been archived or printed. Message: ""Do you want unarchived studies to be deleted?""The user shall have the possibility to protect studies against deletion. Note:Information on archiving status will be set image by image (during transfer). As soon as one image out of study was sent successfully the study as a whole will be considered as archived. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1040 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562325 | Requirement | Measure | NXS_Xray_images_data_integrity | Data integrity of xray image transfer via NXS | For transfer of xray images via NXS the data integrity of these images shall be considered by using a check sum algorithm and a ID/counter of transfered images. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1041 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562326 | Requirement | Measure | NXS_heartbeat | NXS protocol with heartbeat | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1042 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562327 | Requirement | Measure | NXS_heartbeat_extapp | NXS protocol with heartbeat for extapp | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1043 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562328 | Requirement | Measure | NXS_Xray_images_additional_info | Additional information of xray image transfer via NXS | For transfer of xray images via NXS some additional information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1044 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562329 | Requirement | Measure | NXS_Xray_images_additional_info_extapp | Additional information of xray image transfer via NXS for extapp | For transfer of xray images via NXS some additioinal information (like patient quadrupel) shall be incorporated to identify clearly which xray image belong to which patient and the properties of the xray image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1045 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568849 | Requirement | Measure | compression_threshold | Image compression threshold | For static images, the grey scale value of a pixel may vary by +/- 2% of the original grey scale value.For live images, the image quality may be further reduced, keeping the frame rate. The static app UI content shall achieve the above mentioned threshold. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 01.09.2020 | 23.02.2022 08:50 | Sonali Rath |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1046 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569294 | Headline | Measure | image_display_delay_time_ext | Reference - Image display delay time - external | <XREF>568835</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1047 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562330 | Headline | Function |  | Connectivity |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1048 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562331 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1049 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562332 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1050 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562333 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1051 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562334 | Headline | Chapter |  | Perform Clinical Procedure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1052 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562335 | Headline | Function |  | Position Patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1053 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562336 | Headline | Hazard |  | Impaired vision through laser-beam |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation,Inappropriate accessory |  | S1M | VB10 | [12 empty] | NotDetermined
1054 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562337 | Headline | Cause |  | People looking into laser-beam |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz34 | See general justifications in the introduction which apply here. The operator manual advises to take care of the laser light. No additional practical mitigations are available to further reduce or prevent eye exposure to the laser beam. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1055 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562338 | Requirement | Measure | usage_laser_class1 | Usage of laser class 1 | Usage of laser class 1 (according to IEC 60825-1:2014) in monoblock, detector and horizontal laser, and LIDAR scanner. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:21 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1056 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562342 | Headline | Hazard |  | Collision with patient, operator or equipment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1057 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562347 | Headline | Cause |  | Unintended movement activation when wired control element is connected to the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1058 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578761 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1059 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578766 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1060 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562351 | Headline | Cause |  | Movement without intentional activation on remote console |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1061 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578762 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1062 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578767 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1063 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562354 | Headline | Cause |  | Movement does not stop |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1064 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578768 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1065 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578320 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1066 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578763 | Headline | Measure | um_risk_collision_general | Reference - Warning - General risk of collisions | <XREF>563493</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1067 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562358 | Headline | Hazard |  | Electrical shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage,Line voltage,Leakage current |  | S1M | VB10 | [12 empty] | NotDetermined
1068 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562359 | Headline | Cause |  | Protective earth issues |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1069 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562360 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:03 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1070 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562361 | Headline | Measure |  | Reference - Start up instruction | <XREF>562209</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1071 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562362 | Headline | Cause |  | Contact with parts carrying voltage. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1072 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562363 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:04 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1073 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562364 | Headline | Cause |  | Defective insulation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1074 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562365 | Requirement | Measure | IEC60601_1_cl8-8_dielectric_strength | IEC 60601-1 Clause 8.7 and 8.8 | The system shall comply with IEC 60601-1 cl. 8.7 and 8.8.Voltages which are accessible without tool shall not exceed 42,4 V peak a. c. and 60 V d. c.Touch current and patient leakage currents shall comply with cl. 8.7The operator shall be protected by 2 MOOP.The patient shall be protected by 2 MOPP.Applied parts of type B according to cl. 4.6 are:- Cover of the detector incl. FD handle and panel- Cover of the monoblock, excluded monoblock backside- Wired and wireless hand control- Remote control panel- C-shaped mechanics of the systemNo parts are identified which are classified as type BF or CF. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | AccesS,DapS,DSMonS,HICSwitch,HICTouchCon,MainU,MCU,PowerU,TrolleyU,XGen,SysCabl,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1075 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562366 | Requirement | Measure | IEC60601-1_cl9-8_mechanical_strength | IEC 60601-1 Clause 9.8 | Parts of the system that are suspended or supported shall be developed using the tensile safety factors listed in IEC60601-1:2005, Table 21. These parts shall fulfill the requirements of section 9.8 of the standard. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:23 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1076 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562367 | Headline | Hazard |  | Burns |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1077 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562368 | Headline | Cause |  | Hot surfaces |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1078 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562369 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | ompliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:04 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1079 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562370 | Headline | Hazard |  | Injuries because of falling parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Suspending parts,Instability |  | S1M | VB10 | [12 empty] | NotDetermined
1080 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562371 | Headline | Cause |  | Inadequate layout or construction of mechanical system parts. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1081 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562372 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 Standards | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:05 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1082 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562373 | Headline | Measure |  | Reference to mounting of the monoblock | See HM_S1_mec_F4G1U1M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1083 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562374 | Headline | Function |  | Position System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1084 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562375 | Headline | Hazard |  | Crushing through motor-driven movement |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1085 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562376 | Headline | Cause |  | Unintentional activation of movement |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:49 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1086 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562380 | Headline | Measure |  | No additional measure required - Reference to set values motorized speed | <XREF>563486</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:37 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1087 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562381 | Headline | Cause |  | Malfunction of the control element |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1088 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578094 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 30.01.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1089 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562383 | Headline | Cause |  | Falling below the safety distance during vertical movement |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1090 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562384 | Requirement | Measure | stop_vertical_lift | Stop vertical lift | When the system is moved downwards, the movement shall be stopped at a safety distance of min. 120 mm between the floor and the C-arm parts (either detector or monoblock rotated downwards). Note: The crushing distance for feet between C-arm and front side of the wheel cover of the chassis is allowed to be below 120 mm due to the crushing label on the cover. The crushing distance for toes of 50 mm is met. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:41 | Rudolf Siebenhaar |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1091 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562385 | Requirement | Measure | acoustic_signal | Acoustic signal | After reaching the safety distance in <XREF>562384</XREF>, an acoustic alarm signal is triggered, when the C-arm is moved further downwards in direction of crushing hazard. When a distance of min. 25 mm between horizontal unit and the top of the chassisis reached, downwards movement shall be stopped and prevented. Note: Crushing distance for toes between floor and C-arm or C-arm and front side of the wheel cover of the chassis is allowed to be below 50 mm due to acoustic signal and crushing label. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.04.2022 13:28 | Rudolf Siebenhaar |  |  |  | MainU,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1092 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562386 | Headline | Measure |  | Reference - User Manual - Guide Openings | <XREF>562444</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1093 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562387 | Headline | Measure |  | Reference - User Manual - Risk of Crushing | <XREF>562263</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1094 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562388 | Requirement | Measure | docUser_note_before_starting_treatment | User Manual - Check Audio Signal | CAUTIONCause: Audible warning for safety distance not functionalPossible consequences: Risk of crushingRemedy: Check for proper loudspeaker function before beginning the examination. Note when booting the system whether an audible signal is heard. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.02.2022 11:16 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1095 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562389 | Requirement | Measure | docSCM_crushing_zone_switch | Supply Chain Documentation - Assembly Crushing Zone Switch | Ensure location of crushing zone switch during final inspection.Measure the safety distances defined after the intermediate switch off, see <XREF>562384</XREF> and after the final switch off, see <XREF>562385</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 17:08 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1096 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562390 | Requirement | Measure | docUser_meaning_audio_signals | User Manual - Meaning of Audio Signals | CAUTIONCause: Occurrence of audible signals.Possible consequences: Hazards possible, for example due to crushing risks, extended radiation release, and high dose rate.Remedy:- Please note that continued system operation may lead to one of these hazards.- Use the system with extreme caution to avoid these hazards. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1097 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562391 | Headline | Cause |  | Movement speed too high |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1098 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578771 | Headline | Measure | EmergencyStopMovement_s1 | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | HICTouchCon,MCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1099 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562394 | Headline | Cause |  | Overtravel (stopping distance) too long |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1100 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578326 | Headline | Measure | max_overtravel_values | Reference - Maximum overtravel values | <XREF>562465</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 08.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU,MCU |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1101 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562396 | Requirement | Measure | force_limit_ang_orb | Force limit orbital / angular motorized movement | The system shall stop motorized movement if the holding force exeeds 200 N for angular and orbital movements. | ToBeReviewed | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 23.09.2022 13:09 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1102 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562398 | Headline | Hazard |  | Crushing through arbitrary movement of individual axles |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Magnetic fields,Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1103 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562399 | Headline | Cause |  | Weight causes drive to lower autonomously |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1104 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562400 | Requirement | Measure | provide_brake_or_self_locking | Provide brake or self-locking of the drive | The drive of the vertical movement must be designed in a way that prevents an autonoums movement, e.g. to provide a brake or self-locking of the drive. No autonomous movement is given, as long as a lowering within one hour does not exceed 2 mm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:40 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1105 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562401 | Headline | Cause |  | C-arm not weight balanced |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1106 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562402 | Requirement | Measure | design_c_arm_weight_balanced | Design C-arm weight balanced | Design of C-arm shall be weight balanced to prevent autonomous movement with brakes open. Note:A remaining movement may be observable. Range of remaining movement depends on angular/orbital position and shall not exceed 20 degree. Nonetheless, remaining movement shall be uncritical for the user. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 14.03.2022 08:27 | Johannes Schmidt |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1107 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562403 | Headline | Cause |  | Additional weight changes weight balance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1108 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562404 | Requirement | Measure | docUser_note_loss_of_weight_balance | User Manual - Unbalanced C-arm | CAUTIONCause: Changing the weight distributionPossible consequences: Risk of crushingRemedy: Adding weight to the detector or single tank side means a loss in counterweight and can result in unintended movement of the C-arm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1109 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562405 | Headline | Cause |  | Weakening of magnetic force |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1110 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562406 | Requirement | Measure | docServ_periodic_maintenance_magnetic_force | Service Instruction - periodic maintenance | Demagnetization may appear, if one of following actions will affect on OR table holder1) external magnetic field (magnetic hysteresis, coercivity field Hc),2) strong heating (threshold: Curie temperature), and/or3) strong hits/shocks There is no external magnetic field available. Temperature thresholds to demagnetize are 310-340° Celsius (Neodym or Neodym-Iron-Boron combinations, NdFeB) 354 degree Celsius (Nickel, Ni), and 450° Celsius (Ferrite, i.e. Fe2O3 or Fe3O4 combinations), and 770° Celsius (Iron, Fe).These environmental conditions will not appear in lifetime.Remaining issue are (3) strong hits/shocks. These seems to be very unlikely that this kind of force will be applied to the OR table holder, but can be assumed as improbable, but not impossible. Therefore, magnetic force of OR table holder shall be checked during periodic maintenance: The magnets must fix the RCP so that pulling them with up to 150N will not loosen the RCP from the OR table holder."or "The magnets shall fix the RCP so that, when pulling on it with at least 150N, the RCP will not loosen from the OR holder. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1111 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562407 | Headline | Cause |  | Angular / orbital brake not functional |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1112 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562408 | Headline | Measure |  | Reference to Caution Operator Manual - Check brakes | <XREF>561962</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1113 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562409 | Headline | Measure |  | Reference to weight balanced design | <XREF>562402</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1114 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562413 | Headline | Hazard |  | Crushing through breakage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity,Fatigue failure |  | S1M | VB10 | [12 empty] | NotDetermined
1115 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562414 | Headline | Cause |  | Carrying capacity of the unit too small |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | High | A
1116 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562415 | Requirement | Measure | design_movement_unit | Design of the movement unit | The moving positioning units must be designed so that a overload can be carried in compliance with DIN EN 60601-1 (2006), Kap. 9.8 , Table 21. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:23 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1117 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562416 | Headline | Cause |  | Motor-driven movement and collision with fixed parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | High | A
1118 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562417 | Requirement | Measure | design_movement_unit_collision | Design of the movement unit collision | The moving positioning units must be designed so that a overload due to a collision can be handled. Remark:The component parts shall be carrying a overload in compliance with DIN EN 60601-1 (2006). Kap. 9.8, Table 21. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:24 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1119 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562418 | Headline | Cause |  | Breakage or wear of the spindle nut |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1120 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562419 | Requirement | Measure | existing_additional_safety_equipment | Additional safety mechanism - lift column | There shall be a mechanical safety mechanism for the lift column which prevents the C-arm from crashing down in case of a spindle failure.The mechanical safety mechanism shall be designed to fulfill a tensile safety factor >= 4.The overtravel (stopping distance from the time of engagement to the time of no further movement) shall be max. 5 mm.Note: IEC 60601-1, cl. 9.8.4 Systems with MECHANICAL PROTECTIVE DEVICES | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:24 | Rudolf Siebenhaar |  |  |  | MainU | IEC 60601-1, cl. 9.8.4 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1121 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562420 | Requirement | Measure | activation_of_safety_feature_blocks_drive | Activation of safety mechanism blocks lift drive | After the activation of the mechanical safety mechanism <XREF>562419</XREF> of the vertical column, the operation of the lift drive shall not be possible anymore.Note: The drive is not capable of moving again due to the activated mechanism, and an motor overcurrent message is displayed. The column can be manually screwed down with a tool (ratch) only which can be inserted in a whole at the bottom of the chassis (moving it down is necessary for transporting it back to the factory). Repair is only done in factory, not in the field. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:25 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1122 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562421 | Requirement | Measure | labeling_of_safety_equipment | Labeling of the safety equipment | A labeling, safety sign ISO 7010-W001, has to be placed next to the mechanical safety equipment or in a way so the connection to it is obvious. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.10.2022 09:33 | Johannes Schmidt |  |  |  | MainU | IEC 60601-1; cl. 7.2.13 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1123 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562422 | Requirement | Measure | docUser_note_mechanical_safeguard_effected | User Manual - Spindle Nut | WARNING:Cause: Spindle nut broken or worn.Possible consequences: Risk of crushingRemedy: Vertical lift is blocked if the mechanical safeguard was triggered. Contact Siemens Healthineers Customer Service to replace the mechanical safeguard. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:21 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1124 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562423 | Headline | Hazard |  | Risk of crushing or falling parts, impaired mechanical integrity |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,loss of basic safety or essential performance,Unexpected loss of electrical/mechanical integrity |  | S1M | VB10 | [12 empty] | NotDetermined
1125 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562424 | Headline | Cause |  | Overtravel / movement range exceeds range |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1126 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562425 | Requirement | Measure | end_stop | End stop | Every movement axle shall be equipped with mechanical end stops at the end of the movement range. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:27 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1127 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562426 | Headline | Cause |  | Automatic 3D-scan |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1128 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578319 | Headline | Measure | dead_mans_push_buttons | Reference - Continuous activation of movements with redundant signals | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.02.2021 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICButtonCon,HICSwitch,HICTouchCon,MCU,HICForceControl,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1129 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562428 | Headline | Cause |  | Weight of the device |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1130 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562429 | Requirement | Measure | docUser_note_fast_braking_more_difficult | User Manual - Fast Braking | CAUTIONCause: Weight of the devicePossible consequences: Risk of crushingRemedy: Fast braking can be impacted by the weight of the device. Pay attention to your speed as well as the floor covering, slope, and uneven areas. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1131 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562430 | Headline | Cause |  | Mounting RCP on OR table holder |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1132 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562431 | Requirement | Measure | OR_table_holder_crushing_zones | Design without crushing zones | The OR table holder shall have a flat design (e.g. without beadings), so size of emerging gaps between OR table holder and mounted RCP are above threshold (> 25 mm), in accordance to finger gaps for adults in IEC60601-1:2005 clause 9.2.2.2 (Table 20).In case of a non-flat design of OR table holder a protection for all possible finger positions during mounting of RCP shall be available. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 17:19 | Johannes Schmidt |  |  |  | AccesS,MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1133 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562432 | Headline | Hazard |  | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1134 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562433 | Headline | Cause |  | Brake failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1135 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562434 | Requirement | Measure | docUser_note_befor_starting_daily_inspection | User Manual - Daily Inspection | CAUTIONCause: Brake failurePossible Consequences: Crushing due to arbitrary movementsRemedy: Before beginning the examination, perform the daily function and safety checks. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1136 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562435 | Headline | Cause |  | Braking power not sufficient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1137 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562436 | Requirement | Measure | specify_sufficient_braking_power | Specify sufficient braking power | Specify sufficient braking power. This is in compliance with IEC 60601-1:2005; clause 9.4.3 (Instability). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 08.04.2022 14:02 | Rudolf Siebenhaar |  |  |  | MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1138 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562437 | Requirement | Measure | docUser_parking_chassis_incline | User Manual - Parking Chassis on inclined Plane (i.e. Ramp) | CAUTIONCause: Insufficient braking force.Possible consequences: Risk of crushing!Remedy: When parking the C-arm system on an inclined surface drive the vertical colum completely down. Verify that the brakes are locked and that the main unit does not slide away. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 01.06.2022 10:23 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1139 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562439 | Headline | Hazard |  | Risk of crushing during movement execution |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1140 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562440 | Headline | Cause |  | Distance between operator rail and other components too small |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1141 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562441 | Requirement | Measure | compliance_to_DIN_EN_60601_2_32 | Compliance with distances according to IEC 60601-1 | Compliance with DIN EN 60601-1 (2006) Kap. 9.2 Table 20, Acceptable gaps for finger > 25mm | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:33 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1142 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562442 | Requirement | Measure | docUser_note_point_of_crushing | User Manual - Distance Operator Rail | CAUTIONCause: The distance between the handles and other components is too small.Possible consequences: Risk of crushing!Remedy: Be careful around the hazardous locations indicated. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1143 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562443 | Requirement | Measure | docUser_clearance_of_system | User Manual - Appropriate Distance | CAUTIONCause: The distance between the patient or other persons and the system is too small.Possible consequences: Risk of crushing!Remedy: Use the system only if the patient and other persons can be observed during movements. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1144 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562444 | Requirement | Measure | docUser_daily_inspection | User Manual - Guide Openings | CAUTIONCause: C-arm rotation.Possible consequences: Risk of crushingRemedy: Pay particular attention to crushing risks between moving system parts and the corresponding guide openings. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1145 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562445 | Headline | Cause |  | Activation of movement without intervisibility |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:52 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1146 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562446 | Requirement | Measure | docUser_note_intervisibility | User Manual - Observe Patient/Persons during Movement | CAUTIONCause: Activation of system movements without visual contactPossible Consequences: Risk of crushingRemedy: Use the system only if the patient and other persons can be observed during movements. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1147 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562447 | Headline | Hazard |  | Patient cannot be salvaged (e.g. patient is unconscious or anesthetized) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts,Moving and positioning patient |  | S1M | VB10 | [12 empty] | NotDetermined
1148 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562448 | Headline | Cause |  | Orbital drive is failing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 12.10.2022 09:16 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz46 | See general justifications in the introduction which apply here. No additional practical mitigations are available to further reduce or prevent the risk of loss of movement functionality. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1149 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562450 | Headline | Hazard |  | Injuries through OR-instruments |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1150 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562451 | Headline | Cause |  | C-arm movements during operations |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz03 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | B | Med | B
1151 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562452 | Requirement | Measure | docUser_note_collisions_with_other_OR_instruments | User Manual - Collision with OR Instruments | WARNINGCause: C-arm movements during surgeryPossible consequences: Risk of injuriesRemedy: Due to the maneuverability of the C-arm, collisions with other surgical devices are possible if the device is not used as intended and in accordance with the instructions in this manual. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1152 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562453 | Headline | Hazard |  | Misjudgment due to poor image quality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1153 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562454 | Headline | Cause |  | calibration incorrect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1154 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562455 | Requirement | Measure | calibration_cycles | Calibration cycles | Calibration cycles are monitored, when exceeding, notification is sent to operator. Note:DEXI adjustment is not affected by mentioned calibration cycles. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1155 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562456 | Headline | Hazard |  | Injury through dropping parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1156 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562457 | Headline | Cause |  | Collision with other systems |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1157 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562458 | Requirement | Measure | docUser_note_collisions_with_other_systems | User Manual - Collisions due to misuse | CAUTIONCause: Collision with other systemsPossible consequences: Risk of injuryRemedy: Due to the maneuverability of the C-arm, collisions with other systems are possible if the device is not used as intended and in accordance with the instructions in this manual. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1158 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562468 | Headline | Hazard |  | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1159 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562469 | Headline | Cause |  | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 12.10.2022 09:14 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. No additional practical mitigations are available to further reduce or prevent the risk of loss of movement functionality. | Basic | n.a. | none | n.a. | True | empty | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1160 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562470 | Headline | Measure |  | No measure required | No measure required. | ToBeReviewed | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1161 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562471 | Headline | Hazard |  | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1162 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562472 | Headline | Cause |  | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. No additional practical mitigations are available to further reduce or prevent the risk of loss of movement functionality. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1163 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562473 | Headline | Measure |  | No measure required | No measure required. | ToBeReviewed | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1164 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562474 | Headline | Function |  | Wrap c-arm with Sterile Cover |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1165 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562475 | Headline | Hazard |  | Risk of infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory,Infection,Not sterile |  | S1M | VB10 | [12 empty] | NotDetermined
1166 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562476 | Headline | Cause |  | Improperly installed sterile drape |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1167 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562477 | Headline | Measure |  | Reference to User Manual - Cleaning and draping | <XREF>562240</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | empty | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1168 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562478 | Headline | Function |  | Wrap Foot Switch with Covers |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1169 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562479 | Headline | Hazard |  | Risk of infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory,Infection,Not sterile |  | S1M | VB10 | [12 empty] | NotDetermined
1170 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562480 | Headline | Cause |  | Improperly installed sterile drape |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33#Haz42 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1171 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562481 | Headline | Measure |  | Reference to User Manual - Cleaning and draping | <XREF>562240</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1172 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562482 | Headline | Function |  | Wrap RCP with Covers |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1173 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562483 | Headline | Hazard |  | Risk of infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory,Infection,Not sterile |  | S1M | VB10 | [12 empty] | NotDetermined
1174 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562484 | Headline | Cause |  | Damage of sterile drape |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18#Haz10 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1175 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562485 | Requirement | Measure | sharp_edge_sterile_cover | Avoid sharp edges | The control panel (RCP) as well as its holder, e.g. cart needs to be designed without sharp edges which could damage a sterile cover. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 29.07.2020 | 07.07.2022 10:11 | Rudolf Siebenhaar |  |  |  | AccesS,HICTouchCon,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1176 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562486 | Headline | Cause |  | Inappropriate positioning of RCP |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32#Haz10 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1177 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562487 | Requirement | Measure | docUser_avoid_RCP_on_patient | User Manual - Positioning of RCP | CAUTIONCause: Improper positioning of the remote control unit.Possible consequences: Risk of contaminationRemedy: Do not place the remote control unit on the patient. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1178 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562488 | Headline | Cause |  | Self-loosening of OR table holder |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32#Haz10 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | System Test | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1179 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562489 | Requirement | Measure | OR_table_holder_fixation_rail | Fixation on OR table rail system | The OR table holder shall be fixed on OR table rail system. Fixation needs to be designed to prevent unintended self-loosening. Note:As drapery of patient is trapped between OR table holder and RCP, an unintended self-loosening of OR table holder will undrape patient. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 14:08 | Johannes Schmidt |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1180 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562490 | Headline | Function |  | Acquire Image |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1181 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562491 | Headline | Hazard |  | Risk of burnings |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1182 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562492 | Headline | Cause |  | Increasing temperature due to extended use |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1183 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562493 | Requirement | Measure | provide_monoblock_housing | Provide Monobloc Housing | Monobloc needs to be covered with housing to prohibit direct contact of user/patients to parts of monoblock. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:33 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1184 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562494 | Headline | Measure |  | Reference - Pressure switch in tube | <XREF>562232</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1185 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562495 | Requirement | Measure | temperature_monitoring_monoblock | Temperature monitoring monoblock | The temperature of the monoblock shall be monitored. X-ray shall be inhibited when the maximum operating temperature is reached. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1186 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562497 | Headline | Cause |  | Oil leakage out of leaky monoblock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32#Haz16 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1187 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562499 | Requirement | Measure | pressure_compensation | Pressure Compensation Membrane | A pressure compensation membrane for oil expansion has to be integrated in monoblock. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1188 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578321 | Headline | Measure | temperature_monitoring_monoblock | Reference - Temperature monitoring monoblock | <XREF>562495</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1189 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562500 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:11 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1190 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562501 | Headline | Measure |  | Reference - Pressure switch in tube | <XREF>562232</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1191 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562502 | Headline | Cause |  | Error of cooling system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1192 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562503 | Headline | Measure |  | Reference - Temperature Control | <XREF>562495</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1193 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562504 | Headline | Cause |  | Longer operation heats up detector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1194 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562505 | Requirement | Measure | provide_protection_cover | Provide contact protection/access cover for detector | No direct contact possible due to access cover of the detector. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:34 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1195 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562507 | Headline | Hazard |  | Risk of infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Infection |  | S1M | VB10 | [12 empty] | NotDetermined
1196 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562508 | Headline | Cause |  | Oil leakage out of leaky monoblock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32#Haz16 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1197 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562509 | Requirement | Measure | stability_test_energy_input | Stability test | The monobloc must be designed in a way that it withstands excessive energy inputs certain malfunctions Remark:Test of X-ray tube compliance IEC 60601-2-28 (2010) Kap. 9.7 | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1198 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562510 | Headline | Measure |  | Reference - Pressure switch in tube | <XREF>562232</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1199 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562511 | Headline | Measure |  | Reference - Temperature control | <XREF>562495</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1200 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562512 | Headline | Measure |  | Reference - Check leakage of monobloc | <XREF>563289</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1201 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562513 | Headline | Cause |  | Cooling circuit leaky, cooling medium drips into OR-field |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | C
1202 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562514 | Headline | Measure |  | Reference to Cooling Unit Leakage Check | See <XREF>563289</XREF> and <XREF>563364</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1203 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562515 | Headline | Hazard |  | Injury through expelled parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy |  | S1M | VB10 | [12 empty] | NotDetermined
1204 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562516 | Headline | Cause |  | Breakage rotating anode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz37 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1205 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562517 | Headline | Measure |  | Reference - Stability test | <XREF>562509</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1206 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562518 | Headline | Cause |  | Implosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz37 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | High | D
1207 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562519 | Headline | Measure |  | Reference to stability test and constructive measures | See <XREF>562509</XREF> and HM_S1_mec_F4G3U1M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1208 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562520 | Headline | Hazard |  | Excessive temperature, fire, explosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy |  | S1M | VB10 | [12 empty] | NotDetermined
1209 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562521 | Headline | Cause |  | Overcharge of the capacitors for large energy storage / Elko |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz35#Haz15 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1210 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562522 | Headline | Measure |  | Reference - Mounted under covers | <XREF>562126</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1211 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562523 | Requirement | Measure | protection_energy_storage | Protection of the energy storage | Protection of the energy storage by a switch off of the short circuit in <= 20 ms.[HM_S1_esu_F6G3U1M1 in <XREF>578770</XREF>] | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1212 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562524 | Requirement | Measure | rapture_joint | Rupture joint by Overcharge | A rupture joint which breaks at an excessive overcharge must be available at the capacitors.[HM_S1_esu_F6G3U2M1 in <XREF>578770</XREF>] | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1213 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562525 | Requirement | Measure | pressure_refief_valve | Pressure relief valve | Pressure relief valve at the capacitors prevents an excessive pressurization.[HM_S1_esu_F6G3U2M1 in <XREF>578770</XREF>] | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1214 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562526 | Headline | Measure | pressure_refief_valve | Reference to monitoring of the intermediate circuit voltage | See HM_S1_hw_F6G3U2M2 for ESU and Elko, HM_S1_esu_F6G3U2M3 only for ESU in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1215 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562527 | Headline | Cause |  | Short-circuit of electrical components |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36#Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1216 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562528 | Requirement | Measure | supervision_short_circuit | Protection against excessive current and voltage | The system shall provide a protection against excessive current and voltage.Alternatively, a test shall show that a short circuit does not result in a hazardous situation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1217 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562529 | Requirement | Measure | supervision_short_circuit_PowerU | Protection against excessive current and voltage - PowerUnit | The PowerUnit shall provide a protection against excessive current and voltage. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | PowerU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1218 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562530 | Headline | Measure |  | Reference to generator measures to UL listed materials and fuses in intermediate circuit | See HM_S1_ul_F2G1U1M1 and HM_S1_hw_F2G1U1M2 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | PowerU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1219 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562531 | Headline | Hazard |  | Implosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy |  | S1M | VB10 | [12 empty] | NotDetermined
1220 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562532 | Headline | Cause |  | Collision/impact against component |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz37 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-C | 3-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | High | B
1221 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562533 | Requirement | Measure | mechanical_strength_of_the_X_ray_generation | Mechanical strength of the X-ray generation | Mechanical strength of the monobloc has to be verified by the manufacturer | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1222 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562534 | Headline | Measure |  | Reference - Stability test | <XREF>562509</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1223 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562535 | Headline | Cause |  | Breakage rotating anode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz37 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1224 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562536 | Requirement | Measure | housing_has_to_stay_integrally | Housing has to stay integral | Housing has to stay integral, if there is a breakage of the rotating anode. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1225 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562537 | Headline | Measure |  | Reference - Stability test | <XREF>562509</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1226 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562538 | Headline | Measure |  | Reference to constructive measures | See HM_S1_mec_F4G3U1M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1227 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562539 | Headline | Hazard |  | Risk of injury due to bad surface finish |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Sharp edges and points |  | S1M | VB10 | [12 empty] | NotDetermined
1228 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562540 | Headline | Cause |  | Sharp edges |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1229 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562541 | Headline | Measure | develop_design_without_sharp_edges | Develop design without sharp edges | <XREF>562047</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1230 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562542 | Headline | Cause |  | Rough surface finish |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1231 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562543 | Headline | Measure |  | Reference - Specify smooth surface finishes | <XREF>562050</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:39 | Rudolf Siebenhaar | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1232 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562544 | Headline | Hazard |  | Injury due to damaged housing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1233 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562545 | Headline | Cause |  | Transport damages e.g. collision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1234 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562546 | Requirement | Measure | docUser_note_professional_repair_remains | User Manual - Professional Repair | CAUTIONCause: Transport damagePossible consequences: Risk of injuryRemedy: Carry out professional repairs only, otherwise consequential damage to the product cannot be ruled out! | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1235 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562547 | Headline | Hazard |  | Electric shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage |  | S1M | VB10 | [12 empty] | NotDetermined
1236 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562548 | Headline | Cause |  | Intrusion of liquids |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1237 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562549 | Requirement | Measure | compliance_with_a_spillage_test | Spillage of Liquids on System | Spillage of liquids on system shall not result in a hazardous situation. The system shall be designed in accordance to IEC 60601-1 clause 11.6.3 (compliance to spillage test). The spillage test shall be performed with following instructions: Spillage test, scenario 1: * Scenario: capacity of drapes/ max. suction performance exhausted * liquid is seeping continuously from operating area downwards on the system * Quantity: 500 ml; duration: 30 min * Effected parts of system: FD/I.I., monoblock, C-arm (without Chassis) Spillage test, scenario 2: * Scenario: splatter of blood * liquid is splashing from operating area up- and/or sidewards to the system * Quantity: 20 ml; duration: approx. 1 s (i.e. a bolus) * Effected parts of system: FD/I.I., monoblock, C-arm (without Chassis) No liquids need to be spilled on trolley (incl. trolley control panel, TCP). Hints:Spilling of FD is only allowed with coverage of drape (see <XREF>562552</XREF>).Remote control panel (RCP) will be used with drapes only. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:16 | Johannes Schmidt |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined
1238 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562552 | Requirement | Measure | docUser_use_FD_drapes | User Manual - Draping of FD | WARNING:Cause: Fluid penetration and contamination.Possible Consequences: Damage to the flat detectorRemedy: Make sure that the detector is always covered by a disposable cover during the examination. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | System | S1M | VB10 | [12 empty] | NotDetermined
1239 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562553 | Headline | Hazard |  | Unwanted radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1240 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562554 | Headline | Cause |  | No radiation protection behind the entering side of radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1241 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562555 | Requirement | Measure | compliance_with_DIN_EN_60601_1_3 | Compliance with DIN EN 60601-1-3 | Compliance with DIN EN 60601-1-3 (2008); Kap. 11 residual radiation | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:37 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1242 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562556 | Headline | Cause |  | Increased leakage radiation monoblock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1243 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562557 | Headline | Measure |  | Reference to constructive measures and tests to comply with leakage radiation requirements | See HM_S1_mec_F5G1U3M1, HM_S1_manu_F5G1U3M2, HM_S1_mec_F5G1U3M3 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1244 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562558 | Headline | Cause |  | Stay in endangered area |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1245 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562559 | Requirement | Measure | docUser_note_stand_in_a_proper_location | User Manual - Unwanted Radiation | CAUTIONCause: Remaining in hazardous area.Possible consequences: Danger of inadvertent radiation!Remedy: During radiation release, make sure you are in a location subject to as little radiation as possible. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User | S1M | VB10 | [12 empty] | NotDetermined
1246 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562560 | Headline | Cause |  | Accidental release of radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
1247 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562561 | Requirement | Measure | measure_accidental_activation_X-ray | Design of X-ray release controls | The X-ray release controls shall be designed to prevent accidental activation.For the standard footswitch it is acceptable that putting the switch upside down does not activate a function. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 01.08.2022 15:49 | Ingo Hollenborg |  |  |  | HICSwitch,HICHandControl |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1248 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562562 | Requirement | Measure | DisplayRadiationOn_s1 | Display Radiation On | A light signal interface will be provided by the generator when X-ray is on. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.04.2022 15:23 | Rudolf Siebenhaar |  |  |  | TrolleyU,XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1249 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562564 | Headline | Cause |  | Malfunction of exposure release circuit. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1250 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562565 | Requirement | Measure | Add_HW_Release_s1 | Hardware signal | In addition to SW a separate hardware signal shall be checked to be set ""inactive"" during start-up of the x-ray generator. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1251 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580167 | Requirement | Measure | safe_activation_X-ray | Quasi-continuous activation of radiation with redundant activation paths | Radiation shall be possible only when two redundant signals provided by two independent paths are present within certain time constraints (quasi-continuous activation), preventing unintended X-ray activation.When only one of two signals is present, radiation shall be stopped and inhibited for the respective control.?Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 07.06.2021 | 17.06.2022 09:08 | Johannes Schmidt |  |  |  | HICSwitch,MCU,XCU,XGen,HICHandControl | SFSIEC 60601-1; 9.2.3.1 Unintended movementIEC 60601-2-54; 201.9.2.3.1 Unintended movementIEC 60601-1; 9.2.2.5 Continuous activationIEC 60601-2-54; 201.9.2.2.5 Continuous activation |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1252 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562567 | Headline | Measure |  | Reference to generator safety circuit and test at start-up | See HM_S1_hw_F1G1U1M2, HM_S1_sw_F3G2U1M2 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1253 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578773 | Headline | Measure |  | Reference - Emergency stop radiation | <XREF>562135</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 23.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1254 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562568 | Headline | Cause |  | Loss of connection between FD and imaging system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1255 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562569 | Headline | Measure |  | Reference - Imaging system hardware test at startup | <XREF>562637</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | Subsystem Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1256 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562570 | Headline | Cause |  | High radiation dose due to excessive procedure duration and / or heavily inadequate radiation settings without clinical benefi |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz52 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Untrained users | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1257 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562571 | Headline | Measure |  | Reference to Warning - Trained Users | <XREF>561986</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1258 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562572 | Headline | Cause |  | Bad image quality of single image - make single image without having done flluoro before |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz39#Haz40#Haz50#Haz30(HS_121) | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1259 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562573 | Headline | Measure |  | No additional measures required | No additional measures are required. The operators are trained on system behavior and correct operation. The description is included in the operator manual (""advice to Fluoro exposure with > 2fps before acquiring a single image in order to set the exposure parameters and to achieve good image quality). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1260 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562574 | Headline | Cause |  | Failed subtraction / roadmap due to operator error or configuration error (licensing,...) |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1261 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562575 | Headline | Measure |  | No additional measures required | No additional measures are required. The operators are trained on system behavior and correct operation / workflow. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1262 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562576 | Headline | Cause |  | Radiation is not switched off or switched off too late (e. g. pulses) due to HW / SW errors |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1263 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562577 | Headline | Measure |  | Reference to measures for radiation switch off | <XREF>562717</XREF> and HM_S1_hw_F5G1U1M2 (HW radiation timer), HM_S1_hw_F5G1U2M1 (watchdog), HM_S1_hw_F5G1U2M4 (mAs switch off by HW), HM_S1_sw_F5G1U3M1 (CAN monitoring) in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1264 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562578 | Headline | Hazard |  | Wrong interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | misrepresentation of results,insufficient visibility, audibility or tactility,use by unskilled/untrained personnel,slips, lapses and mistakes,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1265 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562579 | Headline | Cause |  | Image orientation mixed up due to image rotation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1266 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562580 | Requirement | Measure | docUser_note_correct_position_of_the_patient | User Manual - Image Orientation | CAUTIONCause: Inverted image orientation.Possible consequences: Risk of an incorrect decision during surgery.Remedy: Prior to releasing radiation, the user must ensure the correct patient position with respect to the X-ray beam. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1267 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562581 | Headline | Cause |  | Image orientation mixed up due to flipped image |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1268 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562582 | Requirement | Measure | docUser_note_label_each_side | User Manual - Label each side | CAUTIONCause: The image orientation may be inverted, e.g. if 2D images are displayed on external viewers.Possible consequences: Risk of an incorrect therapeutic decision.Remedy: The operator must apply due diligence to correctly identify the 2 D image orientation. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 30.05.2022 12:01 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1269 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562583 | Headline | Cause |  | Image orientation mixed up |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27#Haz41 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Data and Image processing | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1270 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562584 | Requirement | Measure | docUser_note_lead_letters | User Manual - Image Orientation | CAUTIONCause: Inverted image orientation.Possible Consequences:´Incorrect intervention decision!Remedy: It is recommended to mark the patient position using lead letters. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1271 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562586 | Headline | Cause |  | Wrong image orientation applied by imaging system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Data and Image processing | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1272 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562587 | Requirement | Measure | FLC_IC_ImageOrientation | Encapsulation for image flip and rotate | Encapsulation and separate software tests of image flip/rotate. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1273 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562588 | Headline | Cause |  | Wrong calibration display |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41 #Haz31 #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Data and Image processing | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1274 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562589 | Requirement | Measure | image_processing_update_calibration | Update calibration | After each calibration the length of the lines in the image is updated. If no calibration has been performed, the length of the lines is displayed in pixels. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1275 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562590 | Requirement | Measure | docUser_note_calibration | User Manual - 2D measurements | CAUTIONCause: Wrong calibration value is entered or length or angle information are not interpreted correctlyPossible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisionsRemedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e. g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1276 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562591 | Headline | Cause |  | Wrong interpretation of results after calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41 #Haz31 #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1277 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562592 | Headline | Measure | docUser_note_calibrated_distance | Reference to Caution for 2D measurements | <XREF>562590</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1278 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562593 | Headline | Cause |  | Wrong mask image for subtraction |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41 #Haz31 #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Untrained users | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
1279 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562594 | Requirement | Measure | FLC_Acquisition_ClearMask | Correct mask calculation | The FLC software triggers the start of the mask calculation phase depending on the system state (subtraction or roadmap). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1280 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562595 | Headline | Cause |  | Movements (either extrinsic or intrinsic) or image manipulations (e.g. AAA line) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz48 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1281 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562596 | Requirement | Measure | docUser_AAA_line_mismatch | User Manual - Mismatch of AAA Line | WARNINGCause: An annotated image with graphic overlay lines is displayed. Following system and/or table or patient movements and/or image modifications, subsequent images may not match the overlay line displayed due to projective changes. Possible Consequences: Incorrect interpretation of the display and resulting incorrect decision during operation! Remedy: The graphic overlay line is displayed until the user actively turns it off.This is done independently of any changes (for example, movements, operating mode changes, image modifications). The user must ensure that the displayed overlay line matches the current, underlying image (for example, display of anatomical structures via the label function and/or use of contrast agent). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1282 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562599 | Headline | Cause |  | Image orientation reset due to application change |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) |  |  |  |  | This behaviour shall be checked for Genesis. It is more reasonable to change the orientation to default only when a new patient is registered, but not to trigger it when changing the application. | #Haz27 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1283 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562600 | Requirement | Measure | docUser_image_orientation_reset | User Manual - Image orientation reset | CAUTIONCause: The image orientation is reset when changing the program.Possible consequences: Risk of an incorrect decision during surgery!Remedy: Be aware that the image orientation is reset when changing the program. All previous settings are lost. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 20.09.2022 14:26 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1284 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562601 | Headline | Cause |  | LIH is confused with live image |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1285 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562602 | Requirement | Measure | FLC_Acquisition_ClearLIH | Clearance of LIH during radiation | When the user starts radiation, FLC clears the LIH image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1286 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562603 | Headline | Cause |  | Live image covered by messages or UI elements |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1287 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562604 | Requirement | Measure | FLC_Acquisition_Live_Image_in_foreground | Keep live imaging in foreground | The displayed live image is not covered by dialogs or other UI elements. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1288 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562605 | Headline | Cause |  | Incorrect K-wire display |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1289 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562606 | Requirement | Measure | docUser_icorrect_k-wire_display | User Manual - Incorrect K-wire display | CAUTIONCause: Time gap between display of K-wire including overlay grafic and real moved K-wire. False positive or false negative detection of K-wires.Possible consequences: Risk of confusion and wrong interventional decisionRemedy: Duty of care for operator, that he/she shall be aware of this possible mismatch. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 29.09.2022 12:07 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1290 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562607 | Headline | Cause |  | overlay hide K-wire on image |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.09.2022 16:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1291 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562608 | Requirement | Measure | docUser_overlay_hide_k-wire | User Manual - overlay hide K-wire on image | CAUTIONCause: Overlay itself may hide K-wire in live X-ray image (depends on thickness of K-wire and overlay configuration).Possible consequences: Risk of wrong interventional decisionRemedy: Target Pointer can be switched-off (toggle ON/OFF) to allow free view on X-ray live image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1292 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562609 | Requirement | Measure | visualize_top_of k-wire | Visibility of top of K-wire | Visibility of the top of K-Wire happens via colored transparent lines overlaying the image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1293 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562610 | Headline | Hazard |  | Loss of images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display |  | S1M | VB10 | [12 empty] | NotDetermined
1294 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562611 | Headline | Cause |  | Insufficient image storage space |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1295 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562612 | Requirement | Measure | NotReady_s1 | Storage capacity exhausted | When the image storage limit is reached, the system shall inhibit further radiography. Note: If the limit is reached during an acquisition, this acquisition will be continued until finished and stored due to a small buffer.?Fluoroscopy shall not be stopped or inhibited, and a message shall be displayed to the operator. LIH functionality shall not be impaired in any case. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 20.09.2022 14:51 | Johannes Schmidt |  |  |  | FLC,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1296 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562613 | Requirement | Measure | DiskCapacity1_s1 | Display storage space | On monitor the actual storage capacity is shown and updated when an image is stored.The image system displays the storage space still available as relative part of the assignable capacity.A message shall be shown to the user when the available storage capacity is equal or below 10%. Note: The accuracy of the displayed storage capacity should be +/- 5% of the used split ration for the icons. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1297 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562614 | Headline | Cause |  | Images / series lost in imaging system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1298 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569132 | Requirement | Measure | FLC_ImageStore | Immediate image storage on hard disk | Acquired images are always stored immediately on hard disk. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | empty | empty | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1299 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578317 | Requirement | Measure | full_storage_rate_radiography | Full storage rate radiography | For radiography, the storage rate shall not be changeable by the customer. Every frame shall be stored. Replay of scenes shall apply the same behaviour as live (fps). Note, the default storage rate shall be set to 100%. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 07.02.2021 | 19.09.2022 14:50 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1300 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562616 | Headline | Cause |  | Reducing the storage rate or being confused about replay |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1301 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562617 | Headline | Measure |  | Reference - Immediate image storage on hard disk. | <XREF>569132</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1302 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562618 | Headline | Hazard |  | Unwanted Injection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,use by unskilled/untrained personnel,slips, lapses and mistakes |  | S1M | VB10 | [12 empty] | NotDetermined
1303 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562619 | Headline | Cause |  | Unwanted radiation release |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1304 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562620 | Headline | Measure |  | No measure required | no measure required due to risk classification | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1305 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562621 | Headline | Cause |  | Triggersignal failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1306 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562622 | Headline | Measure |  | No measure required | no measure required due to risk classification | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1307 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580664 | Headline | Hazard |  | Unintended injection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 17.06.2022 11:20 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
1308 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580665 | Headline | Cause |  | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz31 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. The operator manual advises to check the injector status prior to using it in combination with the system and contains a description how the combination works. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1309 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580666 | Headline | Measure | docUser_injector_readiness | Reference - User Manual - Injector Readiness | <XREF>562625</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 17.06.2022 11:20 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1310 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 581162 | Requirement | Measure | injector_trigger_dedicated_modes | Injector trigger in dedicated modes and conditions only | The system shall trigger the injector only in the following modes and conditions:- Subtraction- Roadmap: Only when the operator has not performed a subtraction before (then automatically a subtraction scene is performed when roadmap is selected), or when no valid peak opacification image is available (e. g. after changing magnification)- Digital Cine Mode Note: That means when a subtraction has been performed before roadmap is selected, then the injector shall not be triggered when roadmap is performed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU,XCU |  |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1311 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 581195 | Requirement | Measure | display_injector_armed | Display when injector is armed | The system shall display when the injector is armed, when the conditions under <XREF>581162</XREF> are fulfilled (that means that upon next X-ray release, the injector is triggered). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1312 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562623 | Headline | Hazard |  | Radiation without injection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Foreseeable misuse |  | S1M | VB10 | [12 empty] | NotDetermined
1313 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562624 | Headline | Cause |  | Acquire image, but injector not prepared correctly |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. The operator manual advises to check the injector status prior to using it in combination with the system. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1314 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562625 | Requirement | Measure | docUser_injector_readiness | User Manual - Injector Readiness | CAUTIONCause: Radiation release without without contrast agent injection or unintended injection.Possible consequences: Risk of unnecessary radiation exposure or unintended contrast agent injection.Remedy: Make sure that the injector is ready to administer contrast agent before releasing radiation. Note that every X-ray mode triggers the injector when it is in ready state. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1315 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 581193 | Headline | Cause |  | Injector interface is damaged due to motorized platform movement with plugged injector connection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. The operator manual advises to check the injector status prior to using it in combination with the system. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1316 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 581194 | Requirement | Measure | restrict_move_range_injector | Restricted platform movement range with plugged injector | When the injector is plugged, motorized movements of the mecanum platform shall be restricted to a distance of < 50 cm. When it is reached, the system shall stop motorized movements and display an appropriate error message. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.07.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [5 empty] | #Haz50 #Haz40 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. The system complies with applicable regulations and technical standards. The operator manual advises to check the injector status prior to using it in combination with the system. This limits the probability of occurrence of this hazardous situation. Additional mitigations are not deemed to further reduce or prevent the risk. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | System Integration Test | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VxFUTURE | [12 empty] | NotDetermined | E | Low | E
1317 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 595189 | Requirement | Measure | um_unplug_injector_during_movement | Unplug injector cable during movement | CAUTIONCause: A strain might be applied to the injector cable when plugged during movements.Possible consequence: The injector cable can get damaged or unplugged during a chassis movement of the system.Remedy: Unplug the injector cable if you perform longer movements that might put tension on the cable. | InWork | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 28.02.2022 | 27.07.2022 15:29 | Johannes Schmidt | [7 empty] | empty | System Integration Test | EXPORT | User Manual | empty | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | System | S1M | VB10 | [12 empty] | NotDetermined
1318 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562626 | Headline | Hazard |  | Radiation interrupted |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1319 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562627 | Headline | Cause |  | Exceeding maximum delay time |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1320 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568834 | Requirement | Measure | image_display_delay_time | Image display delay time | An image display delay time of max. 200 ms is determined to be practicable based on the image processing capabilities of the system and to ensure sufficient reaction time for the operator for all imaging modes except Single Image. Note, the achieved image display delay time will then be below the average human reaction time of ~ 300 ms according to Sigmund Exner. For Single Image, which includes DR (Digital Radiography), DR-SUM (Digital Radiography with several pulses) and 3D Scan, as well as DCM and Subtraction, the image display delay time might be up to 4 s. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 08.06.2022 17:00 | Johannes Schmidt |  |  |  |  | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | EXPORT | Verification&Validation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1321 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568835 | Requirement | Measure | image_display_delay_time_ext | Image display delay time - external | If an external device is utilized to display images acquired with the CIOS system via the NXS interface, the image display delay time shall be evaluated as part of the vendor risk management. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 31.08.2020 | 24.06.2022 09:34 | Johannes Schmidt |  |  |  |  | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1322 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596894 | Requirement | Measure | docUser_image_display_delay | Caution in user manual to image display delay | CAUTION Cause: Radiography modes are misused for real-time imaging on purpose by the operator. Possible consequences: With Radiography modes the image display delay time may be longer than in Radioscopy modes and may exceed limits for accepted delays for real-time imaging. Remedy: For real-time imaging do only select Radioscopy modes such as Fluoro and Road. Do not select Radiography modes (e.g. single image including DR, DR-Sum, 3D Scan, DCM, Subtraction) for real-time imaging. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 08.06.2022 | 24.06.2022 09:34 | Johannes Schmidt |  |  |  |  | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1323 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562629 | Headline | Hazard |  | Restricted imaging functionality due to low-level arcing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1324 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562630 | Headline | Cause |  | Arcing in the X-ray tube |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause081 | #Haz40 #Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1325 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562631 | Headline | Measure |  | No measure required | No additional risk control measures required. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1326 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562632 | Headline | Hazard |  | Restricted imaging functionality due to medium-level arcing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1327 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562633 | Headline | Cause |  | Arcing in the X-ray tube |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause082 | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1328 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562634 | Headline | Measure |  | No measure required | No additional risk control measures required. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1329 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562635 | Headline | Hazard |  | Image quality is reduced |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1330 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562636 | Headline | Cause |  | FLC hardware is defective |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause081 | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1331 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562637 | Requirement | Measure | FLC_PowerUp | Imaging system hardware test at startup | The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state (see <XREF>578746</XREF>). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 22.10.2021 11:45 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1332 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 599844 | Requirement | Measure | um_high_energy_isotopes | User Manual - High-Energy-Isotopes | Caution Cause: Flat detector hardware is exposed to high-energy isotopes at low proximity, e.g. by means of Cobalt-60 radiation sources during BrachytherapyPossible consequences: Flat detector hardware damaged, leading to low image quality or image artifactsRemedy: Keep the detector system and electronics in a safe distance (> 3 m) from high-energy isotopes (> 1 MeV), e.g., during the insertion or removal process of the radiation probe in Brachytherapy. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 30.08.2022 | 29.09.2022 14:35 | Johannes Schmidt | [7 empty] | Basic | empty | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | System | S1M | VB10 | [12 empty] | NotDetermined
1333 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562638 | Headline | Cause |  | Numerical calculation error or parameterization issue in imaging pipeline |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Cause081 | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Availability in critical functions | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1334 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562639 | Requirement | Measure | FLC_IC_ImageQuality_ConsistencyCheck | Consistency check in imaging pipeline | Consistency check in imaging pipeline and notification of host in case of error. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1335 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562640 | Headline | Cause |  | Detector heats up out of calibration range |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1336 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562641 | Headline | Measure | FDTemperature_Switchoff_s1 | Reference - System Warning - Temperature of the Detector | <XREF>562226</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FDmediumH,FDmediumC,FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1337 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562642 | Headline | Measure |  | Reference to Caution in Operator Manual - Detector temperature | <XREF>563301</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FDmedium,FDmediumV,FDsmall,FDsmallV,FDmediumH,FDsmallH,FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1338 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562643 | Headline | Function |  | Patient browser |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1339 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562644 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1340 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562645 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1341 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562646 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1342 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562647 | Headline | Function |  | Patient registration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Data and Image processing,Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
1343 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562648 | Headline | Hazard |  | Misinterpretation of patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
1344 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562649 | Headline | Cause |  | Patient-ID already in use |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz26#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1345 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562650 | Requirement | Measure | docUser_selectPatient | User Manual - Patient ID | CAUTIONCause: Patient ID already exists.Possible Consequences: Misinterpretation of data of the examined patient if the selection is based solely on the patient ID.Remedy:-Always use four-way identification [patient name, sex, date of birth, patient ID].- Avoid mixing up patient data by using four-way identification. Selecting patient data based on the patient ID alone is insufficient. Ensure the proper display of parameters in the Worklist and in the Patient list. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 27.07.2022 15:49 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1346 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562651 | Headline | Hazard |  | Emergency login not possible |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
1347 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562652 | Headline | Cause |  | Authenticated login, access not possible |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator has the possibility to operate the system in emergency situations without a password. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1348 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562653 | Requirement | Measure | FLC_EmergencyLogin | Emergency login without password protection | Emergency user login is possible without a password.Note: Due to PSS requirements, disk encryption is configurable: It is an option and must be activated by service on request. Consequences and boundary conditions are described in operator manual, admin guide and MDS2 statement. With these conditions, it is acceptable to restrict the effectivity of above mentioned measure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1349 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562654 | Headline | Function |  | APS management |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1350 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562655 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1351 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562656 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1352 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562657 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1353 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562658 | Headline | Function |  | Administration, selection and changing of applications |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1354 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562659 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1355 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562660 | Headline | Cause |  | Communication problems |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | EXPORT | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Lack of integrity of data | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1356 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562661 | Requirement | Measure | CommunicationError | Reliable data transfer standards and protocols | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 14:33 | Johannes Schmidt |  |  |  | XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1357 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562662 | Requirement | Measure | NGS_speed_detection | NGS protocol with speed detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a link speed detection. Its also valid when NGS is used in the context of OpenApps.Incase the detected speed is below the threshold limit, the user shall be provided with an audible or visual indication. | Approved | Unassigned | Sonali Rath | Sonali | Rath | Judith Welzel |  |  | 29.07.2020 | 26.01.2022 14:51 | Sonali Rath |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1358 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562663 | Requirement | Measure | NGS_link_loss_detection | NGS protocol with link loss detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. Its also valid for OpenApps.If the link loss is detected, it shall be indicated to the user and an automatic release of mouse events shall occur. | Approved | Unassigned | Sonali Rath | Sonali | Rath | Judith Welzel |  |  | 29.07.2020 | 26.01.2022 14:44 | Sonali Rath |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1359 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562664 | Requirement | Measure | NGS_latency_detection | NGS protocol with latency detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a latency detection. This means measuring the time between sending a frame till receiving this frame and in case of exceeding a configurable limit to indicate it with an audible or visual indicator. VxFUTURE: when using NGS in the OpenApps context, the latency detection shall be configurable specific to the app (i.e. either deactivated or threshold value adjusted based on the Risk Managmenent of the individual partner apps). | Approved | Unassigned | Sonali Rath | Sonali | Rath | Judith Welzel |  |  | 29.07.2020 | 13.06.2022 10:07 | Sonali Rath |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1360 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562665 | Requirement | Measure | NGS_update_rate_detection | NGS protocol with update rate detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a update rate detection between two frames. In case of exceeding a configurable time an audible or visual indication shall be done. VxFUTURE: This can be configured individually for partner apps (either deactivated or a threshold value configured based on the respective partner app's Risk Manamagement proposal) in case of OpenApps. | Approved | Unassigned | Sonali Rath | Sonali | Rath | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 14:56 | Johannes Schmidt |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1361 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562666 | Requirement | Measure | NGS_payload_integrity | NGS protocol with data integrity | For transfer of frames via NGS the data integrity of these frames shall be considered by using a check sum algorithm (e.g CRC32). In case of an incorrect checksum is detected an audible or visual indication shall be done. When NGS is used in context of OpenApps, the temporal integrity of an image series that builds on one another must also be monitored using a checksum. Incase of a missing image or incorrect CRC, an audible or visual indication shall be available. | Approved | Unassigned | Sonali Rath | Sonali | Rath | Judith Welzel |  |  | 29.07.2020 | 17.01.2022 12:16 | Sonali Rath |  |  |  | HICTouchCon |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1362 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569295 | Headline | Measure | image_display_delay_time | Reference - Image display delay time | <XREF>568834</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FLC | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1363 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569296 | Headline | Measure | image_display_delay_time_ext | Reference - Image display delay time - external | <XREF>568835</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 15.09.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC | IEC 60601-2-43; cl. 201.12.4.102 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1364 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569238 | Requirement | Measure | um_image_quality_streaming | Caution - Reduced image quality streaming | CAUTION?Cause: When streaming via external systems, image quality (e.g., resolution, frame rate or flickering) might be affected by connector type or external monitor hardware and settings.?Possible Consequences: Reduced image qualityRemedy: In order to avoid loss of image quality ensure that the connection cable specifications and display parameters of the connected external monitors fulfill the required specifications as described in the chapter Technical Specifications ------ Not part of Caution: Note, the required specifications for the external monitor are medical grade, 4K, color display, 3840x2160 pixels, 60Hz, and 4:4:4 Chroma Subsampling. The cable requirements are a short HDMI2.0 (or better) video cable (<3m) between CIOS Move and external monitor or fiber optic cable. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 14.09.2020 | 16.09.2022 15:05 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1365 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562667 | Headline | Cause |  | Lossy organprogram |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1366 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562668 | Headline | Measure |  | Reference - User Manual - Wrong exposure parameters | <XREF>562722</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | ReadyForTCReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1367 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562669 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1368 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562670 | Headline | Cause |  | Communication network disturbance -incorrect parameters |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1369 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562671 | Headline | Measure |  | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1370 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562672 | Headline | Cause |  | Operating error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1371 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562673 | Headline | Measure |  | Reference - User Manual - Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1372 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 568855 | Requirement | Measure | Indication_of_loading_factors | Indication of loading factors | The system indicates the loading factors as per standard requirements, see 06S456591, and following, 06S434271, and following in <XREF>561879</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 01.09.2020 | 21.06.2022 14:59 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1373 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562674 | Headline | Hazard |  | Dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1374 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562675 | Headline | Cause |  | Incorrect parameters due to communication/network disturbance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1375 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562676 | Headline | Measure |  | Reference - Ready for radiation | <XREF>578746</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1376 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562677 | Headline | Measure |  | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1377 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562678 | Headline | Function |  | Display of patient relevant data on user interface |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1378 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562679 | Headline | Hazard |  | Mismatch of patient data, incorrect treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Incorrect or inappropriate output,data acquisition, evaluation and display,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
1379 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562680 | Headline | Cause |  | Disturbance in processing, incorrect patient data displayed |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1380 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562681 | Requirement | Measure | patientdatadisplay_s1 | Patient Data Display | The patient name and an additional configurable information entity (either patient ID or request ID or accession number or date of birth) of the selected study shall be displayed. Note: In case of limited number of characters the additional information shall be shown at least partially so that different patients with same names can be distinguished. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1381 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562682 | Headline | Measure |  | Reference to measures for patient or image data refresh and data consistency | <XREF>563323</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1382 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569126 | Requirement | Measure | FLC_Correction_display | Refresh display after patient or image data correction | When patient or image data has been corrected, the imaging system refreshes the display. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1383 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569127 | Requirement | Measure | FLC_DataConsistency | Imaging system data consistency | When the image is processed for transfer, a consistency-check is done between the study UID in the job list and the study UID in the image header. For data exchange between FLC and RIS the DICOM3 standard is used. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1384 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569128 | Requirement | Measure | FLC_SingleStudy | Transfer images of current patient only | Only images of the selected patient can be transferred to the hardcopy camera | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 03.09.2020 | 19.05.2022 13:47 | Klaus-Peter Fichtner (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1385 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562683 | Headline | Function |  | Process patient exam list |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1386 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562684 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1387 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562685 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1388 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562686 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1389 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562687 | Headline | Function |  | Presetting of collimator position |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1390 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562688 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1391 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562689 | Headline | Cause |  | Communication network disturbance -incorrect parameters |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1392 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562690 | Headline | Measure |  | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1393 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562691 | Headline | Cause |  | Use Presettings of Collimator Position for Automatic Virtual Shutter |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1394 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562692 | Headline | Measure |  | No measure required | Note:Risk rating is in accordance to <XREF>562695</XREF> as this feature will effect measure <XREF>562696</XREF>. Nonetheless, no measure is required, as this feature can be switched off easily (via toggle button).Risk assessment was necessary due to feature "Automatic Virtual Shutter" (a.k.a. "set black", 04S407692 in <XREF>561854</XREF>). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1395 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562693 | Headline | Function |  | Exposure field |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1396 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562694 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1397 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562695 | Headline | Cause |  | Position of collimator leaves out of control due to stepping errors |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
1398 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562696 | Requirement | Measure | search_run | Perform Calibration of Collimator Leaves | Accuracy of collimator leaves position shall remain unchanged during examination. It is possible to perform a search run to calibrate position of collimator leaves. In case a search run will be implemented, it has to be ensured, that no search run will be triggered and/or performed during acquisition (i.e. while releasing X-ray). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | ColS,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1399 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562697 | Headline | Function |  | Wireless Remote control for movement functions |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1400 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562698 | Headline | Hazard |  | Injury by crushing | Collision with operator, patient or equipment | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1401 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562707 | Headline | Cause |  | Other wireless devices act as remote control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 12.10.2022 11:49 | Johannes Schmidt | [5 empty] | #Haz07 | Risk control measures are in place and reduce the risk as far as possible without adversely affecting the benefit-risk-ratio. Compliance to applicable regulations, technical standards and other design requirements limits the probability of occurrence of the scenario. Service activities ensure to maintain the integrity of the design and detect degradations. The operator manual warns to take care of potential crushing zones and requires the appropriate carefullness when releasing system movements, especially when additional parts are mounted at the C-arm, e. g. for navigation. Additional mitigations are not deemed to further reduce or prevent the risk of collision. Post market surveillance and complaint handling are in place to identify specific issues. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1402 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562708 | Requirement | Measure | Remote_Handheld_interference | Remote control and system are paired | Remote control and system are paired. No other wireless devices sending on the same frequency can interfere. | Deferred | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 20.09.2022 12:05 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | ReadyForTCReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
1403 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562709 | Headline | Hazard |  | Interfere with other devices | Interference with life supporting device and possible malfunction. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Electro-magnetic influence (EMC) |  | S1M | VB10 | [12 empty] | NotDetermined
1404 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562710 | Headline | Cause |  | Electromagnetic field disturbs life-supporting devices |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz45 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1405 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562711 | Headline | Measure |  | Reference to Operator Manual - Interference with life-supporting devices or critical equipment | <XREF>562061</XREF>. | Deferred | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | empty | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1406 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562712 | Headline | Measure |  | Reference - EMC compliance - emissions, immunity and pass / fail criteria | <XREF>562015</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1407 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562713 | Headline | Function |  | Dose control |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1408 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562714 | Headline | Hazard |  | Unnecessary radiation, dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1409 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562715 | Headline | Cause |  | Dose controlling |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1410 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562716 | Requirement | Measure | serv_DoseAdjustment | IQAP measurement | IQAP measurement after each dose adjustment. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1411 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562717 | Requirement | Measure | Time_Limit_s1 | Switch off radiation | Switch off radiation after max time. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1412 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562718 | Headline | Cause |  | Incorrect parameters |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1413 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562719 | Requirement | Measure | docUser_system_calibration | Admin Guide - System calibration | CAUTIONCause: Incorrect calibration parameters degrade image quality.Possible consequences: Risk of unnecessary radiation exposure due to incorrect calibration parameters!Remedy: Calibrate the system after 24 months to preserve image quality. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 19.07.2022 15:31 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Admin Guide | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1414 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569225 | Headline | Measure | radiation_ready | Reference - Ready for radiation | <XREF>578746</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 11.09.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1415 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562721 | Requirement | Measure | FLC_SystemHandling | Display clinical procedure | The selected clinical procedure shall be shown on the monitor during examination. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1416 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562722 | Requirement | Measure | um_Exposure_Parameters | User Manual - Wrong exposure parameters | CAUTIONCause: Incorrect exposure parametersPossible consequences: Risk of unnecessary X-ray exposure due to incorrect exposure parameters.Remedy: Ensure the exposure parameters are correct. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1417 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562723 | Headline | Measure |  | User Manual - Reduced Image Quality due to Filtering | <XREF>562727</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1418 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562724 | Requirement | Measure | FLC_IC_BrightnessValueCalculation | Encapsulation of brightness value calculation | Encapsulation and separate software tests for brightness value calculation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1419 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562725 | Headline | Cause |  | Reduced image quality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1420 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562726 | Headline | Measure |  | User Manual - System calibration | <XREF>562719</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1421 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562727 | Requirement | Measure | um_Processing_Parameters | User Manual - Reduced Image Quality due to Filtering | CAUTIONCause: Processing images by changing the image display. Possible consequences: Loss of image quality by artifacts. Remedy: Use care when changing the image display settings. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.09.2022 12:36 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1422 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562728 | Headline | Cause |  | Incorrect parameters due to error in generator or tube |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1423 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562729 | Headline | Measure |  | Reference to measures to generator self-tests during start-up / initialisation and filtration test | See HM_S1_sw_F5G2U1M1 (RAM test at start-up), HM_S1_sw_F5G2U1M2 (integrity check program code),HM_S1_sw_F5G2U1M3 (integrity check of EEPROM at start-up), HM_S1_sw_F5G2U2M1 (self-test of measuring circuit kV / mA at initialisation), HM_S1_hw_F5G2U2M2 (kV min / max monitoring), HM_S1_manu_F5G2U3M1 (HVL test) in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1424 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578746 | Requirement | Measure | radiation_ready | Ready for radiation | If a relevant component is not ready for radiation, the system shall block any radiation release request. Note:Relevant components are imaging system and X-Ray generator. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 22.02.2021 | 23.09.2022 13:50 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1425 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562731 | Headline | Measure |  | Reference to measure to ensure changing application is possible by operator | <XREF>562721</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1426 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562732 | Headline | Measure | um_Exposure_Parameters | Reference - User Manual - Wrong exposure parameters | <XREF>562722</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1427 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562733 | Headline | Measure |  | Reference - User Manual - Reduced Image Quality due to Filtering | <XREF>562727</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1428 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562734 | Headline | Measure | FLC_IC_BrightnessValueCalculation | Reference - Encapsulation of brightness value calculation | <XREF>562724</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1429 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562735 | Headline | Function |  | Grid selection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1430 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562736 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1431 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562737 | Headline | Cause |  | Operating error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. The operator manual advises to pick the correct imaging parameters and settings and to perform regular checks. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1432 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562738 | Requirement | Measure | um_Removable_Grid | User Manual - Incorrect Grid Status | CAUTIONCause: Incorrect grid status: inserted, removed.Possible consequences: X-ray exposure without clinical benefit.Remedy:- Select the application according to your clinical procedure.- Make sure that the current grid status meets your image requirements. The imaging system indicates whether a grid has been inserted or removed.- If necessary, insert or remove the grid manually. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1433 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562739 | Headline | Hazard |  | Crushing between the housings of the detector case. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1434 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562740 | Headline | Cause |  | Changing the grid |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz08 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1435 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562741 | Requirement | Measure | um_ChangingGridWithCare | User Manual - Changing the grid | CAUTIONCause: Changing the grid.Possible consequences: Crushing hazard between the detector housing parts.Remedy: Note the hazardous zones/locations between the detector housing parts. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1436 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562742 | Headline | Function |  | Store fluoro |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1437 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562743 | Headline | Hazard |  | Loss of images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
1438 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562744 | Headline | Cause |  | Insufficient image storage space |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz40 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1439 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562745 | Headline | Measure |  | Reference - Display storage space | <XREF>562613</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1440 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562746 | Headline | Function |  | Position FOV |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1441 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562747 | Headline | Hazard |  | Dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1442 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562748 | Headline | Cause |  | Incorrect collimator setting |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. The operator manual advises to pick the correct imaging parameters and settings and to perform regular checks. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1443 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562749 | Requirement | Measure | manu_CollimatorAdjust_s1 | Adjustment of the collimator plate | Adjustment of the collimator plate is designed so that the edge is visible in the fluoro image. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1444 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562750 | Requirement | Measure | um_CollimatorCheck | User Manual - Collimator Check | CAUTIONCause: Wrong collimator setting.Possible Consequences: Dose not as required!Remedy: The collimators need to be set so that at least one edge of the leaves is visible in the fluoroscopy image. This must be checked daily. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1445 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562751 | Requirement | Measure | serv_CollimatorAdjust_s1 | Adjustment of the collimator plate | Regular check of the collimator plate adjustment per the Maintenance Plan by prescribed check. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1446 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562752 | Headline | Measure |  | Reference - Perform calibration of collimator leaves | <XREF>562696</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1447 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562753 | Headline | Hazard |  | Eye injury (visual disturbance) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1448 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562754 | Headline | Cause |  | Unintended look into the laser beam for a lengthy time |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz34 | See general justifications in the introduction which apply here. The operator manual advises to take care of the laser light. No additional practical mitigations are available to further reduce or prevent eye exposure to the laser beam. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1449 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562755 | Headline | Measure |  | Usage of laser class 1 | <XREF>562338</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1450 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562756 | Headline | Function |  | Zoom format change |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1451 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562757 | Headline | Hazard |  | Dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1452 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562758 | Headline | Cause |  | Error in FD zoom change |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1453 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562759 | Headline | Measure |  | Adjustment of the collimator plate | <XREF>562749</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1454 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562760 | Headline | Function |  | Exchange grid |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1455 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562761 | Headline | Hazard |  | Dose not as desired |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1456 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562762 | Headline | Cause |  | Incorrect grid status, inserted, removed |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. The operator manual advises to pick the correct imaging parameters and settings and to perform regular checks. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
1457 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562763 | Requirement | Measure | Removable_Grid_s1 | Detect grid status | It has to be detected at the system if grid is inserted or not. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC,MainU,XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1458 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562764 | Headline | Measure |  | User Manual - Incorrect Grid Status | <XREF>562738</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1459 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562765 | Headline | Function |  | Release patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1460 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562766 | Headline | Hazard |  | Infection: A patient is infected by virus/bacteria due to bad disinfection of the system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection |  | S1M | VB10 | [12 empty] | NotDetermined
1461 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562767 | Headline | Cause |  | Insufficient cleaning of system between patients |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:38 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz29 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1462 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562768 | Requirement | Measure | OpeningsInCovers_s1 | Reducing openings and gaps | Covers are designed that openings and gaps are reduced. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 10.03.2022 13:39 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1463 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562769 | Headline | Measure |  | User Manual - Cleaning of Wheels | <XREF>563241</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1464 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562770 | Headline | Hazard |  | Delayed or no treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
1465 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562771 | Headline | Cause |  | Blood or other liquids damage system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33#Haz42 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1466 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562772 | Headline | Measure |  | Reference - Spillage of Liquids on System | <XREF>562238</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1467 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562773 | Headline | Cause |  | Cleaning and disinfection substances damage system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1468 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562774 | Headline | Measure |  | User Manual - Insufficient cleaning | <XREF>562240</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1469 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562775 | Headline | Hazard |  | Undesired radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1470 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562776 | Headline | Cause |  | Cleaning person accidentally steps on footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1471 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562777 | Requirement | Measure | docUser_system_switch_off | User Manual - System switch-off | CAUTIONCause: Inadvertent actuation of the foot pedal.Possible consequences: Risk of radiationRemedy: Switch off system before cleaning. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1472 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562778 | Headline | Measure |  | EmergencyStop | see <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:23 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1473 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562779 | Headline | Function |  | Select patient from patient list |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1474 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562780 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1475 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562781 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1476 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562782 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1477 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562783 | Headline | Function |  | Image display |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1478 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562784 | Headline | Hazard |  | Incorrect treatment basis |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
1479 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562785 | Headline | Cause |  | Confusing of, mix-up of images/series from different patients. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz26#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1480 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562786 | Requirement | Measure | StoreTogetherHcopyMon | Always Display Patient Data | The system shall show patient name during examination. Hint:Please note, this measure neither includes exported images nor print-outs (hardcopy) as hazard and cause is related to time period of examination in accordance to intended use (IFU). Anonymizing of image data sets is not in scope of this hazard, either. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1481 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562787 | Requirement | Measure | FLC_Images_for_registered_patient | Correct image assignment | FLC assigns acquired 2D images to the patient who was registered at the time the images have been acquired. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1482 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562788 | Headline | Cause |  | Images are sorted or numbered incorrectly |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz26#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1483 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562789 | Requirement | Measure | FLC_StoreImagenumber | Image number storage in header | Storage of the image number in the image header of the 2D images. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1484 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562790 | Headline | Function |  | Close Patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1485 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562791 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1486 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562792 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1487 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562793 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1488 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562794 | Headline | Function |  | Disconnect Trolley and Main Unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1489 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562795 | Headline | Hazard |  | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1490 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562796 | Headline | Cause |  | Erroneous release of stored energy (ESU) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | A
1491 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562797 | Requirement | Measure | electrical_separation_access_connections | Electrical separation of accessible parts not protected by MOOP | The system shall safely separate the trolley-chassis cable pins (X10) and the X-ray Generation subsystem as soon as the trolley and the main unit are disconnected. Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1492 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562798 | Requirement | Measure | check_separation_gen_manufacturing | Check of electrical separation functionality during manufacturing - generator | The functionality of the measure implemented with <XREF>562797</XREF> shall be tested with 100% coverage during manufacturing. Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:26 | Johannes Schmidt |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1493 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562799 | Requirement | Measure | serv_check_separation | Check of electrical separation functionality during regular maintenance | Service Documentation:WARNINGCause: Undetected error in electrical separation of X10 pins.Possible consequences: Electrical shock when trolley-chassis cable is disconnected and X10 pins are touched.Remedy: Perform a check of the contactors according to description during maintenance. Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1494 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 573393 | Headline | Hazard |  | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.11.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1495 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 573394 | Headline | Cause |  | Erroneous release of stored energy (ELKO buffer) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.11.2020 | 26.09.2022 16:58 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz64 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | B | Med | B
1496 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 573395 | Headline | Measure | electrical_separation_access_connections | Reference - Electrical separation of accessible parts not protected by MOOP | <XREF>562797</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.11.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1497 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 573396 | Headline | Measure | check_separation_gen_manufacturing | Reference - Check of electrical separation functionality during manufacturing - generator | <XREF>562798</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.11.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1498 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 573397 | Headline | Measure | serv_check_separation | Reference - Check of electrical separation functionality during regular maintenance | <XREF>562799</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 16.11.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1499 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562800 | Headline | Function |  | Quick Clean |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1500 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562801 | Headline | Hazard |  | Contamination through improper cleaning |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Infection,Other agents,Pollution |  | S1M | VB10 | [12 empty] | NotDetermined
1501 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562802 | Headline | Cause |  | Usage of improper cleaning agents |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz29 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,Environment | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1502 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562803 | Requirement | Measure | docUser_note_appropriate_detergents_and_germicide | User Manual - Improper Cleaning Agents | CAUTIONCause: Use of unsuitable cleaning agents.Possible consequences: Risk of contamination.Remedy: Only use the recommended cleaning agents and disinfectants. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1503 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562804 | Headline | Function |  | Retrieve and modify patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1504 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562805 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1505 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562806 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1506 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562807 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1507 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562808 | Headline | Function |  | Image delete |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1508 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562809 | Headline | Hazard |  | Loss of image data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
1509 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562810 | Headline | Cause |  | Unintentional deletion of images/series |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz22 #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1510 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562811 | Headline | Measure |  | Reference - Prompt by local deleting of data | <XREF>562324</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1511 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562812 | Headline | Measure |  | Security dialog for any deletion by the user | <XREF>562818</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1512 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562813 | Headline | Cause |  | Destruction of the patient directory and loss of patient data and images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz22 #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1513 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562814 | Requirement | Measure | FLC_FileSystem | File system of imaging system | Usage of a standard file system (NTFS). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1514 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569130 | Requirement | Measure | FLC_Recovery | Imaging system recovery | Availability of a recovery procedure to reconstruct the patient directory via the study instance UID which is unique. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1515 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562815 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss of data,Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1516 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562816 | Headline | Cause |  | Unintended local deleting of patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz22 #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
1517 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562817 | Headline | Measure |  | Reference - Prompt by local deleting of data | <XREF>562324</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1518 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562818 | Requirement | Measure | SecurityDialog | Security dialog for any deletion by the user | To reduce probability of unintentional loss of data, the deletion of images/series is performed by a security dialog. If non archived images/series are selected for deletion the user has to confirm this in addition. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1519 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562819 | Requirement | Measure | FLC_ArchiveFlag | Storage commitment archive flag | Storage commitment: The archive flag ('A') is set only after a storage commitment has been received.Note: This is valid in case the DICOM node is capable of sending a storage commitment. The flag is set on study and image level . | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1520 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562820 | Headline | Function |  | Calibrate detector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1521 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562821 | Headline | Hazard |  | Image artifacts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1522 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562822 | Headline | Cause |  | Bad calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1523 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562823 | Headline | Measure |  | Automatic procedure to perform calibration correctly | <XREF>562177</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1524 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562824 | Headline | Measure |  | Reference - User Manual - Calibration Procedure | <XREF>562179</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1525 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562825 | Headline | Measure |  | Reference - User Manual - System calibration | <XREF>562719</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1526 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562826 | Headline | Hazard |  | Unwanted radiation exposure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1527 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562827 | Headline | Cause |  | Patient or other person in exam room during calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1528 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562828 | Headline | Measure |  | Reference - Service training | <XREF>562186</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1529 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562829 | Headline | Function |  | Stop button |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1530 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562830 | Headline | Hazard |  | Unit stops too late in a critical situation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
1531 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562831 | Headline | Cause |  | In a panic situation, the user does not find the emergency stop switch quickly enough |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1532 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562832 | Headline | Measure |  | User Manual - Emergency Stop Location | <XREF>562284</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1533 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562833 | Headline | Function |  | Native language support |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1534 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562834 | Headline | Hazard |  | User may be confused or distracted |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | user interface |  | S1M | VB10 | [12 empty] | NotDetermined
1535 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562835 | Headline | Cause |  | Translation incorrect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 | See general justifications in the introduction which apply here. Translations are performed by qualified personnel only. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1536 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562836 | Headline | Measure |  | Reference - Seperate testcase for each foreign user language | <XREF>562838</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1537 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562837 | Headline | Cause |  | Error in translation or implementation of foreign languages |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 | See general justifications in the introduction which apply here. Translations are performed by qualified personnel only. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1538 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562838 | Requirement | Measure | NLS_language_test | Separate testcase for each foreign user language | Define a separate testcase and perform test of UI (text strings and symbols) for each foreign user language. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Ease of Use | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1539 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562839 | Headline | Function |  | Image processing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1540 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562840 | Headline | Hazard |  | Incorrect basis for interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Incorrect or inappropriate output,Data and Image processing,user interface,Inadequate user information |  | S1M | VB10 | [12 empty] | NotDetermined
1541 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562841 | Headline | Cause |  | Wrong configured system language |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz22 #Haz40 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-D | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | F
1542 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562842 | Requirement | Measure | hm_DateNumberFormats | Service: Check date and number formats | Make sure that the date and number formats as well as distance units (inch, cm/mm) are correctly configured for the customer. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 15.06.2022 10:51 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1543 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562843 | Headline | Function |  | Sorting images and modifying patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1544 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562844 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1545 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562845 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1546 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562846 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1547 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562847 | Headline | Function |  | Reporting exam data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1548 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562848 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1549 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562849 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1550 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562850 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1551 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562851 | Headline | Function |  | Transfer of image data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1552 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562852 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
1553 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562853 | Headline | Cause |  | Network disturbance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz50#Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-D | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1554 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562854 | Headline | Measure |  | Reference - Use of reliable data transfer standards | <XREF>562323</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1555 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562855 | Headline | Measure |  | Reference - Prompt by local deleting of data | <XREF>562324</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1556 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562856 | Headline | Function |  | Virus Protection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1557 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562857 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Lack of integrity of data,Loss of data,unknown aspects or quality of third-party software,lack of data sec, particularly vulnera. to tampering,Malicious software,Input from AT SOP 418 [11] |  | S1M | VB10 | [12 empty] | NotDetermined
1558 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562858 | Headline | Cause |  | System gehts infected by malware |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1559 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562865 | Requirement | Measure | ProvideIntrusionDetection_s1_FLC | Intrusion detection solution | An intrusion detection solution based on whitelisting is provided by the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1560 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562866 | Headline | Measure | FLC_ProvideHotfixes | Reference - Provide SVPs regularly | SVPs are regularly provided according to SOP 409 <XREF>578747</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1561 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562867 | Headline | Function |  | Help functionality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1562 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562868 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1563 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562869 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1564 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562870 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1565 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562871 | Headline | Function |  | Measure Area Dose Product |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1566 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562872 | Headline | Hazard |  | Unnecessary dose |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1567 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562873 | Headline | Cause |  | Wrong interpretation of applied dose, dose reporting affected |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz21 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1568 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562874 | Requirement | Measure | serv_Check_DMD_s1 | Maintenance of the dose area product meter | Request for maintenance of the dose area product meter on a regular base. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1569 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562875 | Headline | Function |  | IHE Profiles |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1570 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562876 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1571 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562877 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1572 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562878 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1573 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562879 | Headline | Function |  | Data Model |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1574 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562880 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1575 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562881 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1576 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562882 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1577 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562883 | Headline | Function |  | Worklist |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1578 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562884 | Headline | Hazard |  | Patient data mixed up |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1579 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562885 | Headline | Cause |  | Mixing of data caused by RIS data transfer |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1580 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562886 | Headline | Measure |  | Reference - Imaging system data consistency | <XREF>569127</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1581 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 569129 | Requirement | Measure | FLC_CheckWorklist | Check worklist data | Plausibility and consistency check of worklist data. Worklistdata are marked with a flag. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 03.09.2020 | 14.08.2021 11:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1582 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562887 | Headline | Function |  | Storage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1583 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562888 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1584 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562889 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1585 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562890 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1586 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562891 | Headline | Function |  | Archive to offline media |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1587 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562892 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1588 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562893 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1589 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562894 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1590 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562895 | Headline | Function |  | Query/retrieve |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1591 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562896 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1592 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562897 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1593 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562898 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1594 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562899 | Headline | Function |  | MPPS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1595 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562900 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1596 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562901 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1597 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562902 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1598 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562903 | Headline | Function |  | DICOM dose reporting |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1599 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562904 | Headline | Hazard |  | Wrong dose reported by MPPS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1600 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562905 | Headline | Cause |  | Misleading Information |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz21 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1601 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562906 | Headline | Measure |  | No measure required | Note:This has no impact on interventional decision, but inappropriate actions may be taken based on the dose reading, impact on subsequent radiation events (underreporting...). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1602 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562907 | Headline | Function |  | Standard display function |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1603 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562908 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1604 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562909 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1605 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562910 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1606 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562911 | Headline | Function |  | Verification |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1607 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562912 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1608 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562913 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1609 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562914 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1610 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562915 | Headline | Function |  | Help functionality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1611 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562916 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1612 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562917 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1613 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562918 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1614 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562919 | Headline | Function |  | General requirements |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1615 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562920 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1616 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562921 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1617 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562922 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1618 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562923 | Headline | Function |  | 3D Scan and acquisistion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1619 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562924 | Headline | Hazard |  | Functional Hazards 3D Image Quality is reduced |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1620 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562925 | Headline | Cause |  | Loss of 2D images throughout the 3D scan |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to avoid patient movement during a 3D scan, to lock relevant brakes to avoid unintended system movement during a 3D scan, and to use the whole available orbital range for the scan. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1621 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562926 | Requirement | Measure | save_all_2D_images_data_set_s1 | Save all 2D images data set | Save for all 2D images the corresponding data set of acquisition angles | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1622 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562927 | Requirement | Measure | docUser_reduced_accuracy_of_3D_reconstruction | User Manual - Reduced accuracy of 3D reconstruction | CAUTIONCause: Insufficient geometric accuracy of 3D reconstruction due to an incomplete and/or angulated 3D scan and/or 3D scan started away from 0-positions.Possible consequences: Wrong interventional decision or treatment!Remedy:- Make sure that a complete 3D scan of the entire orbital range is performed.Note, for partial scans it is only required to ensure that the scan is performed within the selected start- and endposition.- Make sure that the C-arm is in 0-position for the angulation axis. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 15:23 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1623 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562928 | Headline | Cause |  | C-arm cable influences c-arm geometry | External C-arm cable influences the movement of C-arm during 3D scan. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
1624 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562929 | Requirement | Measure | cable_consistent_push_and_pull_force_s1 | No external C-arm cable | To reduce mechanical and movement influence for a 3D scan there shall be no external C-arm cable. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.03.2022 14:46 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1625 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562930 | Headline | Hazard |  | Patient risk for biological hazard |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1626 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562931 | Headline | Cause |  | C-arm cable moves into sterile area |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz10 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-D | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | F | Low | D
1627 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562932 | Headline | Measure |  | Reference - C-arm cable avoids sterile area for all positions | <XREF>562929</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:43 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1628 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562933 | Headline | Hazard |  | Risk of injury during motorized orbital movement | Risk of injury during motorized orbital movement of 190° | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1629 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562934 | Headline | Cause |  | Collision of C-arm with people and goods |  | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | F | Low | E
1630 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562935 | Requirement | Measure | docUser_3D_scan_collision_people_and_goods | Reference to workflow | CAUTIONCause: Change of c-arm position after manual test drivePossible consequences: Risk of crushing!Remedy: After the manual test drive the position of the C-arm must not be changed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.07.2022 16:36 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1631 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562936 | Requirement | Measure | motorized_3Dscan_only_after_test_s1 | Motorized operation only after manual test drive | Release for motorized 3D scan after manual test drive | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1632 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562937 | Headline | Measure |  | Reference - Continuous activation | <XREF>562377</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1633 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562938 | Headline | Measure |  | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1634 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562939 | Headline | Measure |  | Reference to set max. speed 3D scan | See <XREF>563486</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:37 | Johannes Schmidt |  |  |  | MCU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1635 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562940 | Headline | Hazard |  | Patient cannot be salvaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1636 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562941 | Headline | Cause |  | Orbital drive is failing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz46 | See general justifications in the introduction which apply here. No additional practical mitigations are available to further reduce or prevent the risk of loss of movement functionality. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1637 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562942 | Headline | Measure |  | C-arm can be moved manually | <XREF>562449</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1638 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562943 | Headline | Hazard |  | Treatment on the wrong side |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1639 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562944 | Headline | Cause |  | Image orientation confused | Image orientation confused due to wrong user input | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Med | C
1640 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562945 | Headline | Measure | docUser_patient_positioned_correctly | Reference to User Manual - Correct Position of Patient | <XREF>562959</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1641 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562946 | Headline | Hazard |  | 3D scan is unavailable |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1642 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562947 | Headline | Cause |  | Orbital drive fails or other error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51 | See general justifications in the introduction which apply here. No additional practical mitigations are available to further reduce or prevent the risk of loss of movement functionality. | Basic | System Test | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1643 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562948 | Headline | Measure |  | No measure required |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1644 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562949 | Headline | Function |  | 3D Imaging |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1645 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562950 | Headline | Hazard |  | Multi Planar Reconstruction (MPR) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1646 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562951 | Headline | Cause |  | Artifacts that might be mistaken for anatomical structures |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1647 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562952 | Requirement | Measure | docCS_image_quality_test_s1 | Artifacts Image Quality test for MPR | Service DocumentsCAUTIONCause: Artifacts in Multi Planar Reconstruction (MPR) 3D imaging, image quality testPossible Consequences: Artifacts could lead to wrong decisions or treatments.Remedy: Perform 3D relevant calibration (Tune UP/Projection Geometry + IZeroCalibration)and IQ test during the regular maintenance interval. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1648 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562953 | Requirement | Measure | docUser_artifacts_that_might_be_mistaken | Artifacts Maintenance recommendation in the user manual | CAUTIONCause: Missing regular maintenancePossible consequences: Insufficient image quality!Remedy: It is recommended to maintain the system regularly to achieve the best possible image quality. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1649 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562954 | Headline | Cause |  | Alignment error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1650 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562955 | Requirement | Measure | docCS_alignment_error_quality_test_s1 | Alignment error Image Quality test for MPR | Image Quality test in service workflow | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1651 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562956 | Headline | Measure |  | Reference - Artifacts Maintenance recommendation in the user manual | <XREF>562953</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1652 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562957 | Headline | Hazard |  | Wrong orientation of the reconstructed 3D volume |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:03 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1653 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562958 | Headline | Cause |  | Wrong input of patient orientation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz27#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | empty | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1654 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562959 | Requirement | Measure | docUser_wrong_input_of_patient_orientation | User Manual - Verify Patient Position and Orientation | CAUTIONCause: Wrong input of patient orientation.Possible consequences: Risk of injury to the patient!Remedy: Please verify the correct patient to system relation and patient orientation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1655 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562961 | Requirement | Measure | picture_and_p_matrices_same_orientation_s1 | Orientation of Image and P-matrices | Encapsulation and separate software tests for volume orientation calculation according to patient position.In the MPR and VRT view, images shall be displayed with an orientation indicator according to the 3D volume information. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1656 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562962 | Headline | Hazard |  | 3D Imaging incorrect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1657 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562963 | Headline | Cause |  | Acquired set of 2D image assigned to wrong P-matrix |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1658 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562964 | Requirement | Measure | correct_assignment_of_2D_s1 | Assure correct assignment | Assure an correct assignment of the 2D images to the P-Matrix throughout the imaging application. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1659 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562965 | Headline | Cause |  | Wrong assignment of P-matrices to scan protocol |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1660 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562966 | Requirement | Measure | correct_assignment_of_scan_protocol_s1 | Correct assignment of scan config | Assure correct assignment of scan config to the correct P-Matrix | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1661 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562967 | Headline | Function |  | 3D reconstruction |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1662 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562968 | Headline | Hazard |  | Functional Hazards 3D Imaging insufficient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1663 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562969 | Headline | Cause |  | Artifact level too high |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1664 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562970 | Headline | Measure |  | Reference - Artifacts Maintenance recommendation in the user manual | <XREF>562953</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1665 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562971 | Headline | Hazard |  | Impaired image acquisition | Impaired image acquisition leads to useless basis for interventional decision and therefore additional strain on the patient | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1666 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562972 | Headline | Cause |  | Patient moves |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to avoid patient movement during a 3D scan, to lock relevant brakes to avoid unintended system movement during a 3D scan, and to use the whole available orbital range for the scan. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1667 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562973 | Requirement | Measure | docUser_patient_moves | User Manual - patient movement during 3D | CAUTIONCause: Patient movement during the 3D scanPossible consequences: Insufficient image quality!Remedy: Patient movement has to be avoided during the 3D scan. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 12.10.2022 13:03 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1668 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562974 | Headline | Cause |  | System/C-armswings | System/C-arm swings due to start up pressure of drive | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1669 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562975 | Requirement | Measure | ramp_function_s1 | Ramp function | Drive starts with ramp function | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1670 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562976 | Headline | Cause |  | System shifts during scan |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to avoid patient movement during a 3D scan, to lock relevant brakes to avoid unintended system movement during a 3D scan, and to use the whole available orbital range for the scan. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1671 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562977 | Requirement | Measure | docUser_system_shifts | User Manual - System shift during 3D | CAUTIONCause: System shifts during the 3D scanPossible consequences: Impaired 3D image acquisition!Remedy: All brakes need to be closed and the system shall not be touched by user to avoid that the system shifts. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.07.2022 16:42 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1672 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562978 | Headline | Cause |  | C-arm angulation does not correspond to calibration area |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1673 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562979 | Requirement | Measure | allowed_angulation_s1 | Retrieval of C-arm angulation | Start of 3D scan only possible if C-arm angulation is inside of the defined allowed 3d scan angulation area (e.g. +/- 15°). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XCU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1674 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562980 | Headline | Cause |  | Image quality for image reconstuction impaired | Image quality for image reconstuction impaired due to insufficient long term stability of the C-arm | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1675 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562981 | Requirement | Measure | system_calibration_via_calibration_phantom_s1 | System calibration via calibration phantom | System calibration via calibration phantom. The system checks the date since the last calibration. After 24 month without recalibration the user receives a warning notice telling him to perform a recalibration. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1676 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562982 | Requirement | Measure | docCS_inspection_of_image_quality_s1 | Inspection of image quality | Service manual:Inspection of image quality during maintenace intervall | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1677 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562983 | Headline | Cause |  | Insufficient geometric fidelity when using the viewer |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1678 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562984 | Requirement | Measure | docUser_insufficient_geometric_fidelity | Statement about fidelity in user manual | CAUTIONCause: Limited geometric fidelity of 3D reconstructionPossible consequences: Reduced 3D image quality or insufficient or incorrect basis for measurements, diagnostic or therapeutic decisionsRemedy: Note that the geometric fidelity of 3D reconstructions can only be maintained within a defined volume. Perform length and angle measurements in 3D images only within this volume and verify the results before diagnostic or therapeutic decisions are taken. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1679 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562985 | Requirement | Measure | docCS_insufficient_geometric_fidelity | Inspection of fidelity | Service DocumentsCAUTIONCause: Reduced accuracy in 3D for functions Measure and AnnotatePossible Consequences: Incorrect interventional decisions or treatment.Remedy: Inspect the geometric accuracy (in IQAP test) during regular maintenance interval. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1680 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562986 | Requirement | Measure | FLC_VDM_image_processing_Scale | Encapsulation 3D image processing scale | Encapsulation and separate software tests for the handling of the model matrix. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 28.07.2022 09:19 | Johannes Schmidt |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1681 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562987 | Headline | Cause |  | Test procedure for inspection of IQ insufficient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1682 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562988 | Requirement | Measure | docCS_inspection_of_iq_insufficient | Application of an appropriate test procedure | Service DocumentsCAUTIONCause: 2D image quality tests are not performed prior to 3D image quality tests.Possible Consequences: Impaired image quality could lead to wrong decisions or treatments and / or additional doseRemedy: Perform 2D image quality tests before performing the 3D image quality tests (3D IQ test procedures can only be performed, when 2D tests are all succeeded). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1683 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562989 | Requirement | Measure | FLC_VDM_image_processing_CW_Calculation | Encapsulation 3D handling of brightness / contrast related data | Encapsulation and separate software tests of handling the Brightness/Contrast related data (rescale slope, rescale intercept). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 05.07.2022 09:04 | Dominik Tal (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1684 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562990 | Headline | Function |  | 2D/3D Navigation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1685 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562991 | Headline | Hazard |  | Additional strain on patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1686 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562992 | Headline | Cause |  | Disturbance of data transfer |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1687 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562993 | Requirement | Measure | use_defined_tcp_ip_protocol_s1 | Use of a defined protocol | Use of a defined protocol (NaviLink) based on TCP/IP | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1688 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562994 | Headline | Hazard |  | Risk of injury by C-arm |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1689 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562995 | Headline | Cause |  | Unintended shifting of the C-arm | Unintended shifting of the C-arm unit when releasing the brake due to the additional weight for the angulation and orbital movement. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:48 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1690 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562996 | Requirement | Measure | docUser_unintended_shifting_of_the_c_arm | User Manual - Non-Counterbalanced C-arm during Navigation | CAUTIONCause: Additional weight after installation of parts for navigationPossible consequences: Risk of injuryRemedy: Note that the C-arm is no longer balanced after additional components are attached. It can move on its own after brakes are released. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1691 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562997 | Headline | Cause |  | Incorrectly mounted navigation components impair accuracy |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1692 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562998 | Requirement | Measure | docUser_incorrectly_mounted_navigation_system | User Manual - Ensure Correct Mounting of Navigation System | CAUTIONCause: Incorrectly attached navigation system.Possible consequences: Risk of injury!Remedy: Before navigation-assisted operation, the user must ensure sufficient accuracy of the navigation system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1693 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 562999 | Headline | Cause |  | Mechanical connection between detector housing and adapter fails |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1694 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563000 | Headline | Measure | proof_of_fourfold_stability_s1 | Reference - Design fit for fourfold weight | <XREF>563023</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:36 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1695 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563001 | Headline | Hazard |  | Risk of injury during operation/transportation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1696 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563002 | Headline | Cause |  | Tipping over of the device |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09#Haz42 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1697 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563003 | Headline | Measure |  | Reference - Design of system | <XREF>562067</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
1698 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563004 | Requirement | Measure | docUser_tipping_over_of_the_device | User Manual - Maximum Weight | WARNINGCause: Heavy weight of external navigation components.Possible consequences: Risk of injury!Remedy: Note that stability in accordance with DIN EN/UL standards is guaranteed only if the external system (navigation) on the detector does not exceed the calculated maximum weight. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
1699 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563005 | Headline | Hazard |  | Injury due to surface condition |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1700 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563006 | Headline | Cause |  | Sharp edges |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-D | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Med | D
1701 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563007 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 external | <XREF>563042</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | empty | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1702 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563008 | Headline | Hazard |  | Risk of injury due to breakage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1703 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563009 | Headline | Cause |  | Overloading of C-arm bearing and supporting elements for the detector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | System Integration Test | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1704 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563010 | Headline | Measure |  | Reference - Design of the movement unit | <XREF>562415</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1705 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563011 | Requirement | Measure | docUser_overloading_of_c_arm | User Manual - Break Resistance | CAUTIONCause: Overloading of C-arm.Possible consequences: Risk of injury!Remedy: Note that the braking action is guaranteed only if the external system (navigation) on the detector does not exceed the calculated maximum weight. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1706 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563012 | Headline | Hazard |  | Scattered radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1707 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563013 | Headline | Cause |  | Unnecessary radiation exposure when additional objects in beam path. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1708 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563014 | Requirement | Measure | docUser_additional_objects_beam_path | User Manual - Scattered Radiation | CAUTIONCause: Additional objects in beam path.Possible consequences: Risk of unnecessary radiation exposure!Remedy: Note that additional objects in the beam path increase scatter radiation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1709 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563015 | Headline | Hazard |  | C-arm collides with people or goods |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1710 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563016 | Headline | Cause |  | Additional weight |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1711 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563017 | Requirement | Measure | docUser_additional_weight | User Manual - Mounting of Additional Parts | CAUTIONCause: Changing the weight distributionPossible consequences: Risk of crushing!Remedy: Adding weight to the detector or single tank side means a loss in counterweight and can result in unintended movement of the C-arm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1712 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563019 | Headline | Hazard |  | Markers of different nav. Systems can fall off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1713 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563020 | Headline | Cause |  | Insufficient dimensioning of the mechanical interface |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1714 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563021 | Requirement | Measure | docNav_insufficient_dimensioning | Navigation Interface - Marker Frame Fixation | The dimensioning of the mechanical interface at the marker frame must be sufficient to enable a tight and secure fixation of the marker frame at the FD. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 15:44 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1715 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563022 | Headline | Cause |  | Reference ring loosenes |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1716 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563023 | Requirement | Measure | design_fit_for_fourfold_weight_s1 | Design fit for fourfold weight | Design fit for fourfold stability; according to IEC 60601-1-1:2005, clause 9.8.2 (tensile safety factor). | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:36 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1717 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563024 | Headline | Cause |  | Markers insufficiently fixed |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1718 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563025 | Requirement | Measure | docNav_markers_insufficiently_fixed | Navigation Interface - Marker Fixation | The markers attached to the marker frame (e.g. spheres) must be fixed sufficiently to prevent markers from falling down. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1719 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563026 | Headline | Hazard |  | Risk of injury during operation/transportation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1720 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563027 | Headline | Cause |  | Tipping over of the device in general |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz09 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | B | High | A
1721 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563028 | Headline | Measure |  | Reference - Design of system | <XREF>562067</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1722 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563029 | Headline | Measure | docNav_tipping_over_of_the_device | Reference - Navigation Interface - Maximum Weight | <XREF>563734</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Verification&Validation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1723 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563033 | Headline | Hazard |  | Wrong treatment due to insufficient precision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1724 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563034 | Headline | Cause |  | Dirty/Damaged Mechanical Detector Interface |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1725 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563035 | Requirement | Measure | docUser_visual_inspect_navi_interface | User Manual - Visual Inspection/Cleaning | CAUTIONCause: Insufficient navigation accuracy due to contaminated or defective detector interface.Possible consequences: Insufficient navigation accuracy, risk of injury.Remedy: Prior to using the navigation system check the required cleanliness and functionality. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1726 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563036 | Headline | Hazard |  | Risk of injury due to breakage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1727 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563037 | Headline | Cause |  | Overloading of C-arm bearing and supporting elements for II |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1728 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563038 | Headline | Measure |  | Reference - Design of the movement unit | <XREF>562415</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1729 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563039 | Headline | Hazard |  | Electric schock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1730 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563040 | Headline | Cause |  | Isolation error of nav. parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1731 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563041 | Requirement | Measure | docUser_isolation_error_of_3d_nav_parts | User Manual - Components Comply With IEC60601-1 | WARNINGCause: Use of insufficiently electrically-protected external navigation components.Possible consequences: Risk of injury!Remedy: Use only accessories corresponding to the CE Declaration of Conformity. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1732 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563042 | Requirement | Measure | docNav_isolation_error_of_nav_parts | Navigation Interface - Comply with IEC60601-1 | The IEC 60601-1:2005 for navigation systems must be adhered to. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1733 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563043 | Headline | Function |  | 3D Navigation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1734 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563044 | Headline | Hazard |  | Collision with catheters |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1735 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563045 | Headline | Cause |  | Operation error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz03 | See general justifications in the introduction which apply here. The operator manual advises to always pay attention for potential collisions, especially in the vicinity of the patient or critical equipment or infrastructure. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1736 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563046 | Requirement | Measure | docUser_manual_drive | User Manual - Dry Run Manually | WARNING:Cause: Collision with e.g. catheters or other vital components in ORPossible consequences: Risk of injury!Remedy: Before starting the 3D scan, you must do a manual test drive to prevent collisions. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1737 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563047 | Headline | Hazard |  | 3D navigation accuracy reduced |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1738 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563048 | Headline | Cause |  | Start position 3D scan |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1739 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563049 | Requirement | Measure | accuracy_of_mechanical_makrer_interface_s1 | The accuracy of the mechanical marker interface | The geometry of the mechanical marker interface has to fulfill the tolerance +/-0,02 mm. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.03.2022 13:36 | Rudolf Siebenhaar |  |  |  | MainU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1740 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563050 | Requirement | Measure | docUser_start_position_3d_scan | User Manual - Verify Accuracy | CAUTIONCause: Absolute position accuracy of the mechanical interface insufficient.Possible consequences: Risk of injuryRemedy: Accuracy has to be verified after 3D navigation system has received the 3D volume data. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1741 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563051 | Headline | Hazard |  | Wrong treatment due to insufficient precision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1742 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563052 | Headline | Cause |  | Damage during transport |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1743 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563053 | Headline | Measure |  | Reference - User Manual - Visual Inspection/Cleaning | <XREF>563035</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1744 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563054 | Headline | Cause |  | Calibration-phantom interface damaged |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | E
1745 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563055 | Headline | Measure |  | Reference - User Manual - Visual Inspection/Cleaning | <XREF>563035</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1746 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563056 | Headline | Cause |  | Recalibration for 3D-mode |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1747 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563057 | Requirement | Measure | docUser_recalibration_for_3d_mode | Admin Guide - Navigation system test after recalibration | CAUTIONCause: Recalibration for 3D modePossible consequences: Risk of injuryRemedy: After recalibrating in 3D mode, a test must be carried out using the navigation system. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 19.07.2022 16:30 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Admin Guide | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1748 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563058 | Requirement | Measure | docCS_recalibration_for_3d_mode | Service Documentation - Navigation system test after recalibration | Service DocumentationCAUTIONCause: After a geometry calibration for 3D mode, or after modifications impairing the operation or accuracy of the´navigation system, no recalibration with the relevant navigation system is performed.Possible consequences: Inaccurate navigation results, risk of injury for the patient.Remedy:- After a 3D geometry calibration or any modification impairing the operation or accuracy of the navigation system, a recalibration needs to be performed with the relevant navigation system.- The operator shall be informed verbally and / or in writing, that the accuracy of the navigation system must be checked.- The relevant parts that influence the navigation accuracy are listed in the Replacement of Parts document. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1749 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563059 | Headline | Cause |  | Faulty navigation system during calibration |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1750 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563060 | Requirement | Measure | docUser_faulty_navigation_system_during_calibration | User Manual - Check Precision of Navigation | CAUTIONCause: Faulty navigation system during calibrationPossible consequences: Risk of injuryRemedy: Before performing an operation via navigation, the user must ensure that the navigation accuracy is adequate. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1751 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563061 | Headline | Hazard |  | Wrong interventional decision or treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1752 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563062 | Headline | Cause |  | Wrong selection by user |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:53 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to perform the required checks of the navigation equipment and functionality prior to use it for navigation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1753 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563063 | Requirement | Measure | docUser_wrong_selection_by_user | User Manual - Choose Correct Navigation | CAUTIONCause: Wrong selection of navigation system by user.Possible consequences: Risk of injury!Remedy: Before starting a scan, make sure that the correct navigation unit is selected. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1754 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563064 | Headline | Hazard |  | Additional strain on patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1755 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563065 | Headline | Cause |  | Incorrect registration of the marker |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) |  |  |  |  | Follow-up: 06H561768 | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1756 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563066 | Requirement | Measure | no_scan_without_confirmation_s1 | No scan is possible without confirmation | Def. Handshake for registration in interface description: No scan is possible without confirmation from the nav. system | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC | Follow-up: 06H561768 |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1757 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563067 | Requirement | Measure | docNav_no_scan_without_confirmation_s1_navisystem | Navigation Interface - Scan Confirmation | A defined handshake for registration shall be used to ensure, that no scan is possible without confirmation from the navigation system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  |  | Follow-up: 06H561768 |  |  | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1758 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563068 | Headline | Cause |  | Rescan due to user error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1759 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563069 | Headline | Measure |  | No additional measures required | No additional measures are required. The operators are trained. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1760 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563070 | Headline | Cause |  | Rescan due to communication issue |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1761 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563071 | Headline | Measure |  | No additional measures required | No additional measures are required. The operators are trained. When communication is established again, the operator can rescan. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1762 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563072 | Headline | Hazard |  | Short-term delay |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,use by unskilled/untrained personnel,data transfer,Untrained users |  | S1M | VB10 | [12 empty] | NotDetermined
1763 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563073 | Headline | Cause |  | Communication issues |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40#Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1764 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563074 | Headline | Measure |  | No additional measures required | No additional measures are required. In case the handshake fails, no scan is possible with navigation system. However 3D without navigation is possible. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1765 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563075 | Headline | Cause |  | User error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1766 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563076 | Headline | Measure | no_scan_without_confirmation_s1 | Reference - No scan is possible without confirmation | <XREF>563066</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1767 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563077 | Headline | Measure | docNav_no_scan_without_confirmation_s1_navisystem | Reference - Navigation Interface - Scan Confirmation | <XREF>563067</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1768 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563078 | Headline | Function |  | Application Hosting |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1769 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563079 | Headline | Hazard |  | Influence on performance or functionality of hosted application |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Lack of integrity of data,Loss of data,unintended interactions within and among SW sys,unknown aspects or quality of third-party software,lack of data sec, particularly vulnera. to tampering,Malicious software,Input from AT SOP 418 [11] |  | S1M | VB10 | [12 empty] | NotDetermined
1770 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563080 | Headline | Cause |  | Unauthorized update or execution of unauthorized software |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1771 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563081 | Requirement | Measure | ProvideIntrusionDetection_s1 | Intrusion detection solution | An intrusion detection solution based on whitelisting is provided by the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1772 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563082 | Headline | Hazard |  | Application Hosting PC does not start |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1773 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563083 | Headline | Cause |  | Application PC does not start due to HW problem, boot problem |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 17:05 | Klaus-Peter Fichtner (EXT) | [6 empty] | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1774 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563084 | Headline | Measure |  | None | The risk depends on the application. Application cannot be executed. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | AppHost |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1775 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563085 | Headline | Function |  | Wireless footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1776 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563086 | Headline | Hazard |  | Unwanted radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Inadequate identification |  | S1M | VB10 | [12 empty] | NotDetermined
1777 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563087 | Headline | Cause |  | User mistakenly uses a wireless footswitch from another system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1778 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563088 | Requirement | Measure | docServ_labeling_wireless_footswitch | Label the wireless footswitch with location | In Service DocumentationCAUTIONCause: Radiation released from a footswitch in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy: Before performing any service activities with the wireless footswitch, check that the footswitch belongs to the system and check that the indicating label is in place and properly visible. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1779 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563089 | Requirement | Measure | docUser_labeling_wireless_footswitch | User Manual - Label the wireless footswitch with location | CAUTIONCause: Radiation released from a footswitch in another roomRisk: Non-clinical exposure to ionizing radiation.Remedy: As part of the system installation, the wireless footswitch is labeled according to which system it belongs to.When utilizing the wireless footswitch, check that it belongs to the system according to the indicating labels. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 23.06.2022 13:33 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1780 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563090 | Requirement | Measure | wireless_footswitch_pairing | Wireless footswitch and system are paired | An initial pairing has to be done. This pairing shall remain, even in the case of reboot, reconnection of the wireless footswitch or the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1781 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563091 | Requirement | Measure | label_wireless_footswitch_with_location | Label - wireless footswitch with location | A label shall be attached on wireless footswitch to which system he belongs to. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Judith Welzel |  |  | 29.07.2020 | 18.05.2022 10:21 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1782 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563092 | Headline | Cause |  | Send/Receive modules of wireless footswitch are not reliable |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
1783 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563093 | Requirement | Measure | wireless_footswitch_save_modules | Safe hardware modules and signal transmission line | Usage of safe hardware modules and signal transmission line (e.g redundant hardware contacts and/or CRC for transmission line) | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1784 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563094 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1785 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563095 | Headline | Cause |  | x-ray indicators and patient not visible for the user |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1786 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563096 | Requirement | Measure | docUser_release_x-ray_close_to_patient | User Manual -Do not release x-ray out of vieweing distance | CAUTIONCause: X-ray release may be possible from remote locations.Risk: Unnecessary dose.Remedy: Do not release X-ray if you are not in viewing distance to the patient. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1787 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563097 | Headline | Hazard |  | Delayed or no treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Availability in critical functions | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1788 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563098 | Headline | Cause |  | Interference with other wireless Footswitch |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50#Haz40#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1789 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563099 | Requirement | Measure | docUser_release_x-ray_close_to_patient_wireless | User Manual - Interference wireless footswitch | CAUTIONCause: The wireless footswitch may be interfered by other equipment, including portable and fixed RF communication equipment, even if such equipment meets the applicable emissions requirements.Risk: Fluoro or acquisition may sporadically be interrupted.Remedy:- The operator must assure that other wireless device in the 2,4 GHz ISM band should not be operated in the vicinity of appr. 5m around the system.- Please observe and verify normal operation of the wireless footswitch prior to using it. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1790 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563100 | Headline | Cause |  | Battery empty |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1791 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563101 | Requirement | Measure | docUser_battery_empty | User Manual -Do not release x-ray out of vieweing distance | CAUTIONCause: Weak batteries of wireless footswitch.Risk: Delay in the clinical procedure.Remedy: Make sure the batteries are charged prior to starting an examination. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1792 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563102 | Requirement | Measure | wireless_footswitch_battery_capacity | High capacity battery | Usage of a battery with the possibility to work with wireless footswitch for 48 hours at least (see table 12 use case 4 for use profile). The user shall be informed via LEDs about low capacity (usage<=24 hours under conditions of table 12 use case 4) to recharge or exchange the battery. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1793 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563103 | Headline | Chapter |  | Service System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1794 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563104 | Headline | Function |  | Remote service |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1795 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563105 | Headline | Hazard |  | Malfunction of the examination routine |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1796 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563106 | Headline | Cause |  | Access via remote service |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to check functioanlity after the update. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1797 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563107 | Headline | Measure |  | Reference - Remote access | <XREF>563179</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1798 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563108 | Headline | Cause |  | Install Software on wrong system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1799 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563109 | Requirement | Measure | conf_MNP_SPS_check_s1 | Check system requirements before installation starts | It is required that the Software Package Source checks the needed system requirements before the installation starts. Therefore an installation on a wrong system is impossible. Note, see 11108300 EKO 000 Configuration Management Plan | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 15:50 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Configuration Management | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1800 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563110 | Headline | Function |  | Software modification incl. Bugfix |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1801 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563111 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Malicious software |  | S1M | VB10 | [12 empty] | NotDetermined
1802 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563112 | Headline | Cause |  | Impermissible or faulty manipulation of SW from network. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1803 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563113 | Headline | Measure |  | Reference - User Manual - Impermissible or faulty manipulation from network | <XREF>562154</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1804 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563114 | Headline | Cause |  | Software virus |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1805 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563116 | Headline | Cause |  | Virus scanner degrades performance and system reaction |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1806 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563118 | Headline | Cause |  | Virus patterns and scan engines are not up to date |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:55 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1807 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563120 | Headline | Cause |  | Non approved SW/HW |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 #Haz42 #Haz52 #Haz01 | See general justifications in the introduction which apply here. The operator manual warns to not perform non-approved modifications of the system, incl. combining with non-approved accessories or other devices. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1808 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563121 | Headline | Measure |  | Reference - User Manual - Use of non-approved parts | <XREF>562152</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1809 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563122 | Requirement | Measure | serv_Service_SW_Download | Use of unauthorized software | Use of unauthorized software or manipulations and alterations of the software can cause the product to function incorrectly and cause injury to the patient and damage to the product. Use software with correct version approved by Siemens. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1810 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563123 | Headline | Function |  | Adjustment (e.g. dose rate adjustment) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1811 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563124 | Headline | Hazard |  | Unwanted radiation released only by SW (closed HW loop in Service Mode, e.g. during calibration) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1812 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563125 | Headline | Cause |  | SW or communication error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. The operator manual gives advice how to react in the rare case of occurrence. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1813 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563128 | Requirement | Measure | serv_ClosedHWloop_s1 | Open HW-loop after Service | Service DocumentsCAUTIONCause: Hardware loop still active after service (radiation release switch S3) and rebootPossible Consequences: X-ray is prohibited, delay of procedureRemedy: Set the radiation release switch S3 to the correct position after service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1814 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563129 | Headline | Hazard |  | Increased radiation exposure during service activities |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output |  | S1M | VB10 | [12 empty] | NotDetermined
1815 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563130 | Headline | Cause |  | Radiation safety guidelines are not followed |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | A | Low | A
1816 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563131 | Requirement | Measure | docCS_radiation_safety | Service Manual - Radiation safety | Service Documents WARNING Cause: Non-compliance with radiation safety guidelines Possible consequences: Increased probability of deterministic and stochastic radiation effects. Remedy:- Follow the radiation safety guidelines and the local regulations for radiation safety- Increase the distance to the radiation source- Release radiation only when required- Use low radiation parameters- Use radiation protection, such as protection clothing | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1817 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563132 | Headline | Function |  | Replace electrical parts / Replace mechanical parts |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1818 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563133 | Headline | Hazard |  | Environmental hazard |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Inadequate Disposal and scrapping |  | S1M | VB10 | [12 empty] | NotDetermined
1819 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563134 | Headline | Cause |  | Improper disposal of hazardous materials |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz38 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1820 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563135 | Headline | Measure | serv_Dismantling_s1 | Dismantling | <XREF>562246</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 10:02 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1821 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563136 | Headline | Hazard |  | Eye injury |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation,Insufficient control of processes,Insufficient training of personnel |  | S1M | VB10 | [12 empty] | NotDetermined
1822 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563137 | Headline | Cause |  | Technician looks into laser |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz56 | See general justifications in the introduction which apply here. The service documentation warns to take care of laser light. The technicians are trained. No additional practical mitigations are available to further reduce or prevent eye exposure to the laser beam. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1823 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563139 | Headline | Hazard |  | Thermal energy and high temperature |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1824 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563140 | Headline | Cause |  | Parts too warm |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1825 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563141 | Requirement | Measure | docServ_TCP_hot_surface | Service Instruction - Hot Surface of TCP backside | Service DocumentsCAUTIONCause: Hot surface of TCP backsidePossible Consequences: Risk of burningsRemedy: While replacing trolley control panel (TCP), please note, surface on backside of TCP maybe hot. Handle with care! | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1826 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563142 | Headline | Hazard |  | ESD |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation | Patient,System | S1M | VB10 | [12 empty] | NotDetermined
1827 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563143 | Headline | Cause |  | ESD related damages of electronic components |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Deterioration in function,Loss or deterioration of function,Availability in critical functions | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1828 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563144 | Requirement | Measure | docCS_ESD | Service Instruction - ESD | Service Documentation CAUTIONCause: ESDPossible Consequences: Damage of electronic components leading to deterioration in system functionality.Remedy: Please follow the ESD regulations during all work steps. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1829 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563145 | Headline | Hazard |  | Spring jumps out when replacing the standard monitor holder on systems with a standard monitor column (non-motorized) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy,Moving parts | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1830 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563146 | Headline | Cause |  | Removing the snap ring without holding the spring |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1831 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563147 | Requirement | Measure | docCS_spring_monitor_holder | Service Instruction - Hold spring of monitor holder | Service DocumentationCAUTIONCause: Removing the snap ring of the monitor holder without holding the spring.Possible Consequences: Spring can jump out, risk of injury.Remedy: While removing the snap ring, hold the spring with the other hand. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1832 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563148 | Headline | Hazard |  | Inaccurate navigation reesult |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Release of stored energy,Moving parts | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined
1833 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563149 | Headline | Cause |  | Removing and remounting the FD cover and optional mechanical interface of the navigation system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function,Loss or deterioration of function | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1834 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563150 | Requirement | Measure | docCS_inform_customer_navigation_accuracy | Service Instruction - Inform customer to check navigation accuracy | Service DocumentationCAUTIONCause: Removing and remounting the FD cover and optional mechanical interface of the navigation system.Possible consequences: Inaccurate navigation results, risk of injury for the patient.Remedy:- After remounting the FD cover and the mechanical interface for the navigation system, check the correct mounting (clearance).- Inform the operator verbally and / or in writing, that the accuracy of the navigation system must be checked. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1835 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563151 | Headline | Function |  | UPS operation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1836 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563152 | Headline | Hazard |  | Electrical shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage,Insufficient control of processes,Insufficient training of personnel |  | S1M | VB10 | [12 empty] | NotDetermined
1837 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563153 | Headline | Cause |  | Power is not switched off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1838 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563154 | Requirement | Measure | serv_WarnSetUPSPowerMayBePresent_s1 | Service Instruction - Power Presence in System | Service DocumentsWARNINGCause: Power input connector of the imaging system PC is not unplugged or UPS is not switched off during service.Possible Consequences: Risk of death or serious injury due to electric shock when electrified components are touched.Remedy: Prior to beginning service activities, unplug the power input connector in the imaging system PC or switch off the UPS. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1839 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563155 | Requirement | Measure | manu_WarnStickerRemovePlugForService_s1 | Warning labeling | Warning labelingWarning labeling at the power inlet of the image system.Unplug the power input connector for service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1840 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563156 | Requirement | Measure | label_unplug_power_input_connector | Label - Unplug the power input connector | A label shall be attached on Trolley housing to unplug the power input connector for service. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.04.2022 15:24 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1841 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563157 | Headline | Cause |  | Inappropiate devices connected to UPS |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 14:59 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | B | Med | B
1842 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563158 | Requirement | Measure | um_UPS | User Manual - UPS | WARNINGCause: Unsuitable devices connected to the UPS.Possible consequences: Risk of electrical shock or damage to the system.Remedy: Do not connect any additional devices to the UPS sockets. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1843 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563159 | Requirement | Measure | docCS_UPS | Service Instruction UPS | Service DocumentsWARNINGCause: Inappropriate devices connected to UPSPossible consequences: Risk of electrical hazard or damage to the systemRemedy: Do not connect any additional device to the UPS power outlets | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1844 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563160 | Headline | Function |  | Installation General |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1845 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563161 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1846 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563162 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1847 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563163 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1848 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563164 | Headline | Function |  | Software installation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1849 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563165 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1850 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563166 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1851 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563167 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1852 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563168 | Headline | Function |  | Service General |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1853 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563169 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1854 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563170 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1855 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563171 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1856 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563172 | Headline | Function |  | Report unexpected event |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1857 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563173 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1858 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563174 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1859 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563175 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1860 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563176 | Headline | Function |  | Remote service |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1861 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563177 | Headline | Hazard |  | Malfunction of the examination routine |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Moving parts,Incorrect or inappropriate output,user interface |  | S1M | VB10 | [12 empty] | NotDetermined
1862 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563178 | Headline | Cause |  | Access via remote service |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to check functioanlity after the update. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1863 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563179 | Requirement | Measure | RemoteService | Remote access | Only selected service functionality shall be available for remote service in limited access mode.No customer functions shall be available during remote service in full access mode.Remote service shall be only possible when approved by the operator.An indication shall be provided to the user that remote service with full access has been enabled.The user shall still be able to release radiation. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1864 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563180 | Headline | Hazard |  | Remote software update |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Moving parts,Incorrect or inappropriate output,data acquisition, evaluation and display,user interface |  | S1M | VB10 | [12 empty] | NotDetermined
1865 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563181 | Headline | Cause |  | Installing the hotfix or software update |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to check functioanlity after the update. | Basic | n.a. | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1866 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563182 | Requirement | Measure | um_RemoteSoftwareUpdate_192229 | User Manual - Remote Software Update | CAUTIONCause: Software installation interferes with system functions.Possible consequences: Interruption of patient examination and loss of patient data.Remedy: Do not start the installation while working with patients. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1867 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563183 | Headline | Cause |  | Failed installation of software |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to check functioanlity after the update. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1868 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563184 | Requirement | Measure | um_RemoteUpdateFailure_192230 | User Manual - Remote Update Failure | CAUTIONCause: Failure installing the update package.Possible consequences: A failed software installation means that the system is in an undefined state.Remedy: Stop using the system and notify the UPTIME Service Center. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None | Inadequate specification of accessories |  | S1M | VB10 | [12 empty] | NotDetermined
1869 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563185 | Headline | Function |  | Cables |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1870 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563186 | Headline | Hazard |  | Electric Shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1871 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563187 | Headline | Cause |  | Moving Components |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1872 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563188 | Headline | Measure |  | Reference - Cable conduit for movable cables | <XREF>562037</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1873 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563189 | Headline | Function |  | Service documentation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1874 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563190 | Headline | Hazard |  | No hazard on system level |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1875 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563191 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1876 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563192 | Headline | Measure |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1877 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563193 | Headline | Function |  | Maintenance, service / trouble shooting |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1878 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563194 | Headline | Hazard |  | Impaired vision through laser-beam |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Non-ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1879 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563195 | Headline | Cause |  | Service staff looking into laser-beam |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz34 | See general justifications in the introduction which apply here. The service documentation warns to take care of laser light. The technicians are trained. No additional practical mitigations are available to further reduce or prevent eye exposure to the laser beam. | Basic | n.a. | none | n.a. | Not Determined | 1-A | 1-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1880 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563196 | Headline | Measure |  | Reference - Usage of laser class 1 | <XREF>562338</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | ReadyForTCReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1881 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563197 | Headline | Hazard |  | Excessive temperature, fire, explosion |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Flammability,High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1882 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563198 | Headline | Cause |  | Improper exchange of batteries |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 #Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1883 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563199 | Requirement | Measure | label_exchange_battery | Warning Label - Exchange Battery UPS | A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 30.03.2022 14:46 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1884 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563200 | Requirement | Measure | label_exchange_battery_2 | Warning Label - Exchange Battery HICSwitch | A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1885 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563201 | Headline | Cause |  | Short-circuit of batteries during exchange |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 #Haz14 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | C | Med | C
1886 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563202 | Requirement | Measure | safe_battery_cases | Safe battery cases TrolleyU | Use battery cases or housings that prevent an accidental short-circuit | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 30.03.2022 14:48 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
1887 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563203 | Requirement | Measure | safe_battery_cases_2 | Safe battery cases HICSwitch | Use battery cases or housings that prevent an accidental short-circuit | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | HICSwitch |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
1888 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563204 | Requirement | Measure | docCS_exchange_batteries | Service Caution - Servicing battery components | Service Documentation CAUTION Cause: Incorrect handling of battery components during service, for example, wearing metal objects such as rings, watches or bracelets, or using tools with non-insulated handles when servicing the battery components, could lead to an accidental short-circuit of the battery contacts. Potential consequences: Risk of burns due to high short-circuit current and of destroying the battery component. Remedy: Before servicing the battery components, remove watches, rings, bracelets and all other metal objects from the hands and arms. Use tools with insulated handles. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | User,Service Personnel,System,Environment | S1M | VB10 | [12 empty] | NotDetermined
1889 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563205 | Headline | Hazard |  | Risk of Infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection |  | S1M | VB10 | [12 empty] | NotDetermined
1890 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563206 | Headline | Cause |  | Opening Chassis / Trolley |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1891 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563207 | Requirement | Measure | docServ_self_protection | Service Instruction - Measures to prevent infection due to contaminated system parts - service | Warning in Service InstructionWARNINGCause: Service, disassembly or refurbishment activities on hardware parts of Cios systems which had been in clinical use and are potentially contaminated.Possible consequences: Increased risk of infection due to residues of infectious substances such as bodily fluids or deposits of organic matter colonized with microorganismsRemedy: Consequently apply precautionary measures for handling potentially infectious material, especially:- Wear protective gloves, closed and long-sleeved protective clothing, protective shoes, and protective glasses.- In case of severe contamination or strong infectiousness, in case the contamination may be dusting or emit droplets, apply additional measures such as face masks and disposable protective clothing, for example overalls and caps.- After finishing the activities on contaminated hardware parts, clean and disinfect potentially contaminated body parts, especially hands and face, as well as clothing and equipment. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1892 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563208 | Headline | Function |  | Imaging System, Field service |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1893 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563209 | Headline | Hazard |  | Wrong interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Workflow and configuration |  | S1M | VB10 | [12 empty] | NotDetermined
1894 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563210 | Headline | Cause |  | Service parameters changed by unauthorized user |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41 #Haz51#Haz31 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Workflow and configuration | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1895 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563211 | Headline | Measure |  | Reference - User Manual - Trained Users | <XREF>561986</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1896 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563212 | Headline | Measure |  | Reference - Service training | <XREF>562186</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1897 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563213 | Headline | Function |  | 3D Service functions |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1898 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563214 | Headline | Hazard |  | Wrong interventional decision or treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1899 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563215 | Headline | Cause |  | Unprecise 3D calibration due to a badly positioned calibration phantom |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1900 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563216 | Requirement | Measure | serv_unprecise_3d_calibration | Correct Position of 3D-Calibration Phantom | Service DocumentationCAUTIONCause: Imprecise 3D calibrationPossible consequences: Wrong interventional decision or treatmentRemedy: Make sure that the calibration phantom is positioned correctly when calibrating the 3D system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1901 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580671 | Requirement | Measure | docuser_unprecise_3d_calibration | Correct Position of 3D-Calibration Phantom | Admin GuideCAUTIONCause: Imprecise 3D calibrationPossible consequences: Incorrect interventional decision or treatmentRemedy: When calibrating the 3D system, make sure that the calibration phantom is correctly positioned. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 06.07.2021 | 21.09.2022 11:53 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Admin Guide | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1902 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563217 | Headline | Cause |  | Wrong 3D-calibration due to wrong transformation matrices |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-A | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Low | C
1903 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563218 | Requirement | Measure | docCS_wrong_3d_calibration | measuring instrument boards are positioned correctly | Service DocumentsCAUTIONCause: Incorrect 3D calibration (wrong transformation matrices) due to incorrect position of measuring instrumentsPossible Consequences: Wrong interventional decision or treatment.Remedy: Make sure that the measuring instruments (navigation references like marker frame and marker plate) are positioned correctly in the measuring volume of the 3D camera when calibrating the 3D system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1904 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563219 | Headline | Hazard |  | Risk of injury due to structure of surface |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1905 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563220 | Headline | Cause |  | Sharp edges |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz18 | See general justifications in the introduction which apply here. | Basic | System Test | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | C | Med | C
1906 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563221 | Requirement | Measure | docNav_sharp_edges | Note in interface description | Note in interface description for nav. system: Attachable hardware shall have no sharp edges. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 21.06.2022 15:54 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Supplier | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1907 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563222 | Headline | Chapter |  | Shutdown System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1908 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563223 | Headline | Function |  | Shut down system & restart |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1909 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563224 | Headline | Hazard |  | Electrical safety, Electrical shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage,Line voltage,Leakage current,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
1910 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563225 | Headline | Cause |  | Electrical energy is present after shutdown. |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1911 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563226 | Requirement | Measure | LabelPowerPresent_s1 | Warning label | Warning Labels show where voltage is present after shutdown. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | XGen,PowerU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1912 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563227 | Headline | Measure |  | Reference - Service Documentation - Voltage Presence | <XREF>563280</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1913 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563228 | Headline | Hazard |  | Loss of data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Software failure,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
1914 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563229 | Headline | Cause |  | System shutdown because of a software failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50 #Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1915 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563230 | Headline | Measure |  | Reference - Protect shutdown | <XREF>563232</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1916 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563231 | Headline | Cause |  | User accidentally shuts off system during examination |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50 #Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | C
1917 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563232 | Requirement | Measure | shutdown_protect_s1 | Protect shutdown | UPS must be available. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 04.04.2022 15:25 | Rudolf Siebenhaar |  |  |  | TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1918 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563233 | Headline | Cause |  | System restart after X-ray exposure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz40 #Haz50 #Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
1919 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563234 | Headline | Measure |  | No measure required |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1920 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563235 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
1921 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563236 | Headline | Cause |  | System automatically switches on/off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51 #Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. In case of a recoverable function, the operator manual advises the to restart the system. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1922 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563237 | Headline | Measure |  | Reference - Display system status | <XREF>562150</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | N.A. | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1923 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563238 | Headline | Function |  | Clean System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1924 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563239 | Headline | Hazard |  | Risk of infection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection,Other agents,Pollution |  | S1M | VB10 | [12 empty] | NotDetermined
1925 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563240 | Headline | Cause |  | Improper cleaning |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz29 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1926 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563241 | Requirement | Measure | docUser_note_improper_cleaning | User Manual - Cleaning of Wheels | CAUTIONCause: Improper cleaningPossible consequences: Risk of infection!Remedy: After dealing with patients with highly infectious diseases, clean the unit wheels, in addition to regular cleaning. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined
1927 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563242 | Headline | Cause |  | Contaminated device wheels |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz29 | See general justifications in the introduction which apply here. The operator manual advises to clean and drape the system appropriately and to position the system and its components with the appropriate carefullness. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1928 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563243 | Headline | Measure |  | Reference - User Manual - Cleaning of Wheels | <XREF>563241</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined
1929 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563244 | Headline | Cause |  | Insufficient cleaning |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz29 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1930 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563245 | Requirement | Measure | clean_compliant_IEC60601_1 | Compliant with IEC60601-1 | The enclosure shall be designed in a way that it can be easily cleaned and wipe disinfected. The enclosure must be compliant with IEC60601-1:2005 chapter 11.6.6 & 7.9.2.12 regarding cleaning and disinfection requirements. Fit criterion:All different enclosure surfaces must be compliant with MFR 890/009e - Resistance of surfaces to disinfection and cleaning agents. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 20.06.2022 09:39 | Johannes Schmidt |  |  |  | GridS,HICTouchCon,MainU,TrolleyU |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1931 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563246 | Headline | Function |  | Image data storage |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1932 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563247 | Headline | Hazard |  | Unnecessary radiation |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1933 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563248 | Headline | Cause |  | Image data loss |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40#Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1934 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563249 | Requirement | Measure | controlled_shutdown | The imaging system performs a controlled shutdown | The imaging system automatically performs a controlled shutdown, before the power of the Uninterrupted Power Supply runs out. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | Accepted | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User | S1M | VB10 | [12 empty] | NotDetermined
1935 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563250 | Headline | Chapter |  | Additional hazards from subsystems |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1936 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563251 | Headline | Function |  | Movements during patient positioning |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1937 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563252 | Headline | Hazard |  | Crushing of patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Moving parts |  | S1M | VB10 | [12 empty] | NotDetermined
1938 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563253 | Headline | Cause |  | Contact of Patient with System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1939 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563254 | Requirement | Measure | um_SqueezingZones | User Manual - Crushing zones | CAUTIONCause: Movements of the C-arm during patient positioning.Possible Consequences: Crushing hazard for the patient!Remedy: Be careful around the hazardous locations indicated. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1940 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563255 | Headline | Function |  | System Performance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1941 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563256 | Headline | Hazard |  | Delayed/ no treatment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
1942 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563257 | Headline | Cause |  | Communication problems |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz07 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1943 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563259 | Headline | Cause |  | Degraded system performance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1944 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563263 | Headline | Function |  | Virus Protection |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1945 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563264 | Headline | Hazard |  | Wrong or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
1946 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563265 | Headline | Cause |  | Degraded system performance |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The system complies with applicable regulations and technical standards. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual advises that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
1947 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563266 | Requirement | Measure | verify_internet_access | Verify access to internet | There shall be only restricted access to internet possible for a clinical user (e.g. only some apphost related vendor URLs shall be allowed to access, but a normal internet usage shall not be possible).The service user shall be able to access internet via a special service environment | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | InPreReview | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1948 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563267 | Requirement | Measure | verify_BIOS_access | Verify access to BIOS | Access to BIOS needs to be password-protected. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 11.10.2022 17:41 | Bence David Fodor (EXT) |  |  |  | AppHost,FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1949 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563268 | Headline | Function |  | Documentation Safety Requirements for X-ray Generator |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1950 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563269 | Headline | Hazard |  | Hazards from fire or smoke |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Flammability |  | S1M | VB10 | [12 empty] | NotDetermined
1951 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563270 | Headline | Cause |  | Overheating |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz36 | See general justifications in the introduction which apply here. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | D | Med | D
1952 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563271 | Requirement | Measure | docCS_wrong_fuses_xgen | Service Documentation usage of wrong fuses | Service Safety Instruction CAUTIONCause: Usage of wrong fusesPossible Consequences: Risk of fireRemedy: Ensure the correct usage of fuses taking into account the label next to the fuses | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1953 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563272 | Headline | Measure |  | Reference to warning label next to fuses to use correct fuses | See HM_S1_hw_F2G1U2M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1954 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563273 | Headline | Measure |  | Reference to temperature monitoring of the ESU discharge resistors and inspection | See HM_S1_hw_F2G1U3M1 and HM_S1_hw_F2G1U3M2 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1955 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563274 | Headline | Hazard |  | Environmental hazard |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inadequate Disposal and scrapping |  | S1M | VB10 | [12 empty] | NotDetermined
1956 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563275 | Headline | Cause |  | Improper disposal of hazardous materials |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz38 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1957 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563276 | Requirement | Measure | docCS_disposal_esu | Improper disposal of ESU (ultra capacitor cells with harmful and flammable liquid) | Which parts must be removed prior to scrapping and how these parts are disposed in an environmental friendly manner are documented in the Disposal Instructions (Service Documents). CAUTION Cause: ESU contains harmful and flammable liquid (Acetonitrile) and is not disposed correctlyPossible Consequences: Risk of injury and environmental damage.Remedy: For proper disposal, send the ESU back to the manufacturer, Siemens Healthineers. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 27.06.2022 11:24 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1958 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563277 | Headline | Measure |  | Reference to material request form and label on ESU | HM_S1_serv_F2G2U1M1 and HM_S1_esu_F2G2U1M2 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1959 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563278 | Headline | Hazard |  | Electric shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High Voltage,Line voltage |  | S1M | VB10 | [12 empty] | NotDetermined
1960 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563279 | Headline | Cause |  | Voltage still present after switching-off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1961 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563280 | Requirement | Measure | docCS_voltage_presenence_xgen | Service Documentation - Voltage Presence | Service DocumentsWARNINGCause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting.Possible Consequences: Risk of death, serious physical injury or damage to property.Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task.Read and observe the warning labels attached next to conducting parts. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1962 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596998 | Requirement | Measure | docSCM_voltage_presenence_xgen | Factory Documentation - Voltage Presence | Factory Documents WARNING Cause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting.Possible Consequences: Risk of death, serious physical injury or damage to property.Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task.Read and observe the warning labels attached next to conducting parts. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 23.06.2022 | 23.06.2022 14:00 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1963 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563282 | Headline | Measure |  | Reference to discharge circuits | See HM_S1_hw_F6G1U4M1 in <XREF>578770</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1964 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563283 | Headline | Cause |  | ESU requires discharge after switch-off |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The system has tool-removable covers and is designed with appropriate MOOP / MOPP. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1965 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563284 | Requirement | Measure | docCS_esu_discharge | Service Documentation ESU discharge | Service Safety Instruction WARNING Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e. g. main unit covers are removed. Possible consequences: Risk of death or serious injury due to electric shock when electrified components are touched.Remedy: Discharge the ESU immediately after removing main unit covers before starting to work on relevant parts of the main unit, especially when removing or working on the ESU itself.If one of the fuses F21 or F20 is broken, the fast discharging procedure is not working. Verify that no voltage on the fuse holder is present before replacing broken fuses! | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1966 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563285 | Requirement | Measure | docSCM_esu_discharge | Supply Chain Documentation ESU discharge | Supply Chain Safety Instruction Warning: Cause: If the ESU is not discharged completely, electrical voltage is present. Live-line working can cause electric shock. Possible consequences: Risk of death, serious physical injury or damage to property. Remedy: Discharge the ESU prior to removing it. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1967 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563286 | Headline | Measure |  | Reference to ESU discharge circuit, display of voltage, warning labels | SeeHM_S1_hw_F6G1U1M2HM_S1_esu_F6G1U1M3HM_S1_hw_F6G1U1M4HM_S1_hw_F6G1U1M5HM_S1_hw_F6G1U1M6HM_S1_mec_F6G1U1M1,HM_S1_mec_F6G1U2M1,HM_S1_mec_F6G1U2M2,HM_S1_hw_F6G1U3M1,HM_S1_manu_F6G1U3M2in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1968 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580657 | Headline | Cause |  | Discharge unit defective, intermediate circuit not discharged completely |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 26.09.2022 17:01 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-A | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Service Personnel | S1I,S1F | VA10,VA20,VA30 | [12 empty] | NotDetermined | B | Med | B
1969 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580658 | Requirement | Measure | docCS_check_discharge_intermediate_circuit | Warning - Service Documentation - Check discharge of intermediate circuit | Service Instruction WARNING Cause: If the intermediate circuit is not discharged completely, electrical voltage is present. Live-line working can cause electric shock. Possible consequences: Risk of death, serious injury or damage to property.Remedy: Check the intermediate circuit with suitable voltmeter on connector board M100. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 15.06.2022 16:17 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1970 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 580659 | Requirement | Measure | docSCM_check_discharge_intermediate_circuit | Warning - Factory Documentation - Check discharge of intermediate circuit | Factory Documentation WARNING Cause: If the intermediate circuit is not discharged completely, electrical voltage is present. Live-line working can cause electric shock.Possible consequences: Risk of death, serious injury or damage to property.Remedy: Check the intermediate circuit with suitable voltmeter on connector board M100. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 05.07.2021 | 15.06.2022 16:18 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1971 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563287 | Headline | Hazard |  | Oil leaks out of the monoblock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Infection,Not sterile,Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
1972 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563288 | Headline | Cause |  | Component error |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32 #Haz16 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,System | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1973 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563289 | Requirement | Measure | docCS_leakage_monobloc | Check leakage of monobloc | Service DocumentationCAUTIONCause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil)Possible consequences: Functional disturbances, leaking of liquids from the systemRemedy: Check the monoblock and the cooling unit for leakage of liquids during service activities and replace defective parts. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1974 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563290 | Headline | Measure |  | Reference to measures against leakage of oil | HM_S1_manu_F3G1U1M1, HM_S1_hw_F3G1U1M2, HM_S1_hw_F3G1U1M3, HM_S1_manu_F3G1U1M4, HM_S1_hw_F3G1U2M1, HM_S1_mec_F3G1U3M1 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1975 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 596995 | Requirement | Measure | docSCM_leakage_monobloc | Check monobloc leakage | Factory Documentation CAUTION Cause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil)Possible consequences: Functional disturbances, leaking of liquids from the systemRemedy: Check the monoblock and the cooling unit for leakage of liquids during service activities and replace defective parts. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 23.06.2022 | 23.06.2022 13:56 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1976 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563291 | Headline | Hazard |  | Leaking cooling liquid |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Infection,Not sterile,Loss or deterioration of function |  | S1M | VB10 | [12 empty] | NotDetermined
1977 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563292 | Headline | Cause |  | Component error or service error (Refill of Cooling Unit with connected Compensation Bag) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz32 #Haz16 | See general justifications in the introduction which apply here. The service documentation warns how to refill the cooling unit correctly. The technicians are trained. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,System | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1978 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563293 | Requirement | Measure | docCS_refilling_cooling_unit | Service Documentation - Disconnect Compensation Bag | CAUTIONCause: If the compensation bag is not disconnected from the cooling unit, the compensation bag can lacerate.Possible Consequences: Risk of burning or infection.Remedy: Prior to refilling the cooling unit, disconnect the compensation bag. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 22.06.2022 16:06 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1979 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563294 | Headline | Measure |  | Reference to measures against leakage of cooling liquid | See HM_S1_mec_F3G1U4M1, HM_S1_manu_F3G1U4M2, HM_S1_manu_F3G1U4M3, HM_S1_serv_F3G1U4M4, HM_S1_hw_F3G1U4M5 in <XREF>578770</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1980 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563295 | Headline | Hazard |  | Unnecessary radiation |  | ToBeReviewed | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation,Incorrect or inappropriate output | Patient,System | S1M | VB10 | [12 empty] | NotDetermined
1981 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563296 | Headline | Cause |  | Missing error signal about image loss |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz51 #Haz41#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1982 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563297 | Requirement | Measure | missing_HV_ON_signal | Missing HV_ON signal | The system must ensure that missing HV_ON signal shall not lead to a patient harm.[See also HM_S1_hw_F5G3U1M1 in <XREF>578770</XREF>] | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 16.09.2022 12:44 | Javier Lopez Leal |  |  |  | FLC |  |  |  | Basic | System Integration Test | none | n.a. | Not Determined | Not determined | Not determined | NotForTC | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Ionizing radiation |  | S1M | VB10 | [12 empty] | NotDetermined
1983 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563298 | Headline | Function |  | Documentation user manual for Imaging System |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1984 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563299 | Headline | Hazard |  | Risk of burnings |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
1985 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563300 | Headline | Cause |  | Too high temperature of the detector |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz14 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System,Environment | S1M | VB10 | [12 empty] | NotDetermined | D | Med | C
1986 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563301 | Requirement | Measure | docUser_to_high_detector_temp | User Manual - Detector Temperature | CAUTIONCause: Detector temperature is not within recommended range.Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of X-ray radiation exposure without any diagnostic value!Remedy: Regularly check the temperature icon on the display. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1987 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563302 | Headline | Measure |  | Reference to System Warning - Temperature of the Detector | <XREF>562226</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | none | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1988 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563303 | Headline | Hazard |  | Malfunction of imaging system during operation due to defective HW |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1989 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563304 | Headline | Cause |  | Defective FLC-Hardware |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:41 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz50 #Haz40 #Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1990 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563305 | Requirement | Measure | docUser_defective_FLC_HW | User Manual - Defective FLC HW | CAUTIONCause: Defective imaging system hardwarePossible consequences: Risk of X-ray radiation without a stored imageRemedy: In case of image failures stop exposures and switch off the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1991 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563306 | Headline | Hazard |  | Risk of incorrect interventional decision caused by wrong image processing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1992 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563307 | Headline | Cause |  | Wrong LUT |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20 #Haz41 #Haz51 #Haz31 | See general justifications in the introduction which apply here. The operator manual advises to pick the correct imaging parameters and settings and to perform regular checks. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
1993 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563308 | Requirement | Measure | docUser_wrong_LUT | User Manual - Wrong LUT | CAUTIONCause: Wrong or extreme image processing parameters.Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of incorrect decision during intervention or artifacts.Remedy:- Apply image processing parameters carefully, such as edge enhancement and windowing, and always check them carefully.- Always check the image quality before reaching an interventional decision. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1994 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563309 | Headline | Hazard |  | Delayed or no interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data acquisition, evaluation and display,data transfer,Data and Image processing |  | S1M | VB10 | [12 empty] | NotDetermined
1995 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563310 | Headline | Cause |  | Plugging/Unplugging USB devices interfere with performance of imaging system |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The operator manual advises to plug and unplug USB devices not during critical tasks. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
1996 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563311 | Requirement | Measure | docUser_unplugg_USB | User Manual - Plug/Unplug USB Devices | CAUTIONCause: Risk of connecting/disconnecting USB devices during the exposure.Possible consequences: System instability, impact on or disturbance of running acquisition tasks, loss of data.Remedy: Do not insert or disconnect USB devices during exposure or other critical processes. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1997 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563312 | Headline | Cause |  | Unauthorized installation of SW/HW |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12#Haz42 #Haz52 #Haz01 | See general justifications in the introduction which apply here. The operator manual warns to not perform non-approved modifications of the system, incl. combining with non-approved accessories or other devices. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
1998 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563313 | Headline | Measure |  | Reference to User Manual - Unauthorized Software | <XREF>562152</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
1999 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563314 | Headline | Cause |  | Network failure |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz50#Haz30 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-C | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | E
2000 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563315 | Requirement | Measure | docUser_network_failure | User Manual - Network Failure | CAUTIONCause: Network connection may not be available continuously.Possible consequences: Reference images not available during intervention.Remedy: Before beginning the intervention, the user should load the required images from the network archive. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2001 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563316 | Headline | Cause |  | Software lockup |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The system software is developed according to SW development process. The system provides warning messages in the event of specific system malfunctions. The operator manual warns that the functionality of the system shall be checked before starting a procedure and that procedures shall be in place in case of a system failure. In case of a recoverable malfunction, the operator manual advises the to restart the system. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
2002 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563317 | Requirement | Measure | docUser_software_lockup | User Manual - Software Lockup | CAUTIONCause: Software crash.Possible consequences: Working with the imaging system is no longer possible.Remedy: Restart the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2003 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563318 | Headline | Cause |  | Restart missing after Service session |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The service documentation advises how to perform service activities including updates correctly. The technicians are trained. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient,User,System | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
2004 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563319 | Requirement | Measure | docUser_restart_after_service | User Manual - Reboot after Service | CAUTIONCause: No restart after service session.Possible consequences: Changes to the service environment may not go into effect without a system restart.Remedy: Restart the system after a service session. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2005 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563320 | Requirement | Measure | FLC_Restart_Service | Message to request reboot after service | A message is displayed to restart the system when the Service environment is left. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 14.08.2021 11:06 | Yesin Gasmi (EXT) |  |  |  | FLC |  |  |  | Basic | Subsystem Integration Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2006 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563321 | Headline | Hazard |  | Danger of incorrect interventional decision |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Assignment of image, data, objects |  | S1M | VB10 | [12 empty] | NotDetermined
2007 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563322 | Headline | Cause |  | Assignment of series/ scenes/ images to the wrong patient |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz26#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
2008 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563323 | Requirement | Measure | docUser_wrong_patient_assignment | User Manual - Mix-up of Patients | CAUTIONCause: Risk of mistaking patients.Possible consequences: The wrong patient is being treated.Remedy: The user must ensure the correct patient is being treated before he or she is registered on the system. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2009 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563324 | Headline | Cause |  | Scale displayed for non-calibrated images |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz20#Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
2010 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563325 | Headline | Measure | docUser_wrong_calibration_objects | Reference to Caution for 2D measurements | <XREF>562590</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2011 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563326 | Headline | Hazard |  | loss of images or patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
2012 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563327 | Headline | Cause |  | Postponed loss of recorded patient data |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz50#Haz30 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
2013 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563329 | Requirement | Measure | docUser_loss_of_data_USB | User Manual - Loss of Data USB | CAUTIONCause: Disconnecting USB devices during ongoing data transfer.Possible consequences: Loss of data and damage to the data medium possible.Remedy: - Do not disconnect the USB device during active data transfer (such as export).- Use removable devices as transfer media only. Removable devices are not permitted for archiving purposes. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2014 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563331 | Headline | Cause |  | Harddisk or file system defect |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz21 #Haz50 | See general justifications in the introduction which apply here. The operator manual warns to take care of archiving images as appropriate. | Basic | n.a. | none | n.a. | Not Determined | 1-C | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
2015 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563332 | Requirement | Measure | docUser_hard_disk_no_archive | User Manual - Hard Disk Failure | CAUTIONCause: Use the imaging system hard disk as a long-term archive.Possible consequences: Loss of image dataRemedy: Do not use the imaging system hard disk as a long-term archive.Follow the local regulations regarding the archiving of X-ray images, while taking into account the use of DICOM archive nodes. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2016 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563333 | Headline | Hazard |  | loss of image quality |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | data transfer,Loss of data |  | S1M | VB10 | [12 empty] | NotDetermined
2017 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563334 | Headline | Cause |  | Incorrect calibration data used |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz23 #Haz41 | See general justifications in the introduction which apply here. The service documentation advises how to perform the FD calibration correctly. The technicians are trained. | Basic | n.a. | none | n.a. | Not Determined | 2-C | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | E | Low | D
2018 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563335 | Requirement | Measure | docServ_regular_calibration | Note in service documentation | Note in service documentation:Correct FD calibration in regular intervals | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2019 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563336 | Headline | Function |  | User Manual for Touch Interfaces (xCP) |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2020 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563337 | Headline | Hazard |  | Electric Shock |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Line voltage,Flammability |  | S1M | VB10 | [12 empty] | NotDetermined
2021 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563338 | Headline | Cause |  | Intrusion of Liquids and Instruments |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz12 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | System Test | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | User,Service Personnel,System | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
2022 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563339 | Headline | Measure |  | Reference - User Manual - Insufficient cleaning | <XREF>562240</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2023 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563340 | Headline | Measure |  | Reference - Spillage of Liquids on System | <XREF>562549</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2024 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563341 | Requirement | Measure | docUser_xCP_intrusion_instruments | User Manual - Intrusion of Instruments xCP | WARNING:Cause: Instruments falling into the gap between the control unit and its housing.Possible Consequences: Short circuit, electric shock, and/or fire!Remedy: Make sure no instruments, conductive material, or fluids penetrate into the openings/joints between the control unit and its housing. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2025 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563342 | Headline | Hazard |  | Overheating |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | High temperature |  | S1M | VB10 | [12 empty] | NotDetermined
2026 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563343 | Headline | Cause |  | Inappropriate Draping |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. The operator manual advises the user to drape the system appropriately and to clean the system correctly. | Basic | System Test | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | System | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
2027 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563344 | Requirement | Measure | docUser_RCP_specified_drapes | User Manual - Use Specified Drapes Only | CAUTIONCause: Remote control unit covered by unsuitable covers.Possible Consequences: Overheating of the remote control unit!Remedy: Only use the sterile covers specified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2028 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563345 | Headline | Cause |  | Hot Surface of RCP |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz13 | See general justifications in the introduction which apply here. The Operator Manual advises the operator how to drape the system correctly and to prevent long-term contact between the patient and heated surfaces. It is not practical to implement additional control measures to further reduce the probability of the hazardous situation. | Basic | System Test | none | n.a. | Not Determined | 1-B | 1-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
2029 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563346 | Requirement | Measure | docUser_RCP_avoid_patient_contact | User Manual - Avoid Patient Contact with RCP | CAUTIONCause: The surfaces of the remote control unit become hot.Possible Consequences: Risk of burns!Remedy:- Make sure that the patient does not touch the remote control unit for more than 10 minutes or that after 10 minutes can no longer come in contact with it.- As a rule, contact between the patient and the remote control unit should be avoided. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2030 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563347 | Headline | Measure |  | User Manual - Positioning of RCP | <XREF>562487</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2031 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563348 | Headline | Measure |  | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF> | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:38 | Johannes Schmidt | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2032 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563349 | Headline | Hazard |  | Risk of injury |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Test | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Inappropriate accessory |  | S1M | VB10 | [12 empty] | NotDetermined
2033 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563350 | Headline | Cause |  | Falling Down of RCP |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz55#Haz32#Haz41#Haz19 | See general justifications in the introduction which apply here. The operator manual warns to take care when attaching and detaching parts to and from the system. | Basic | System Test | none | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | D
2034 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563351 | Requirement | Measure | docUser_RCP_fixation_RCP_holder | User Manual - Falling Down | CAUTIONCause: Remote control unit falls down.Possible Consequences: Risk of injury/crushing!Remedy: It is recommended to keep the remote control unit in the holder provided. If a holder is not available, place the remote control unit on the instrument table. | Approved | Unassigned | Klaus-Peter Fichtner (EXT) | Klaus-Peter | Fichtner (EXT) | Judith Welzel |  |  | 29.07.2020 | 08.03.2022 09:17 | Klaus-Peter Fichtner (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | User Manual | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VxFUTURE | [12 empty] | NotDetermined
2035 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563352 | Headline | Chapter |  | Accessories and combinations | Even if a system is considered safe, it might cause new or additional hazards and risks when used together with an accessory or in combination with other devices.Risks related to accessories and combinations are covered in above functional analysis, e. g. in <XREF>562990</XREF>, <XREF>563078</XREF>, <XREF>562474</XREF>, <XREF>562478</XREF>, <XREF>562482</XREF>, <XREF>562624</XREF>, <XREF>562322</XREF>, <XREF>562660</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2036 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563353 | Headline | Chapter |  | Processes employed | Processes employed throughout the product lifetime like design, purchasing, production, modification, packing, transport, installation, service, refurbishing and disposal can be the source of failures or hazards when not carried out appropriately. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2037 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563354 | Headline | Function |  | Refurbishment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2038 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563355 | Headline | Hazard |  | Biological contamination or infection during refurbishment activities |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Insufficient control of subcontractors,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2039 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563356 | Headline | Cause |  | Contact with contaminated system parts during deinstallation or cleaning for refurbishment |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33 | See general justifications in the introduction which apply here. | Basic | n.a. | none | n.a. | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
2040 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563357 | Requirement | Measure | infection_refurbishment | Measures to prevent infection due to contaminated system parts - refurbishment | Warning in instruction for refurbishment: Dienstleisterspezifikation <XREF>561903</XREF>. For text, see <XREF>563207</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | System Integration Test | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined |  |  | E
2041 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563358 | Headline | Function |  | Service / maintenance | Risks related to service / maintenance are covered in above functional analysis, e. g. in <XREF>562185</XREF>, <XREF>562202</XREF>, <XREF>562210</XREF>, <XREF>562245</XREF>, <XREF>562954</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2042 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563359 | Headline | Hazard |  | Biological contamination or infection during service activities |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2043 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563360 | Headline | Cause |  | Contact with contaminated system parts during service activities |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz33 | See general justifications in the introduction which apply here. | Basic | n.a. | none | Factory Documentation | Not Determined | 3-B | 3-A | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Bacteria, viruses,Infection | Patient,User,Service Personnel | S1M | VB10 | [12 empty] | NotDetermined | C | Med | B
2044 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563361 | Headline | Measure | infection_service | Reference to measures to prevent infection due to contaminated system parts - service | <XREF>563207</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2045 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563362 | Headline | Hazard |  | Undetected damage of the cooling unit |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2046 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563363 | Headline | Cause |  | Overpressure test of cooling unit assemly not performed correctly |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 26.09.2022 16:51 | Klaus-Peter Fichtner (EXT) | [5 empty] | #Haz41#Haz51#Haz31 | See general justifications in the introduction which apply here. The service documentation contains the instruction how to perform the cooling unit check correctly. | Basic | n.a. | EXPORT | n.a. | Not Determined | 2-B | 2-B | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Deterioration in function | Patient | S1M | VB10 | [12 empty] | NotDetermined | D | Low | C
2047 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563364 | Requirement | Measure | serv_cooling_assembly_check | Service Documentation - Overpressure test | Service DocumentationCAUTIONCause: Overpressure test is performed with cooling unit components and hose line including the flow indicator while connected to the system.Possible consequences: Undetected damage of the cooling unit components.Remedy: Make sure that the cooling unit assembly and the hose line including the flow indicator are disconnected from the system during overpressure test. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 15.06.2022 17:55 | Johannes Schmidt | [7 empty] | Basic | System Integration Test | EXPORT | Service Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | Yes | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2048 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563365 | Headline | Function |  | Manufacturing |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2049 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578701 | Headline | Hazard |  | None | No additional risks besides the risks related to manufacturing are covered in above functional analysis, e. g. in <XREF>562028</XREF>, <XREF>562032</XREF>, <XREF>562094</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2050 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578702 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2051 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578703 | Headline | Measure | infection_service | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2052 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563366 | Headline | Function |  | Supply chain management process |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2053 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578704 | Headline | Hazard |  | None | No risk on system level has been identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2054 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578705 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2055 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578706 | Headline | Measure | infection_service | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2056 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563367 | Headline | Function |  | Packaging and transport |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2057 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578707 | Headline | Hazard |  | None | No risk on system level has been identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2058 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578708 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2059 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578709 | Headline | Measure | infection_service | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2060 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563368 | Headline | Function |  | Disposal |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2061 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578710 | Headline | Hazard |  | None | No additional risks besides the risks related to disposal are covered in above functional analysis, e. g. in <XREF>562010</XREF>, <XREF>563134</XREF>, <XREF>562094</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2062 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578711 | Headline | Cause |  | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None | Insufficient control of processes,Insufficient training of personnel,Bacteria, viruses |  | S1M | VB10 | [12 empty] | NotDetermined
2063 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 578712 | Headline | Measure | infection_service | None |  | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 22.02.2021 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | EXPORT | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2064 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563370 | Headline | Chapter |  | Usability aspects | Inappropriate design of the user interface as well as adverse working situations can lead to intended or unintended use errors, causing hazards to persons and goods. E.g. risk of having certain parts of the user manual or user interface not in the users’ language.Risks related to usability are covered in above functional analysis, e. g. in <XREF>562462 </XREF>, <XREF>562202</XREF>, <XREF>562347</XREF>, <XREF>563480</XREF>.All usability relevant risk control measures are tagged with "Usability: Yes". | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2065 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563369 | Headline | Chapter |  | Warnings and cautions from supplied components | Safety information from the operator manual of suppliers' components, which are integrated into the medical system, has to be checked for applicability and adoption into the system operator manual.Risks related to supplied components are covered in above functional analysis. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 17.06.2022 11:40 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2066 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563371 | Headline | Chapter |  | Hazards to environment | Environmentally relevant risks, such as from emissions under normal use, radiation generation or leakage, or caused by the use of critical materials, or risks occurring during end of life treatment shall be considered in the scope of risk analysis. This risk analysis shall consider at least:• Material supply – starting with the obtaining of natural resources and ending with the supply of materials at the production site,• Production – covers manufacturing, including transportation and ends at handover of the product to the customer• Use – covers the use of the product including servicing• End of life – starts with the permanent taking out of service and de-installation of the product at the customer and ends with the supply of secondary material (material recycling), thermal recycling or disposal of non-recyclable substances. Risks related to the environment are covered in above functional analysis, e. g. in <XREF>563134</XREF>, <XREF>562094</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2067 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563372 | Headline | Chapter |  | Biological and chemical hazards | The devices must be designed and manufactured in such a way as to reduce to a minimum the risks posed by substances leaking from the device. Special attention shall be given to substances which are carcinogenic, mutagenic or toxic to reproduction (MDD).Biological and chemical risks are covered in above functional analysis, e. g. in <XREF>562010</XREF>, <XREF>563134</XREF>, <XREF>562094</XREF>. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:05 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2068 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563373 | Headline | Chapter |  | Hazards through outsourcing of processes | It must be evaluated, whether the outsourcing of the defined processes will introduce hazards to the product. These hazards can be related to e.g. inappropriate controlling of the supplier of outsourced processes.The wear and tear related hazards are managed by component design and their integration into the system according to the requirements about "Service Life" and the worst-case conditions within the scope of the intended use. The system level verification provides a confirmation for the fulfillment of wear and tear related requirements. Thereafter it is monitored throughout the entire product lifetime cycle by means of the CAPA process. Wear and tear related, potentially hazardous component failures are addressed in the system risk analysis in sections "Evaluation of Functions, Hazards, Risks and Measures".No risk on system level has been identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2069 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563374 | Headline | Chapter |  | Hazards by building services | It must be evaluated, whether effects from inappropriate building management services will introduce hazards to the manufacturing of the product. These effects may result through: e.g. temperature, pressure, humidity, pest infestation.No risk on system level has been identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2070 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 563376 | Headline | Chapter |  | Hazards by environmental conditions during usage | It must be evaluated, whether the environmental conditions (as they are defined) will introduce hazards to the use of the product. These conditions are: temperature, pressure and humidity.No risk on system level has been identified. | Approved | Unassigned | Judith Welzel | Judith | Welzel | Judith Welzel |  |  | 29.07.2020 | 07.08.2021 15:04 | Yesin Gasmi (EXT) | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2071 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 595404 | Headline | Chapter |  | Hazard by to visual electromagnetic radiation | The system contains light sources (according to IEC 60601-1:2005 Clause 10.5 visual electromagnetic radiation) such as LEDs on UI devices like hand controls and control panels which are classified in the exempt group (Lamp does not pose any photobiological hazard) according to IEC 62471:2006. This is addressed in 06S434342 in 11108300-EPH-801. Therefore, no risk on system level has been identified. | Approved | Unassigned | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 23.03.2022 | 27.06.2022 14:08 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | n.a. | Not Determined | empty | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined
2072 | PS_S1 | Current Baseline | 06H_Risk_Analysis (06H) | 597012 | Headline | n.a. |  | Biocompatibility evaluation | Introduction IEC 60601-1, cl 11.7 and specifically ISO 10993-1 describe the general principles governing the biological evaluation of medical devices within a risk management framework, including the general categorization of devices based on the nature and duration of their contact with the body.The range of possible chemical and biological hazards is wide and can include short-term effects such as acute toxicity, irritation to the skin, eye and mucosal surfaces, haemolysis and thrombogenicity, as well as long-term or specific toxic effects such as subchronic and chronic toxic effects, sensitization resulting in allergy, genotoxicity, carcinogenicity (tumorigenicity) and effects on reproduction or development, including teratogenicity.Essential for the risk assessment of these hazards is to determine how critical are substances within the components and how likely it is that these substances are transferred to people to trigger a response.Note, the following evaluation is based on <XREF>347633</XREF>, specifically the guidance for biological evaluation of medical devices in acc. with ISO 10993-1. Overview and classification The <XREF>276902</XREF> describes the "nature and duration of body contact" for patients and users with the system. It is rated as a non-contacting medical device, because there is no direct contact between patient and the system required to fulfill the intended use. However, system contact is possible within the following situations:1. "Intended" short term (only seconds or minutes) contact (predominantly with the hand) with intact skin of qualified personnel (Manufacturing, Installation, Service, etc.) or the clinical personal at system control covers or control elements.2. Unintended (e.g. touched accidentally during a procedure) short term (only seconds or minutes) contact at the C-arm area with intact or non-intact skin of the patient. The CIOS systems are out of scope of ISO 10993-1 with the following rationales: Rationale and details for system operators System operators get in contact with user interfaces such as handles, control elements, PC-accessories or covers. These users are qualified manufacturing, installation, service or clinical personnel which are expected to have intact skin at their hands or are wearing gloves during work. Also the contact duration is very short-term considering the duration of contact categories in ISO 10993-1. Thus no detailed biological evaluation is needed for these components. However, in general it shall be shown that the materials of these components are in common use for other consumer products with a similar nature of contact, see <XREF>597019</XREF>.Furthermore, all the components are fulfilling applicable regulations (RoHS, REACH, etc.) which are restricting the use of substances of high concerns in medical devices and products. Rationale and details for patients The mobile C-arm systems are rated as non-contacting medical devices for patients. No direct or indirect contact is required to fulfill the intended use of the system. However, short-time contact is possible between patient skin (intact and in rate cases non-intact) and the C-arm area (C-Arm, detector cover and monobloc cover). Also it cannot be fully excluded that there might be a leakage of bodily fluids of the patient onto the above described system area. Therefore, this area specifically needs to be covered by sterile or sterilizable covers, see <XREF>562240</XREF>, effectively protecting the system and patient. Since there is no reasonably assumable contact between the patient and other parts of the device no further biological evaluation is needed. | Approved | Medium | Johannes Schmidt | Johannes | Schmidt | Johannes Schmidt |  |  | 27.06.2022 | 27.06.2022 18:09 | Johannes Schmidt | [7 empty] | Basic | n.a. | none | Factory Documentation | Not Determined | Not determined | Not determined | TBD | NotDetermined | NotDetermined | NotDetermined | NotDetermined | No | None |  |  | S1M | VB10 | [12 empty] | NotDetermined