## Sheet1
| Chapter Sec | Chapter ID | Chapter Type | Chapter Requirement | Function Sec | Function ID | Function Type | Function Requirement | Hazard Sec | Hazard ID | Hazard Type | Hazard Requirement | Cause Sec | Cause ID | Cause Type | Cause Requirement | Measure Sec | Measure ID | Measure Type | Measure Requirement | Measure Description
0 | 4.1 | 563430 | Chapter | Move trolley | 4.1.1 | 569143 | Function | Move trolley | 4.1.1.1 | 561941 | Hazard | Crushing due to trolley tipping over | 4.1.1.1.1 | 561942 | Cause | Movement over inclined surface or monitor movement changes the center of gravity of the trolley | 4.1.1.1.1.1 | 561943 | Measure | Trolley design prevents overbalance | The trolley must be designed in a way that prevents tipping over when placed in any transport position on an inclined plane at an angle of 10 Â°.
1 | 4.1 | 563430 | Chapter | Move trolley | 4.1.1 | 569143 | Function | Move trolley | 4.1.1.2 | 561944 | Hazard | Crushing through tilting of the device | 4.1.1.2.1 | 561947 | Cause | Increase of tilt moment due to additional weight | 4.1.1.2.1.1 | 561949 | Measure | Design of the trolley | The trolley must be designed in a way that prevents tilting of the trolley when driving it over an inclined plane (i.e. ramp). Â Remark:Compliance with DIN EN 60601-1 tilt test
2 | 4.1 | 563430 | Chapter | Move trolley | 4.1.1 | 569143 | Function | Move trolley | 4.1.1.2 | 561944 | Hazard | Crushing through tilting of the device | 4.1.1.2.2 | 561950 | Cause | Moving the trolley over a doorstep, against a doorframe, wall or other equipment | 4.1.1.2.2.1 | 561951 | Measure | Rough handling test trolley | The trolley shall be designed to comply with IEC 60601-1Â cl. 15. 3. 5. Assumed speed of the trolley: 0,8 m/s +/- 0,1 m/s
3 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.1 | 564256 | Cause | Short circuit of the battery (external) | 4.3.1.1.1.1 | 564257 | Measure | Coded battery connector | The battery connector for charging shall be coded in order to prevent connection with wrong polarity.
4 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.1 | 564256 | Cause | Short circuit of the battery (external) | 4.3.1.1.1.2 | 564258 | Measure | Battery housing preventing accidental short circuit | The housing of the battery shall prevent accidental short circuit.
5 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.1 | 564256 | Cause | Short circuit of the battery (external) | 4.3.1.1.1.3 | 564259 | Measure | Wiring check battery in manufacturing | The battery wiring shall be checked during manufacturing.
6 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.2 | 564267 | Cause | Overcharging of the battery (voltage, current) | 4.3.1.1.2.1 | 564260 | Measure | Prevent overcharging of battery | Overcharging of the battery shall be prevented regarding voltage and current.
7 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.3 | 571858 | Cause | Overcharging of the battery, sleeping error in P path | 4.3.1.1.3.1 | 571859 | Measure | Self-test of battery overcharging protection | Self-test at startup of battery overcharging protection
8 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.4 | 564268 | Cause | Internal short circuits and temperature increase of the battery packs | 4.3.1.1.4.1 | 564261 | Measure | Prevent charging of battery with high temperature | The temperature of the battery unit shall be monitored. When a critical temperature is reached, the charging of the battery shall be blocked.
9 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.4 | 564268 | Cause | Internal short circuits and temperature increase of the battery packs | 4.3.1.1.4.2 | 577242 | Measure | Ventilation battery housing | The battery pack housing shall beÂ ventilated in such a way that in case of gas emission the accumulation of gases is prevented. Note: Gas emission only in error condition, not normal condition. Ventilation prevents accumulation of gases and possible ignition.
10 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.5 | 571860 | Cause | Internal short circuits and temperature increase of the battery packs, sleeping error in P path | 4.3.1.1.5.1 | 571861 | Measure | Self-test of protection against charging of battery with high temperature | Self-test at startup of battery temperature protection
11 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.6 | 564269 | Cause | Mechanical destruction | 4.3.1.1.6.1 | 564262 | Measure | Battery housing mechanical strength | The battery shall be provided with a housing to protect the battery from external mechanical impact.
12 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.6 | 564269 | Cause | Mechanical destruction | 4.3.1.1.6.2 | 564263 | Measure | Battery discharge current limitation | The discharge current of the battery shall be limited (prevent sudden short circuit discharge, e. g. fuse at battery outlet).
13 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.7 | 564270 | Cause | Thermal runaway of battery pack (internal short circuit, overheating) | 4.3.1.1.7.1 | 564265 | Measure | Battery housing fireproof and preventing flame break-out | The battery shall be provided with a fireproof housing (thermal runaway), preventing spread of flames.
14 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.1 | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 4.3.1.1.7 | 564270 | Cause | Thermal runaway of battery pack (internal short circuit, overheating) | 4.3.1.1.7.2 | 564264 | Measure | Thermal isolation of single battery cells | Thermal isolation of single battery cells against spread of thermal runaway chain reaction in other battery cells.
15 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.2 | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 4.3.1.2.1 | 577236 | Cause | Increased temperature of the battery packs | 4.3.1.2.1.1 | 577237 | Measure | Reference - Prevent charging of battery with high temperature | <XREF>564261</XREF>
16 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.2 | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 4.3.1.2.1 | 577236 | Cause | Increased temperature of the battery packs | 4.3.1.2.1.2 | 577241 | Measure | Reference - Prevent overcharging of battery | <XREF>564260</XREF>
17 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.2 | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 4.3.1.2.1 | 577236 | Cause | Increased temperature of the battery packs | 4.3.1.2.1.3 | 577238 | Measure | Reference - Battery housing mechanical strength | <XREF>564262</XREF>
18 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.2 | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 4.3.1.2.1 | 577236 | Cause | Increased temperature of the battery packs | 4.3.1.2.1.4 | 577239 | Measure | Reference - Battery discharge current limitation | <XREF>564263</XREF>
19 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.2 | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 4.3.1.2.1 | 577236 | Cause | Increased temperature of the battery packs | 4.3.1.2.1.5 | 577240 | Measure | Reference - Thermal isolation of single battery cells | <XREF>564264</XREF>
20 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.3 | 569265 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.3.1.3.1 | 569264 | Cause | Activation of control elements by non-authorized persons (e. g. passer-by in the corridor) | 4.3.1.3.1.1 | 564347 | Measure | Lock movement initiation from chassis handle | The system shall provide the possibility to lock all motorized chassis movements activated via chassis handle by pressing a dedicated hardware button at the handle panel.
21 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.3 | 569265 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.3.1.3.1 | 569264 | Cause | Activation of control elements by non-authorized persons (e. g. passer-by in the corridor) | 4.3.1.3.1.2 | 569266 | Measure | Caution - Lock chassis movement or switch off | CAUTIONÂ Cause: System is left unattended (e.g. in standalone mode). Possible consequences: Unauthorized operation of system movements. Remedy: Do not leave the systemÂ unattended when turned on. Use the lock movement function at the chassis handle panel or switch off the system.
22 | 4.3 | 563431 | Chapter | Move main unit chassis in standalone mode | 4.3.1 | 564254 | Function | Move main unit chassis in standalone mode | 4.3.1.4 | 569291 | Hazard | Severe limitation of movement functionality (HS_111) | 4.3.1.4.1 | 569292 | Cause | Chassis battery runs empty, no more motorized movements possible | 4.3.1.4.1.1 | 568840 | Measure | Indicate battery charge for standalone mode | The system shall indicate the battery charge for standalone mode.
23 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.1 | 564116 | Measure | Collision avoidance in fast drive mode | The system shall detect the approximation of solid obstacles at critical collision areas of the system duringÂ fast driveÂ mode (e.g. towards moving direction), e.g. by means of a distance sensor. Â When an approximation is detected, the system shall reduce the speed to OR (operating room) mode speed before an imminent collision takes place. Â During normal operation (e.g. stopping without collision / approximation, emergency stop activation without collision / approximation), the system shall not create false positive collision / approximation detection information. Â The functionality of the sensor shall be tested during boot and monitored during operation (e.g. checking and calibrating by regularly scanning known fixed mounted objects with differing fixed distances to the sensor). Â When an error in the sensor system is detected, the system shall go to operating room speed (deactivate fast drive mode) and indicate the error to the operator.
24 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.2 | 562463 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>Â and to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
25 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.3 | 569215 | Measure | Reference - Set acceleration values | The acceleration of motorized movements under normal conditions is defined inÂ <XREF>569146</XREF>. For maneuverability of the system, the set speed values are deemed to be sufficient. The acceleration (and deceleration)Â is however monitored for mechanical stability <XREF>563680</XREF>.
26 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.4 | 564118 | Measure | Additional restrictions for fast drive mode | The max. fast drive speed (set speed value) shall only be allowed in forward and backward direction under consideration of the speed limits <XREF>563486</XREF> andÂ permissible centrifugal acceleration <XREF>563681</XREF>, and when initiated from the chassis force handle. Furthermore,Â fast drive mode shall only be allowed if the footswitchÂ is disconnected. Â When the system is in fast drive mode, and movements are initiated from other movement controls than from the chassis handle, the system shall apply max. OR mode speed (set speed).
27 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.5 | 564119 | Measure | Timeout for fast drive mode | AfterÂ at most 300 seconds of chassis standstill in fast drive mode, there shall be a timeout deactivating the fast drive mode automatically.
28 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.6 | 564120 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
29 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.7 | 564121 | Measure | Reference - System speed limits | <XREF>563486</XREF>
30 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.8 | 564122 | Measure | Reference - Collision detection | <XREF>563490</XREF>
31 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.9 | 564123 | Measure | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF>
32 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.10 | 564124 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
33 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.1 | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 4.4.1.1.1.11 | 564125 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
34 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.2 | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 4.4.1.1.2.1 | 571908 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
35 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.2 | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 4.4.1.1.2.2 | 578162 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
36 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.2 | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 4.4.1.1.2.3 | 571909 | Measure | Reference - System speed limits | <XREF>563486</XREF>The protection path knows when transport mode is activated.
37 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.2 | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 4.4.1.1.2.4 | 571910 | Measure | Reference - Additional restrictions for transport mode | <XREF>564118</XREF>Only chassis handle movements are allowed to initiate transport speed movements.
38 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.3 | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 4.4.1.1.3.1 | 571914 | Measure | Fast drive button protected against unintended activation | The button to activate the fast drive mode shall be protected against unintended activation, e.g. recessed button.
39 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.3 | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 4.4.1.1.3.2 | 571912 | Measure | Reference - Timeout for transport mode | <XREF>564119</XREF>
40 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.3 | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 4.4.1.1.3.3 | 571913 | Measure | Indicate active fast drive mode | The system shall have an optical indication which shows that the fast drive mode is active.
41 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.4 | 578308 | Cause | Unintentional activation of fast drive mode by technical error | 4.4.1.1.4.1 | 578309 | Measure | Safe activation fast drive mode | The fast drive mode shall be activated by two redundant paths. Â Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design contraints for single fault safety, see <XREF>578311</XREF>.
42 | 4.4 | 564113 | Chapter | Move main unit chassis in transport mode | 4.4.1 | 564112 | Function | Move main unit chassis in transport mode | 4.4.1.1 | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.4.1.1.4 | 578308 | Cause | Unintentional activation of fast drive mode by technical error | 4.4.1.1.4.2 | 571915 | Measure | Check fast drive mode button state | The system shall check the status of both fast drive button signal paths during boot-up. Â When an active fast drive button signal path is detected, the fast drive mode shall be blocked and an error message shall be displayed until a deactivation and regular re-activation of both fast-drive release paths takes place.
43 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.1 | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 4.8.1.1.1.1 | 562344 | Measure | Check movement release signal status at startup | The system shall check the status of both movement release signal pathsÂ at startup. When an active movement release signal pathÂ is detected, movements from the affected control (sharing one enable hardware line) shall be blocked and anÂ error message shall be displayed, until a deactivation and regular re-activation of both movement release paths takes place. ãÂ Note 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Â Â Note 2: The emergency stop of the affected control remains functional.
44 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.1 | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 4.8.1.1.1.2 | 562378 | Measure | Control elements protected against unintended movement or x-ray activation | The control elements shall be designed in such a manner that unintended movement or x-ray activation by the operator is prevented e.g. :- Joysticks are recessed and have an additional movement triggering switch- No locking of buttons in activated position possible- For the wireless and wired remote control, no function shall be activated when the component is dropped or put reversely on a flat surfaceÂ (see also 60601-1 cl. 15. 3. 1). - At least control elements related to switching the system on/off, chassis movement initiation, brake control and x-ray release shall be recessed. - Handle for power assisted chassis movements shall have two switches / buttons / sensors which are physically separated. - The system shall only allow movement initiation from the handles if force is applied within 2 s timeout after the capacitive sensors are activated. Note: In case the timeout passes by without force being applied or when trying to reinitiate a movement after a temporary stop that exceeds the timeout, the operator needs to release the handles and start a new movement.
45 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.1 | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 4.8.1.1.1.3 | 562377 | Measure | Continuous activation of movements with redundant activation paths | Motorized movement shall be possible only when two redundant signals provided by two independent paths are present (continuous activation), preventing unintended movement activation. When only one of two paths is activated, all motorized movements shall be stopped and inhibited for the respective control. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Â Further note: For network interfaces cyclic retriggering (cyclic messages within e.g. 150 ms) are seen as equivalent to continuous activation.
46 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.1 | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 4.8.1.1.1.4 | 562345 | Measure | Emergency stop at every point of motorized movement activation | Emergency stop shall be available at every point of motorized movement activation. Points of movement activation (control locations) are:- control panel at the FD cover- wireless hand control- wired hand control- Backside of Servo-Handle ChassisÂ The emergency stop shall be designed according to IEC 60947-5-1 and IEC 60947-5-5. Â Â When an emergency stop is activated, all motorized movements of the system except the motorized monitor height adjustment shall be stopped (e.g. stop category 1, thereof e.g. safe stop 1 function according to 61800-5-2) and inhibited until it is deactivated again. Â ãNote 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Â Â Note 2: For behaviour regarding radiation, see <XREF>562135</XREF>.
47 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.2 | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 4.8.1.1.2.1 | 571864 | Measure | Self-test of protection module hardware | The protection module hardware shall execute a self-test of its safety relevant hardware at startup. When the test fails, all motorized movements shall remain inhibited.
48 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.2 | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 4.8.1.1.2.2 | 571866 | Measure | Reference - Check movement release signal status at startup | <XREF>562344</XREF>
49 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.2 | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 4.8.1.1.2.3 | 578161 | Measure | Safe power supply of protection module hardware | The protection module hardware shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. Â When the startup test fails or one of the over-/undervoltage protection detects an error, all motorized movements and radiation shall be inhibited.
50 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.1 | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.8.1.1.2 | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 4.8.1.1.2.4 | 578307 | Measure | General principles protective path | The following general principles for designing theÂ protective path shall be considered:- The software of the protective path shall be developed according to software safety class C (IEC 62304). - Common causes (errors or environmental conditions that could affect C and P simultaneously) shall be considered and excluded by design:Â Â ESD / transients (sufficient electrical isolation between P and C path signals), temperature out of specified limits, etc. Â For components which are connected to the protection module, especially the integration requirements of <XREF>580888</XREF> shall be considered. Â Â Note 1: Safe power supply of P and the mechanical mounting of the encoders have already separate requirements <XREF>578161</XREF>, <XREF>578163</XREF>,<XREF>578173</XREF>. Â Note 2: For design contraints for single fault safety, see <XREF>578311</XREF>.
51 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.2 | 569274 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 4.8.1.2.1 | 569275 | Cause | X-ray chain is not functional (e. g. communication error, hardware error) | 4.8.1.2.1.1 | 564352 | Measure | X-ray chain selftest at startup | The system shall perform a cross-component self-test of the X-ray chain at startup of the system:HV_trigger, idexi line from FLC (IC) to Generator interface and HV_ON with the involved subsystems Xgen, GenIF, FLC, XCU. All involved subsystems / components shall report if the test was successful or not. If the test was not successful, a message shall be displayed to the operator (no X-ray blocking). Remark: HWÂ error inÂ HV_Trigger,Â HV_OnÂ then detectable before X-ray release, cross-component selftest (and that DEXI is not sent to generator due to network configuration issue)
52 | 4.8 | 563429 | Chapter | Startup the system via trolley | 4.8.1 | 563478 | Function | Startup the system via trolley | 4.8.1.2 | 569274 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 4.8.1.2.2 | 569297 | Cause | Delayed boot-up or restart time to recover system functions (e. g. after error, SW freeze) | 4.8.1.2.2.1 | 568860 | Measure | Recovery management | The following time restrictions shall be met for recovery management (this is also valid for EMC immunity testing): * Time from initiating a shut down via reset button, booting the system again, until X-ray ready: 120 sï· * Time from initiating a shut down via reset button, until the previous settings are restored again (this includes registering a patient, set loading parameters (kV, mA, p/s), move the system from angular 0Â° to an exemplary previously stored position of angular 180Â°): 4 min * Time from initiating a regular shut down via off button, booting the system again, until X-ray ready: 4 min * Time from power supply interruption, booting the system again, until X-ray ready: 120 sNote: Due to the limited time to recover all system functions, no emergency radiography modeÂ needs toÂ be provided.
53 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.1 | 563627 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or high contact pressure (HS_051) | 4.10.1.1.1 | 563629 | Cause | Mechanical support structures of the wheels break, the system suddenly tips to one side or tips over completely | 4.10.1.1.1.1 | 561959 | Measure | Mechanical design drive and wheels | The gear, motor shaft and castors shall be designed appropriately for mechanical stress according to IEC 60601-1 cl. 9. 8. 2.
54 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.2 | 578748 | Hazard | Limited imaging functionality | 4.10.1.2.1 | 578749 | Cause | Performance of the imaging system is degraded due to system mangement application | 4.10.1.2.1.1 | 578750 | Measure | Boundary conditions for imaging system performance test | During all phases of the imaging system performance test, the MNP has to be installed and be active all the time to verify that MNP does not disturb the overall system performance.
55 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.3 | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 4.10.1.3.1 | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 4.10.1.3.1.1 | 597004 | Measure | Label electromagnetic fields | A label shall be attached at the backside of the chassis, such that it is visible to an operator when moving the system. It shall contain the following safety signs: ISO 7010-P007, ISO 7010-W006, ISO 7010-W001, ISO 7010-M002
56 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.3 | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 4.10.1.3.1 | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 4.10.1.3.1.2 | 597005 | Measure | User Manual: Electromagnetic fields Chassis | CAUTIONÂ Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes,Â chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Â Possible consequences:Â HazardsÂ on people with implants due to electromagnetic disturbances. Â Remedy:Â If you have active implanted cardiac devices keep a distance ofÂ >10 cm to the above mentioned components.
57 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.3 | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 4.10.1.3.1 | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 4.10.1.3.1.3 | 597007 | Measure | Service Doc: Electromagnetic fields Chassis | CAUTION in Service ManualÂ Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes,Â chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Â Possible consequences:Â HazardsÂ on people with implants due to electromagnetic disturbances. Â Remedy:Â If you have active implanted cardiac devices keep a distance of >10 cm to the above mentioned components.
58 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.3 | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 4.10.1.3.1 | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 4.10.1.3.1.4 | 597011 | Measure | Factory Doc: Electromagnetic fields Chassis | CAUTION in Factory DocumentationÂ Cause: The chassis contains electromagnetic components with high magnetic near fields (Drives and brakes of angular and orbital axes,Â chassis wheels), which might have an influence on implant devices such as cardiac pacemakers. Â Possible consequences:Â HazardsÂ on people with implants due to electromagnetic disturbances. Â Remedy:Â If you have active implanted cardiac devices keep a distance of >10 cm to the above mentioned components.
59 | 4.10 | 563628 | Chapter | System operation - general | 4.10.1 | 563631 | Function | System operation - general | 4.10.1.4 | 597013 | Hazard | User has skin irritation | 4.10.1.4.1 | 597014 | Cause | System contains non-biocompatible material | 4.10.1.4.1.1 | 597019 | Measure | Biocompatible material | The clinical user and the service personnel get in regularÂ hand skinÂ contact with the following components of the system: AccesS (RCP cart, Keyboard, Mouse), DSmonS (Monitor), GridS, HICSwitch, HICTouchCon, MainU (Covers), TrolleyU (Covers, X10 cable and plug), HICForceControl, HICHandControl,Â These components shall be made out of biocompatible material or at least it shall be shown that the utilized material is in common use forÂ regular consumer products or that there is an established history of safe use in a clinical environment, e.g. Â from previous CIOS systems.
60 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.1 | 563485 | Measure | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF>
61 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.2 | 578331 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
62 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.3 | 578316 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
63 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.4 | 578332 | Measure | Reference - System speed limits | <XREF>563486</XREF>
64 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.5 | 563488 | Measure | Prerequisites for fast drive mode | Prerequisites for fast drive mode: Vertical axle shall be in the max. downwards position and the C-arm in the upright position (orbital/angular = 0Â° +/- 5Â° tolerance). The correct position shall be detected single fault safe. The trolley-chassis cable shall be disconnected fromÂ the main unit. Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
65 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.6 | 578291 | Measure | Reference - Collision detection | <XREF>563490</XREF>
66 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.7 | 578358 | Measure | Reference - Collision detection at rear side of the chassis and FD | <XREF>563491</XREF>
67 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.8 | 571916 | Measure | Collision detection via position change of other axles | Position changes of motorized axles which are not used for motorized movements shall be detected (although the brakes of orb. and ang. axles e.g. are activated each during other motorized movements, they can be overridden manually). In case a non-expected position change of more than +/- 5Â°Â is detected, all motorized movements shall stop. Hint: The system may be configured to stop at even lower thresholds < 5Â°.
68 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.9 | 563493 | Measure | Warning - General risk of collisions | WARNINGCause: System movements. Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment. Remedy: Ensure that movements are released only if neither the operator, the patient, other persons or other equipment cannot be endangered. Always pay attention to possible collisions during movements. Make sure that nobody is standing inside the danger area. Ensure that collision sensors of the system are not covered by any accessories e.g. , drape fixtures,Â which would unintentionally activate or prevent activation of the collision detection system. Remove any objects from collision area, e.g. injector or infusion stand. Body penetrating objects introduced in the patient (e.g. catheter, biopsy / kyphoplasty needle) must not be actuated by system movements. Special care is required when the system has already detected a collision situation. Use the emergency stop if required.
69 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.10 | 563494 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
70 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.1 | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 4.12.1.1.1.11 | 563495 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
71 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.2 | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 4.12.1.1.2.1 | 571869 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
72 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.2 | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 4.12.1.1.2.2 | 571870 | Measure | Positioning validity of motorized axes | DuringÂ every movement in orbital or in angular direction, a comparisonÂ between their protectiveÂ position sensorsÂ and their position sensors used by the control systemÂ shall be executed. When an error is detected, motorized movements of the affected axle shall be inhibited. Â Note1: This replaces a regular self-test at startup. This is also required since the protective position sensorsÂ do not provide safe positions according to SIL2 or higher. Â Note2:TheÂ movement stop affects all system drives. Movements from the not affectedÂ drive can be initiated directly again.
73 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.2 | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 4.12.1.1.2.3 | 578163 | Measure | Independent power supply of protective position sensors | The protective position sensors (orbital, angular and for mecanum wheels) shall have an independent power supply from the other position sensors. Â Hint: The power supply for one position sensor shall be provided by the c-path, the power supply of the other position sensor shall be provided by p-path.
74 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.2 | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 4.12.1.1.2.4 | 578164 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
75 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.2 | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 4.12.1.1.2.5 | 595374 | Measure | Reference - Prevent incorrect movement execution - mecanum | <XREF>563729</XREF>
76 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.3 | 571872 | Cause | Collision detection fails due to sleeping error in P path | 4.12.1.1.3.1 | 571874 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
77 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.3 | 571872 | Cause | Collision detection fails due to sleeping error in P path | 4.12.1.1.3.2 | 571873 | Measure | Self-test of the collision sensors | The functionality of the collision sensors shall be tested at startup and continuously during operation (no connection). Â When an error in the collision detection system is detected, the system shall stop all motorized movements and show an error message to the operator. SubsequentÂ movements are allowed withÂ collision mode speed. For detailed behaviour, seeÂ collision detectionÂ requirements <XREF>563490</XREF> and <XREF>563719</XREF>.
78 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.3 | 571872 | Cause | Collision detection fails due to sleeping error in P path | 4.12.1.1.3.3 | 571876 | Measure | Reliable design of collision sensors | TheÂ collision sensorsÂ shall be designed reliably for an expected service life of 10 years (e.g. for an electro-mechanical switch e.g. 100. 000 switching cycles).
79 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.3 | 571872 | Cause | Collision detection fails due to sleeping error in P path | 4.12.1.1.3.4 | 563492 | Measure | Warning - Check collision sensors | WARNINGCause: Undetected failure of the collision sensors. Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment. Remedy: Perform a daily check of the collision sensors at the chassis and the C-arm according to the steps below. Check that the collision sensors are functioning and verify that if a collision has been detected all motorized movements are stopped immediately. Messages must be displayed.
80 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.4 | 563645 | Cause | Unintentional activation of movement due to technical failure (without the operator's request / approval) | 4.12.1.1.4.1 | 563646 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
81 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.4 | 563645 | Cause | Unintentional activation of movement due to technical failure (without the operator's request / approval) | 4.12.1.1.4.2 | 578359 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF>
82 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.5 | 563622 | Cause | User does not pay attention to feet in the movement path during movement (intended): System drives over toes | 4.12.1.1.5.1 | 563625 | Measure | Mechanical safety - allowed gaps mecanum | The gaps between the floor and chassis cover around the mecanum wheels shall be < 35 mm according to IEC 60601-1 table 20 (adult foot). In addition to that, all areas of possible crushing between the covers and the floor shall be covered by <XREF>563490</XREF>.
83 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.5 | 563622 | Cause | User does not pay attention to feet in the movement path during movement (intended): System drives over toes | 4.12.1.1.5.2 | 599924 | Measure | Reference - warning label risk of crushing foot | <XREF>562271</XREF>
84 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.1 | 563638 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
85 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.2 | 563639 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
86 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.3 | 563640 | Measure | Reference - System speed limits | <XREF>563486</XREF>
87 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.4 | 563641 | Measure | Reference - Collision detection | <XREF>563490</XREF>
88 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.5 | 563642 | Measure | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF>
89 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.6 | 563643 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
90 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.6 | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 4.12.1.1.6.7 | 563644 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
91 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.7 | 563674 | Cause | System does not move according to control parameters: Resulting system speed too high due to technical error | 4.12.1.1.7.1 | 563676 | Measure | Reference - System speed limits | <XREF>563486</XREF>
92 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.8 | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 4.12.1.1.8.1 | 563680 | Measure | System acceleration and deceleration limits | The system shall limit the acceleration and deceleration so that tipping over of the system is prevented. In case the acceleration / deceleration threshold is exceeded, the system shall stop all motorized movements. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
93 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.8 | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 4.12.1.1.8.2 | 563681 | Measure | Limit system centrifugal acceleration | The system shall limit the maximum permissible centrifugal acceleration so that tipping over of the system is prevented. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
94 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.8 | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 4.12.1.1.8.3 | 563682 | Measure | Reference - Plausibility check for sensors in movement control elements | <XREF>563673</XREF>
95 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.9 | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 4.12.1.1.9.1 | 571878 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
96 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.9 | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 4.12.1.1.9.2 | 578166 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
97 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.9 | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 4.12.1.1.9.3 | 571879 | Measure | Reference - Self-test of the protective position sensors | <XREF>571870</XREF>
98 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.9 | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 4.12.1.1.9.4 | 571880 | Measure | Reference - Mounting of protective position sensors | <XREF>571870</XREF>
99 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.9 | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 4.12.1.1.9.5 | 578165 | Measure | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF>
100 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.10 | 580816 | Cause | System does not move according to control parameters: During quickstop, the system violates trajectory trueness | 4.12.1.1.10.1 | 596883 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
101 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.11 | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 4.12.1.1.11.1 | 563719 | Measure | Behaviour in collision | When the operator wants to move the system out of a collision (all movements have been stopped), the system shall behave as follows:Movements shall only be possible with reduced speed, after new activation of movement. Â The reduced speed is maintained as long as the movement is activated (and not as long as collision sensor is activated, no "automatic" acceleration). When the movement is stopped, and re-initiated and the collision state has been removed, the subsequent movement shall be performed normally again. Â In case additional collision sensorÂ are activated, the system shall also behave as described above (stop, reduced speed after new movement activation, check real system position changes). Â Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
102 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.11 | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 4.12.1.1.11.2 | 563722 | Measure | Torque limits mecanum | The system shall restrict the torque of the mecanum drives when collision detection is active. If a torque exceeding the limit is detected, the mecanum drives shall stop. Â Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>.
103 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.11 | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 4.12.1.1.11.3 | 563723 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
104 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.11 | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 4.12.1.1.11.4 | 576754 | Measure | Reference - Override angular / orbital brakes manually | The orbital and angular axles can be overriden manually, see <XREF>562449</XREF>. Thus, the system may be able to retract, depending on the initial axle positions.
105 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.12 | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 4.12.1.1.12.1 | 571894 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
106 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.12 | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 4.12.1.1.12.2 | 578167 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
107 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.12 | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 4.12.1.1.12.3 | 571895 | Measure | Reference - Self-test of the protective position sensors | <XREF>571870</XREF>
108 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.12 | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 4.12.1.1.12.4 | 571896 | Measure | Reference - Mounting of protective position sensors | <XREF>571870</XREF>
109 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.12 | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 4.12.1.1.12.5 | 578168 | Measure | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF>
110 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.13 | 571881 | Cause | Collision sensor is active, operator moves further into collision by technical error (e. g. joystick deflection sensor) | 4.12.1.1.13.1 | 571882 | Measure | Reference - Behaviour in collision | <XREF>563719</XREF>
111 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.13 | 571881 | Cause | Collision sensor is active, operator moves further into collision by technical error (e. g. joystick deflection sensor) | 4.12.1.1.13.2 | 571883 | Measure | Reference - Torque limits mecanum | <XREF>563722</XREF>
112 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.14 | 563747 | Cause | System does not stop when user releases control element or stops too late | 4.12.1.1.14.1 | 563748 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
113 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.14 | 563747 | Cause | System does not stop when user releases control element or stops too late | 4.12.1.1.14.2 | 563749 | Measure | Safe stop motorized movements | The braking process (regular release of move enable or quickstop)Â shall be monitored. Â If the system detects an increased delay in reaction or a wrong braking ramp, the system shall stop, e.g. by implementing a safe stopÂ 2 function according to IEC 60204-1 or equivalent. The fast braking process (quickstop) shall be implemented single fault safe. When the system detects an increased delay in reaction or a wrong braking ramp, the system shall stop, e.g. by implementing a safe stop 2 function according to IEC 60204-1 or equivalent. Â Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
114 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.14 | 563747 | Cause | System does not stop when user releases control element or stops too late | 4.12.1.1.14.3 | 563750 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF>
115 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.1 | 571885 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
116 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.2 | 578169 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
117 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.3 | 571886 | Measure | Reference - Self-test of the protective position sensors | <XREF>571870</XREF>
118 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.4 | 578170 | Measure | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF>
119 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.5 | 571887 | Measure | Reference - Mounting of protective position sensors | <XREF>571870</XREF>
120 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.6 | 576755 | Measure | Self-test STO functionality | A self-test of the ability to trigger STO via the protection module shall be performed at startup. In case the test is not passed, the motorized movements affected by the failed STOÂ (e.g. mecanum wheel base)Â shall remain blocked. Â The motorcontrollers shall have a certified STO functionality according to SIL2.
121 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.15 | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 4.12.1.1.15.7 | 576756 | Measure | Reference - Self-test of the brakes | <XREF>564087</XREF>
122 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.16 | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 4.12.1.1.16.1 | 563759 | Measure | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF>
123 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.16 | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 4.12.1.1.16.2 | 563760 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
124 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.16 | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 4.12.1.1.16.3 | 563761 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
125 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.16 | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 4.12.1.1.16.4 | 563762 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
126 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.1 | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.1.17 | 563763 | Cause | System does not stop when user releases control element: Overtravel is increased due to slip , e. g. fluid on the floor | 4.12.1.1.17.1 | 563765 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
127 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.1 | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 4.12.1.2.1.1 | 563727 | Measure | Reference - System acceleration and deceleration limits | <XREF>563680</XREF>
128 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.1 | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 4.12.1.2.1.2 | 563728 | Measure | Reference - Limit system centrifugal acceleration | <XREF>563681</XREF>
129 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.1 | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 4.12.1.2.1.3 | 563729 | Measure | Prevent incorrect movement execution - mecanum | The system shall check if the current mecanum drive information is a valid mecanum movement (fulfill the mecanum equations). When an incorrect movement is detected, the system shall stop motorized movements. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. Â For design constraints for single fault safety, see <XREF>578311</XREF>.
130 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.1 | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 4.12.1.2.1.4 | 563730 | Measure | Prevent tip over by mechanical design | System shall not tip over by mechanical design in the following normal and fault conditions:- emergency stop during max. transport speed in transport position- emergency stop during curve movement in the worst case axle positions and movement speed- system is stopped with max. permissible external additional components mounted in the worst case axle positions and OR speed- one mecanum drive blocks completely during movement in the worst case axle position and movement speed- all mecanum drives block completely during movement (brakes activated) in the worst case axle position and movement speed- emergency stop at OR speed and inclined surface of 7Â° (no transport mode speed)Â Note:During normal stopping (releasing movement control element), the system shall not lift at one side at all. During other stopping scenarios, it is allowed that the rear mecanum wheels of the system lift temporarily. The short duration of the lifting combined with the expectedÂ position of the operator with respect to the possible trapping zone prevent a crushing of the operator foot with the system during this scenario.
131 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.2 | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 4.12.1.2.2.1 | 571889 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
132 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.2 | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 4.12.1.2.2.2 | 578171 | Measure | Reference - Safe power supply of protection module hardware | <XREF>578161</XREF>
133 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.2 | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 4.12.1.2.2.3 | 571890 | Measure | Reference - Self-test of the protective position sensors | <XREF>571870</XREF>
134 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.2 | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 4.12.1.2.2.4 | 571891 | Measure | Reference - Mounting of protective position sensors | <XREF>571870</XREF>
135 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.2 | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 4.12.1.2.2.5 | 578172 | Measure | Reference - Safe power supply of protective position sensors | <XREF>578163</XREF>
136 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.3 | 563731 | Cause | System tips over during movement or during braking: Safety mechanism is activated (e. g. quickstop is not initiated at all) | 4.12.1.2.3.1 | 563732 | Measure | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF>
137 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.4 | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 4.12.1.2.4.1 | 563734 | Measure | Weight restriction mounting external components | The weight of additional components mounted at the C-arm (like navigation components) shall be restricted to max. Â 2 kg.
138 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.4 | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 4.12.1.2.4.2 | 563735 | Measure | Warning - Weight limit for additionally mounted components | WARNINGCause: Additional components are attached to the C-arm. Possible consequences: System movements behave differently, C-arm may be counterbalanced, risk of crushing or tipping over is increased. Remedy: Only mount approved components to the C-arm. The max. additional weight is max. Â 2 kg. Ensure that the floor is free of obstacles and is not slippery.
139 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.4 | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 4.12.1.2.4.3 | 563770 | Measure | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF>
140 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.4 | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 4.12.1.2.4.4 | 578343 | Measure | Warning - Inclined surface restriction | WARNINGCause: System is movedÂ on a steep ramp, over a step or uneven ground. Â Possible consequences: System movements behave differently. The C-arm may be counterbalanced or collide with the floor. Risk of crushing or tipping over is increased. Remedy: The system is designed for steps of max. 2 cm and ramp pitches of max. 5Â° when in transport position. If moving on steeper ramps the lifting column of the C-arm needs to be raised enough to avoid collision with the floor and the system must be moved with care. Note, do not raise the lifting column unnecessarily high because the risk of tipping over is increased.
141 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.4 | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 4.12.1.2.4.5 | 596712 | Measure | Warning - Service instruction for inclined surfaces | Warning - Service Instruction for inclined surfacesÂ WARNINGCause: ãããããããã System is moved on a steep ramp, over a step or uneven ground. Possible consequences: ããSystem movements behave differently. The C-arm may be counterbalanced or collide with the floor. Risk of crushing or tipping over is increased. Remedy: ããããããããThe system is designed for steps of max. 2 cm and ramps of max. 7Â°. Note, when the system is in transport position ramp pitches of max. 5Â° are possible. If moving on steeper ramps the lifting column of the C-arm needs to be raised enough to avoid collision with the floor and the system must be moved with care. Do not raise the lifting column unnecessarily high because the risk of tipping over is increased.
142 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.5 | 563736 | Cause | System tips over during movement or during braking process: Curve radius too small in relation to speed | 4.12.1.2.5.1 | 563737 | Measure | Reference - Limit system centrifugal acceleration | <XREF>563681</XREF>
143 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.6 | 563766 | Cause | System tips over during movement or during braking process: Normal braking ramp is too harsh | 4.12.1.2.6.1 | 563767 | Measure | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF>
144 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.6 | 563766 | Cause | System tips over during movement or during braking process: Normal braking ramp is too harsh | 4.12.1.2.6.2 | 563768 | Measure | Reference - System acceleration and deceleration limits | <XREF>563680</XREF>
145 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.2 | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 4.12.1.2.7 | 580818 | Cause | System lifts at one side during braking process | 4.12.1.2.7.1 | 580819 | Measure | Reference - Prevent tip over by mechanical design | <XREF>563730</XREF>
146 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.1 | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 4.12.1.3.1.1 | 563634 | Measure | Only one active movement control element at a time | The system shall be designed in such a way that if a control element (e.g. , button or handle) is actively used for a specific movement, operation ofÂ any movement from a different control element is inhibited until theÂ active control element is deactivated. Â Note:- Emergency stops and in particular collision sensors are not affected and always overrule the action of control elements if activated. -Â In case a movement has been terminated e.g. by means of a collision, a new movement can be initiated from a different control device even if the control element that triggered the initial movement is still activated, e.g. a still pressed button or touched handle. -Â Initiating more than one movement by pressing more than one button simultaneously at the same device shall also not be possible. -Â Other system functionalities apart from system movements such as activating the laser light shall not be affected by this requirement and can be performed simultaneously during movements.
147 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.1 | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 4.12.1.3.1.2 | 563635 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
148 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.1 | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 4.12.1.3.1.3 | 569307 | Measure | No movement take over across controls | When another control buttonÂ is activated (not controlling the movement) while the movement controlled by another control buttonÂ is stopped, the movement shall be stopped until a deactivation and reactivation of the movement command took place.
149 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.1 | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 4.12.1.3.1.4 | 569308 | Measure | Redundant signals from same control button | Only movement release signal pairs from the same control button or from control buttons which share one enable line shall trigger movements. Note: Chassis handle panel transport position button and chassis handles share the enable line.
150 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.1 | 563648 | Measure | Movement preview for position movements | For all position movements, the system shall have an optical indication of the planned system movement (movement preview), as soon as the position movement is selected (not initiated) (e.g. pictogram, color marking displaying the difference between the current pose and the target pose with respect to the system orientation). ãExceptions: During image acquisition in any mode no preview is required. The movement preview shall be shown after the image acquisition has finished. For movement to the transport and 0/0Â position, no preview is required. ãNote 1: The operator has the possibility to check the subsequent movement. ãNote 2: Position movements are e.g. position storage, park & return, default positions. ãNote 3: Only a preview is required. The preview is required before start of movement. During movement, no display is required (functionally it might be reasonable though in order to see whether the desired position has been reached, at least towards the end of the movement sequence), since the operator observes the mechanical movement of the system.
151 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.2 | 563650 | Measure | Warning - Collisions during position movements without test run | WARNINGCause: Movement to a predefined position. Risk: Risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment. Remedy: Prior to initiating a movement to a predefined position, check the planned movement path via the movement preview and ensure that the path is free of any obstacles. Ensure that the patient is fixated at the table and that the table is positioned correctly (tilt etc. ). Observe the movement continuously.
152 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.3 | 563651 | Measure | Orientation indication for remote movement controls | For control modules where the operator can switch his / her orientation towards the system, the affected control shall indicate the currently selected location of the operator in relation to the system movements. For controls or buttons which have several functions, it shall be clearly indicated to the operator, which function is currently selected. Â Note: The operator has the possibility to check the subsequent movement.
153 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.4 | 564191 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
154 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.5 | 563652 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
155 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.6 | 563655 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
156 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.7 | 563654 | Measure | Reference - System speed limits | <XREF>563486</XREF>
157 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.8 | 563656 | Measure | Reference - Collision detection | <XREF>563490</XREF>
158 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.9 | 563657 | Measure | Reference - Collision detection at rear side of the chassis | <XREF>563491</XREF>
159 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.10 | 563658 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
160 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.2 | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 4.12.1.3.2.11 | 563659 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
161 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.3 | 563671 | Cause | System moves in wrong direction due to technical error (intentional movement): force handles | 4.12.1.3.3.1 | 563672 | Measure | Deactivate whole chassis handle when error in movement activation is detected | When an error in the movement activation circuit in one side of the chassis handle is detected, movement activation from the whole chassis handle shall be inhibited. Â Note: This is checked during startup.
162 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.3 | 563671 | Cause | System moves in wrong direction due to technical error (intentional movement): force handles | 4.12.1.3.3.2 | 563673 | Measure | Plausibility check for sensors in movement control elements | The system shall permanently check the plausibility of analogue sensor outputs for movement control of the force handle sensors and the hand controlÂ joystick during movement activation. When a significant implausible deviation is detected, movement initiated from this control location shall be stopped andÂ inhibited and an error message shall be shown as long as the malfunction is present.
163 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.1 | 563684 | Measure | Main Unit movement shock detection | The system shall check the movement for significant shocks. If a significant shock is detected, motorized movements shall be stopped (but not inhibited) and an appropriate error message shall be displayedãto the user, so that the user can decide to keep or delete stored positions. ãThis is relevant in cases when the collision detection of the system has not been activated.
164 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.2 | 563685 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
165 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.3 | 564164 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
166 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.4 | 563687 | Measure | Reference - System speed limits | <XREF>563486</XREF>
167 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.5 | 563688 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
168 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.6 | 563689 | Measure | Cable deflectors and wheel protection | The wheels shall be equipped with cable deflectors and cover at the upper side to prevent entanglement of external objects in the wheels and prevent external mechanical influence (cables, other disturbing parts which could get caught in the wheels).
169 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.7 | 563018 | Measure | Reference - Weight restriction mounting external components | <XREF>563734</XREF>
170 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.4 | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 4.12.1.3.4.8 | 561948 | Measure | Reference - Warning - Weight limit for additionally mounted components | <XREF>563735</XREF>
171 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.1 | 563691 | Measure | Reference - Main Unit shock detection | <XREF>563684</XREF>
172 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.2 | 599809 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
173 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.3 | 563692 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
174 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.4 | 563693 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
175 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.5 | 563695 | Measure | Reference - System speed limits | <XREF>563486</XREF>
176 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.6 | 563696 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
177 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.7 | 563697 | Measure | Reference - Mechanical design drive and wheels | <XREF>561959</XREF>
178 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.8 | 568836 | Measure | Classification for mechanical shock and vibration - operation | The system shall be designed according to class 3M12 according to IEC 60721-3-3 during operation. Note: Only components located in the main unit are relevant, not in the trolley.
179 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.9 | 568848 | Measure | Movement check in manufacturing | There shall be a dedicated check of the movement functionality in manufacturing.
180 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.5 | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 4.12.1.3.5.10 | 568847 | Measure | Movement check after service / maintenance | There shall be a dedicated check of the movement functionality after service / maintenance activity.
181 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.6 | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 4.12.1.3.6.1 | 563699 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
182 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.6 | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 4.12.1.3.6.2 | 599810 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
183 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.6 | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 4.12.1.3.6.3 | 563700 | Measure | Reference - Mechanical design drive and wheels | <XREF>561959</XREF>
184 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.6 | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 4.12.1.3.6.4 | 563702 | Measure | Reference - Cable deflectors and wheel protection | <XREF>563689</XREF>
185 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.7 | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 4.12.1.3.7.1 | 563755 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
186 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.7 | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 4.12.1.3.7.2 | 563756 | Measure | Reference - Weight restriction mounting external components | <XREF>563734</XREF>
187 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.7 | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 4.12.1.3.7.3 | 563757 | Measure | Reference - Warning - Weight limit for additionally mounted components | <XREF>563735</XREF>
188 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.7 | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 4.12.1.3.7.4 | 563752 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
189 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.8 | 571853 | Cause | Brakes have been opened manually, subsequent motorized movement is performed with brakes open, arbitrary movement of axle | 4.12.1.3.8.1 | 571854 | Measure | Activate other brakes for motorized movement | If a motorized movement is requested, the movement shall only be started after the brakes of all non-relevant axes are activated.
190 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.3 | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.1.3.9 | 580821 | Cause | During braking process, orbital and angular axle are moved unintentionally due to inertia (overridden brakes) | 4.12.1.3.9.1 | 580822 | Measure | Reference - Braking force | <XREF>564086</XREF>
191 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.4 | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 4.12.1.4.1 | 563706 | Cause | Chassis movement is blocked: E. g. Brakes cannot be unlocked due to power supply failure, motor not functional | 4.12.1.4.1.1 | 562205 | Measure | Warning - Establish emergency procedures | WARNING:Cause: System functions including X-ray imaging and system movements get disturbed or get unavailable due to component failures or failures of the external power supply. Â Possible consequences: Disturbance or unavailability of system functions. Interventions cannot be started or continued. Remedy: Consider establishing emergency procedures to manage situations related with the failures of the system, e.g. keeping a reserve system available.
192 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.4 | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 4.12.1.4.2 | 563718 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.12.1.4.2.1 | 563721 | Measure | Reference - Behaviour in collision | <XREF>563719</XREF>
193 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.4 | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 4.12.1.4.2 | 563718 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.12.1.4.2.2 | 571892 | Measure | No additional measures are required | In case the collision state cannot be removed by one control (e.g. error in joystick deflection), an alternative control can be used to maneuver the system out of collision zone.
194 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.5 | 580797 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.1.5.1 | 580798 | Cause | Unintended movement activation with activated e-stop: Operator releases manual brakes during active e-stop | 4.12.1.5.1.1 | 580799 | Measure | Reference - Behaviour in collision | No risk identified, no measure required, since with activated e-stop, the STO is still active, so that a restart of the drives is prohibited safely. The operator is allowed to open the brakes and perform manual movements. During motorized movements, it is not possible to open the brakes manually.
195 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.6 | 563705 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 4.12.1.6.1 | 563707 | Cause | System gets damaged due to a previous collision | 4.12.1.6.1.1 | 563708 | Measure | Reference - Collision detection | <XREF>563490</XREF>
196 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.6 | 563705 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 4.12.1.6.2 | 563741 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 4.12.1.6.2.1 | 563742 | Measure | Reference - Prevent tension at trolley-chassis cable | <XREF>563740</XREF>
197 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.7 | 563709 | Hazard | Critical influence on external electric devices (HS_011) | 4.12.1.7.1 | 563710 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 4.12.1.7.1.1 | 563712 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF>
198 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.7 | 563709 | Hazard | Critical influence on external electric devices (HS_011) | 4.12.1.7.1 | 563710 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 4.12.1.7.1.2 | 562061 | Measure | Warning - EMC disturbances of other equipment | Note: Warning that, even when the system complies with 60601-1-2, other interferences could result in functional disturbances:WARNINGCause: The electromagnetic radiation of the system may cause interference. Risk: This interference may disturb or disrupt the operation of nearby equipment. Remedy:- Examinations involving external equipment require that there are healthcare professionals present, who are trained for managing the malfunctions of these external equipment and associated emergency situations. - In case of disturbances in external equipment operation, try to re-orient or re-locate these. - To prevent such disturbances, it may be necessary to shield the external equipment/location.
199 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.8 | 563714 | Hazard | Non-critical influence on external electric devices (HS_010) | 4.12.1.8.1 | 563715 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 4.12.1.8.1.1 | 563716 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF>
200 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.8 | 563714 | Hazard | Non-critical influence on external electric devices (HS_010) | 4.12.1.8.1 | 563715 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 4.12.1.8.1.2 | 563717 | Measure | Reference - Warning - EMC disturbances of other equipment | <XREF>562061</XREF>
201 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.9 | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 4.12.1.9.1 | 563739 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 4.12.1.9.1.1 | 563740 | Measure | Prevent tension at trolley-chassis cable | The system shall detect any tensile load on the trolley-chassis cable. When a hazardous tensile load is detected, motorized chassis movements shall be stopped before any damage to the system or any influence on theÂ movement trajectory occurs. Subsequent movements may only be initiated with collision speed <XREF>576777</XREF>. At the same time, the torque supervision <XREF>563722</XREF> shall be applied. An appropriate message or indication shall be displayed to the operator. ãNote 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Â Note 2: When the trolley is moved (manually always), there is direct feedback to the operator who pushes or pulls the trolley, so that it is not likely that any damage occurs due to that, also based on the field information based on previous systems.
202 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.9 | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 4.12.1.9.1 | 563739 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 4.12.1.9.1.2 | 576777 | Measure | Behaviour with activated trolley-chassis cable tension detection | With activated trolley-chassis cable tension detection (that means after stopping and new movement initiation), the system shall only allow motorized chassis movements with reduced speed. Â The behaviour shall be the same as in a collision detection situation. <XREF>563490</XREF>Â Note: In case the operator cannot relieve the tension with motorized movements, he / she can unplug the cable, move the cable, untangle the cable or the setup,Â or move the trolley.
203 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.9 | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 4.12.1.9.2 | 576778 | Cause | System is moved away from the trolley, trolley-chassis cable tension detection is not working | 4.12.1.9.2.1 | 576779 | Measure | Self-test trolley-chassis cable tension detection | The system shall perform a self-test of the trolley-chassis cable tension detection as long as the cable is plugged in. If an error is detected the behavior shall be the same as if cable tension is detected (see <XREF>576777</XREF> and <XREF>563740</XREF>).
204 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.10 | 563623 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.1.10.1 | 563624 | Cause | Collision is detected, the system is stopped, but the operator is not aware of the situation and is confused about the stop | 4.12.1.10.1.1 | 563626 | Measure | Collision state indication | When a collision is detected, the system shall indicate the state to the operator. Â Note: This refers to collisions detected by collision detection sensors. For Chassis & FD handle see <XREF>563491</XREF>.
205 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.11 | 563743 | Hazard | Severe limitation of movement functionality (HS_111) | 4.12.1.11.1 | 563744 | Cause | Collision sensor fails during any other operation (direction movement, standstill,â¦): permanently activated collision sensor | 4.12.1.11.1.1 | 563745 | Measure | No additional measure required
206 | 4.12 | 563438 | Chapter | Move main unit | 4.12.1 | 563475 | Function | Move main unit - motorized, general | 4.12.1.12 | 576792 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.12.1.12.1 | 576793 | Cause | Regulatory source to skin distance is decreased when collision detection mechanism of X-ray source cover is activated | 4.12.1.12.1.1 | 576794 | Measure | No measure required | IEC 60601-2-54, clause 203. 9 requires a minimum focal spot to skin distance (aka next surface which mechanically prevents that the operator cannot place any anatomy closer to the X-ray source, e.g. X-ray window) of 200 mm. Theoretically, it is possible that the anatomy is positioned directly on the X-ray window so that the focal spot to skin distance would be violated. However, this is not considered as reasonably foreseeable sequence of events. Therefore, it is concluded, that no significant risk is related to the design. No additional measures are required and the design is acceptable.
207 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.1 | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 4.12.2.1.1.1 | 564224 | Measure | Pairing of hand control and system | The hand control and the system shall be paired. The pairing shall be performed by an interaction on the control and on the system side within a defined time to ensure the paring of specified devices only. Only pairs shall operate together.
208 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.1 | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 4.12.2.1.1.2 | 564225 | Measure | Optical indication of system and hand control pairs | There shall be an optical indication on the system and the hand control, e.g. colored symbols, indicating that they belong together.
209 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.1 | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 4.12.2.1.1.3 | 596993 | Measure | Label wireless hand control | Factory Documentation:Â Attach a label on the wireless hand control to which system it belongs to.
210 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.1 | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 4.12.2.1.1.4 | 600715 | Measure | Optical indication of system and hand controls - collision | In Service DocumentationãWARNINGCause:Â Movement released from aÂ remote hand controlÂ in another roomPossible Consequences: Collision of system with persons or objectRemedy: Before performing any service activities with the remote hand control, check that the remote hand control belongs to the system and check that the indicating label is in place and properly visible.
211 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.1 | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 4.12.2.1.1.5 | 600716 | Measure | Labeling wireless hand control - collision | WARNINGCause:Â Movement released from aÂ remote hand controlÂ in another roomPossible Consequences: Collision of system with persons or objectRemedy:Â As part of the system installation, the wirelessÂ hand controlÂ is labeled according to which system it belongs to. When utilizingÂ theÂ wireless hand control, check thatÂ it belongs to the systemÂ according to theÂ indicating labels and ensure there is visual contact with the system.
212 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.2 | 564234 | Cause | Control is dropped and function gets activated | 4.12.2.1.2.1 | 564235 | Measure | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF>
213 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.1 | 564238 | Measure | Deactivate movement and radiation at wireless hand control receiver | The movement and radiationÂ request shall be deactivated at the wireless hand control receiver, when the receiver looses its connection to the sender or when the communication is disturbed (independent of a potential running movement or radiation initiated from this control). If the wireless hand control receiver becomes non-functional, movement and radiation shall be deactivated from this control. In case an emergency stop has been activated before connection loss, it shall remain activated by the receiver. ãNote: Safety architecture "mixed", see <XREF>561860</XREF> and <XREF>543971</XREF>. Â For design constraints for single fault safety, see <XREF>578311</XREF>.
214 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.2 | 595749 | Measure | Deactivate movement and radiation at wired hand control | If the wired hand control is disconnected or becomes non-functional, movement and radiation shall be deactivated from this control. In case an emergency stop has been activated before unplugging or non-functionality of the wired hand control, it shall remain activated.
215 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.3 | 564239 | Measure | Indication of e-stop functionality | The system shall indicate on the wireless hand control and on the user interface (if available, not available in battery mode) if the emergency stop of the wireless hand control is functional. It is acceptable if the wireless hand control uses e.g. , a blinkingÂ LED for displaying the functionality. Â Note, there are two independent time outs, one for the hand control (for detecting the connection status and indicating it) and one at the receiver (for detecting the connection status and indicating it on system side).
216 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.4 | 571957 | Measure | User message about non-functional wireless hand control and instruction to put it in the drawer | A message shall be displayedÂ when the wireless hand control becomes non-functional or connection gets lost which shall instruct the operator to put the control into the drawer, and which shall be confirmed by the operator.
217 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.5 | 595744 | Measure | User message if wired hand control is non-functional | A message shall be displayed when the wired hand control is non-functional and shall instruct the operator to put the wired hand control away. The message shall be confirmed by the operator.
218 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.6 | 564240 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
219 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.7 | 564241 | Measure | Reference - Collision detection | <XREF>563490</XREF>
220 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.3 | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 4.12.2.1.3.8 | 564242 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
221 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.4 | 578630 | Cause | Emergency stop on the wireless hand control is not functional when control goes to sleep mode | 4.12.2.1.4.1 | 578632 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>Emergency stop on wireless hand control is functional when control is in sleep mode (with only slightly increased delay time).
222 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.5 | 564246 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 4.12.2.1.5.1 | 564247 | Measure | Reference - Deactivate movement and radiation at wireless hand control | <XREF>564238</XREF>
223 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.5 | 564246 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 4.12.2.1.5.2 | 564248 | Measure | Reference - Alternative deactivation of emergency stop - hand control | <XREF>564251</XREF>
224 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.6 | 568773 | Cause | Error in wireless data transmission: Delay, loss of data packet(s) | 4.12.2.1.6.1 | 568776 | Measure | Reference - Deactivate movement and radiation at wireless hand control receiver | <XREF>564238</XREF>
225 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.7 | 568774 | Cause | Error in wireless data transmission: Wrong sequence of data packets, repetition of invalid packets | 4.12.2.1.7.1 | 568779 | Measure | Reference - Safe wireless transfer of safety relevant information | <XREF>564349</XREF>
226 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.8 | 568775 | Cause | Error in wireless data transmission: Wrong information transmitted | 4.12.2.1.8.1 | 564349 | Measure | Safe wireless transfer of safety relevant information | The transfer of safety relevant movement and radiation signals / information from wireless remote controls shall be safe and protected against transmission errors (ensure integrity of packets including sequence). When an integrity error is detected, the movement and radiation enable shall be deactivated at the wireless hand control receiver (independent of a potential running movement or radiation initiated from this control). In case an emergency stop has been activated when the error is detected, it shall remain activated by the receiver.
227 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.9 | 576770 | Cause | Error in wireless data transmission: Sleeping error in protection mechanisms of wireless transfer | 4.12.2.1.9.1 | 576771 | Measure | Self-test safe wireless transfer components | A self-test of the components required for protection mechanisms for safe wireless transfer shall be performed at system startup. This includes the protection mechanism of the receiver as well. In case components cannot be self-tested, the component shall be implemented redundantly. Â Note: For system reaction see <XREF>564351</XREF>.
228 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.9 | 576770 | Cause | Error in wireless data transmission: Sleeping error in protection mechanisms of wireless transfer | 4.12.2.1.9.2 | 578173 | Measure | Safe power supply for safe wireless transfer components | The components required for protection mechanisms for safe wireless transfer shall have a safe power supply: The over- / undervoltage protection shall be either redundantly implemented or be tested at startup. The voltage limits shall be defined according to the guaranteed working voltages of the hardware to ensure correct functionality. Â Note: For system reaction see <XREF>564351</XREF>.
229 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.10 | 568777 | Cause | Error in wireless data transmission: Internal error in receiver | 4.12.2.1.10.1 | 564351 | Measure | Safe state wireless hand control receiver | When the wireless hand control receiver detects an internal problem (self-test, watchdog), the receiver shall go to a safe state: X-ray and movement enable deactivated and emergency stop activated. Â ãNote: Safety architecture "mixed", see <XREF>561860</XREF> and <XREF>543971</XREF>. Â For design constraints for single fault safety, see <XREF>578311</XREF>.
230 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.11 | 568778 | Cause | Error in movement release circuit of the wireless control | 4.12.2.1.11.1 | 578324 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
231 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.11 | 568778 | Cause | Error in movement release circuit of the wireless control | 4.12.2.1.11.2 | 578325 | Measure | Reference - Plausibility check for sensors in movement control elements | <XREF>563673</XREF>
232 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.1 | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.2.1.12 | 568780 | Cause | Wireless control comes into reach or connect. reestablished and mov. button is activ. or tech. error in mov. release circuit | 4.12.2.1.12.1 | 568781 | Measure | Block movement when controls connecting | The system shall check the status of both movement release signal paths when controls are connected to the system. If an active movement release signal path is detected at hot plug, movements from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both movement release paths take place.
233 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.2 | 591782 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.2.2.1 | 591783 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates non-correctly calibrated control | 4.12.2.2.1.1 | 591784 | Measure | Reference - Pairing of hand control and system | <XREF>564224</XREF>
234 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.2 | 591782 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.2.2.1 | 591783 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates non-correctly calibrated control | 4.12.2.2.1.2 | 591785 | Measure | Reference - Optical indication of system and hand control pairs | <XREF>564225</XREF>
235 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.1 | 564227 | Cause | Wireless remote control is not functional: Battery runs empty | 4.12.2.3.1.1 | 564228 | Measure | Check battery charge of wireless remote control | The system shall check the battery charge of the wireless remote control at start-up shall warn the operator when the remaining battery charge would not be sufficient for one day of system operation with the expected use profile.
236 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.1 | 564227 | Cause | Wireless remote control is not functional: Battery runs empty | 4.12.2.3.1.2 | 564229 | Measure | Wired remote control mandatory | The system shall always be equipped with a wired remote control when it is equipped with a wireless remote control (option).
237 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.2 | 564230 | Cause | Wireless remote control is not functional: Liquid intrudes into the electronics | 4.12.2.3.2.1 | 564231 | Measure | IP protection class remote controls | The wireless and the wired remote control shall have protection class IPx5 according to EN 60529 with closed battery compartment.
238 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.2 | 564230 | Cause | Wireless remote control is not functional: Liquid intrudes into the electronics | 4.12.2.3.2.2 | 568818 | Measure | Caution - Correct change of batteries wireless remote control | CAUTIONCause: Incorrect exchange of batteries of the wireless remote control. Possible consequences: Wireless remote control not functional. Remedy: Correct exchange of batteries with dry hands and dry control. Do not let the battery compartment open, close it as soon as you are finished, e.g. also for storing the control.
239 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.3 | 564232 | Cause | Wireless remote control is not functional: Control is dropped and gets damaged | 4.12.2.3.3.1 | 564233 | Measure | Reference - Control elements protected against unintended movement activation | <XREF>562378</XREF>
240 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.4 | 564243 | Cause | Emergency stop on the wired / wireless remote control is activated unintentionally | 4.12.2.3.4.1 | 564244 | Measure | Prevent unintended activation of e-stop | The emergency stops shall be protected against unintended activation (for the wired and wireless remote control e.g. mounting within a recessed surface of the surrounding control, but no closed mechanical shroud; for other emergency stops e.g. mounted away from areas which are foreseeably unintentionally touched etc. ).
241 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.4 | 564243 | Cause | Emergency stop on the wired / wireless remote control is activated unintentionally | 4.12.2.3.4.2 | 564245 | Measure | Indication of e-stop activation and location | When an emergency stop is activated, the system shall indicate that anÂ emergency stop is activated and at which location. In battery mode it is acceptable, that only the activation state of emergency stop is indicated, and not the location.
242 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.3 | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 4.12.2.3.5 | 591780 | Cause | Wireless remote control comes into wireless range or becomes functional again with a pressed emergency stop | 4.12.2.3.5.1 | 591781 | Measure | No additional measure required | When the wireless hand control comes into wireless range again or becomes functional again after being out of the range or non-functional, the wireless hand control is not accepted as a valid control when the emergency stop is activated at that time due to regulatory reasons. This prevents that an ongoing movement is interrupted by an unintentional emergency stop activation while the wireless hand control is out of reach or not functional (emergency stop pressed when control is carried in the pocket of a coat, or when the batteries are exchanged etc. ).
243 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.4 | 564249 | Hazard | Severe limitation of movement functionality (HS_111) | 4.12.2.4.1 | 564250 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 4.12.2.4.1.1 | 564252 | Measure | Reference - Deactivate movement and radiation at wireless hand control receiver | <XREF>564238</XREF>
244 | 4.12 | 563438 | Chapter | Move main unit | 4.12.2 | 564220 | Function | Perform movements via wireless hand control | 4.12.2.4 | 564249 | Hazard | Severe limitation of movement functionality (HS_111) | 4.12.2.4.1 | 564250 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 4.12.2.4.1.2 | 564251 | Measure | Alternative deactivation of emergency stop - hand control | There shall be an alternative safe way to deactivate an emergency stopÂ of the wireless and the wired hand control withÂ a deliberate action (two subsequent actions) by the operator. For the wireless hand control: Only if the hand control connection is lost with an activated e-stop by thisÂ control: After the connection loss, an activation and deactivation must take place at another emergency stop. For the wired hand control: Only if the wired hand control functionality is lost with an activated e-stop by this control: Unplugging the wired hand control and thereafter activating and deactivating another emergency stop. Note 1:Â The single actions of unplugging the wired hand controlÂ orÂ putting the wireless hand control in the drawer shall not deactivate the emergency stop. IfÂ a connection is re-established and the e-stop has been released in the meanwhile at another device, the e-stop shall not be present. If the e-stop has not been released in the meanwhile at another device, then the e-stop shall remain present until resolved. Â Â Note 2: This is deemed to sufficiently prevent the scenario, that an e-stop is deactivated unintentionally when e.g. two operators activate their e-stops nearly simultaneously (followed by a deactivation of the second one), since it is deemed very unlikely that the sequence connection loss, and e-stop already is activated, another e-stop is activated for another purpose or unintentionally than to deactivate the e-stop in general.
245 | 4.12 | 563438 | Chapter | Move main unit | 4.12.3 | 569231 | Function | Perform movements via wired hand control | 4.12.3.1 | 569232 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.3.1.1 | 562466 | Cause | Disconnect wired hand control with activated emergency stop, system starts moving | 4.12.3.1.1.1 | 571978 | Measure | Reference - Alternative deactivation of emergency stop - hand control | <XREF>564251</XREF>
246 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.1 | 564079 | Cause | Control system does not brake / brakes too slowy (Quickstop not initiated or wrong braking ramp) | 4.12.4.1.1.1 | 564080 | Measure | Reference - Safe stop motorized movements | <XREF>563749</XREF>
247 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.2 | 564081 | Cause | After activating the emergency stop, movements stop, operator deactivates emergency stop, movement is initiated again | 4.12.4.1.2.1 | 564082 | Measure | No automatic movement release after e-stop deactivation | If the move requestÂ signal is still active after the emergency stop is deactivated again, the system shall block movements until a normal deactivation of the enable signal and re-activation took place. Â See <XREF>564245</XREF> for user messaging for emergency stop. Â Note: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. Â´Â For design constraints for single fault safety, see <XREF>578311</XREF>.
248 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.3 | 562460 | Cause | Emergency stop not functional | 4.12.4.1.3.1 | 563483 | Measure | Emergency stop self-test and failure exclusions | The emergency stop circuit shall be checked during start-up. In case the test is not passed, motorized movements shall remain blocked. Â Failure modes which cannot be self-tested or prevented shall be covered by failure exclusions. ãNote 1: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>. Â Note 2: This mechanism does not cover mechanical failures of the emergency stop switches, therefore see <XREF>563484</XREF>.
249 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.3 | 562460 | Cause | Emergency stop not functional | 4.12.4.1.3.2 | 563484 | Measure | Warning - Check emergency stop | WARNINGCause: Emergency Stop button not functioning. Risk: Risks of injury by crushing or collision, system could stop too late in a critical situation. Remedy: Perform the check of the emergency stop buttons on a daily basis according to the steps below.
250 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.3 | 562460 | Cause | Emergency stop not functional | 4.12.4.1.3.3 | 571900 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
251 | 4.12 | 563438 | Chapter | Move main unit | 4.12.4 | 564077 | Function | Activate emergency stop | 4.12.4.1 | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.4.1.4 | 571898 | Cause | After activating the emergency stop, movement is initiated again, sleeping error in P path | 4.12.4.1.4.1 | 571899 | Measure | Reference - Self-test of protection module hardware | <XREF>571864</XREF>
252 | 4.12 | 563438 | Chapter | Move main unit | 4.12.5 | 564083 | Function | Maintain stable system position | 4.12.5.1 | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.5.1.1 | 564084 | Cause | System moves unintentionally during standstill | 4.12.5.1.1.1 | 564086 | Measure | Braking force | The braking force needs to be sufficient to hold the system in standstill on an inclined plane of 5Â° in non-transport position (worst case position of axles), and of 10Â° in transport position. The braking force needs to be sufficient to hold the system in standstill during 3D. Â During any stopping scenario of the mecanum platform, the orbital and angular axle shall not moveÂ more than +/- 5Â°Â (brakes overridden due to inertia). Â Hint: Definition of ramp inclinations (5Â° and 10Â°) were in alignment with IEC 60601-1, clause 9. 4. 2. 2
253 | 4.12 | 563438 | Chapter | Move main unit | 4.12.5 | 564083 | Function | Maintain stable system position | 4.12.5.1 | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.5.1.1 | 564084 | Cause | System moves unintentionally during standstill | 4.12.5.1.1.2 | 564087 | Measure | Self-test of the brakes | An automatic brake test for mecanum wheels shall be performed at first chassis movement in order to detect if a brake does not work properly. This shall include the ability of the protection module to activate the brakes. The test shall be performed before a motorized movement takes place. Â If the test is not passed, the chassis movements shall remain blocked. Â ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. For design constraints for single fault safety, see <XREF>578311</XREF>.
254 | 4.12 | 563438 | Chapter | Move main unit | 4.12.5 | 564083 | Function | Maintain stable system position | 4.12.5.1 | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.5.1.2 | 571856 | Cause | System moves unintentionally during brake test | 4.12.5.1.2.1 | 594854 | Measure | Reference protect unintentional activation rescue user | <XREF>564146</XREF>
255 | 4.12 | 563438 | Chapter | Move main unit | 4.12.5 | 564083 | Function | Maintain stable system position | 4.12.5.1 | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.12.5.1.2 | 571856 | Cause | System moves unintentionally during brake test | 4.12.5.1.2.2 | 594855 | Measure | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF>
256 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.1 | 564154 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 4.12.6.1.1.1 | 564155 | Measure | Positioning accuracy | The absolute precision for chassis positioning over a travel range of 1 m shall be at least 2 cm. The repetitive accuracy for chassis repositioning over a travel range of 1 m (back- and forth movement) shall be at least 2 cm.
257 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.2 | 564156 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.6.1.2.1 | 564158 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
258 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.2 | 564156 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.6.1.2.2 | 599811 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
259 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.3 | 564159 | Cause | Deviation of target position too large: External influences (slip, obstacles,â¦) | 4.12.6.1.3.1 | 564160 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
260 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.1 | 564162 | Measure | Reference - Movement preview for position movements | <XREF>563648</XREF>
261 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.2 | 564193 | Measure | Timeout for position movement selection | After at most 60 seconds of motorized standstill after selection of a target position for any position movement, there shall be a timeout deselecting the selected target position automatically.
262 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.3 | 569216 | Measure | Movement range limit park and return | The distance between starting and end point of the wheel platformÂ for park and return movements shall be limited toÂ 4 m. The maximum wheel platform rotation shall be limited to 135Â°.
263 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.4 | 564170 | Measure | Reference - Warning - Collisions during position movements without test run | <XREF>563650</XREF>
264 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.5 | 564192 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
265 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.6 | 564163 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
266 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.7 | 564165 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
267 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.8 | 564166 | Measure | Reference - System speed limits | <XREF>563486</XREF>
268 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.9 | 564167 | Measure | Reference - Collision detection | <XREF>563490</XREF>
269 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.4 | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 4.12.6.1.4.10 | 564168 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
270 | 4.12 | 563438 | Chapter | Move main unit | 4.12.6 | 564152 | Function | Perform position movement: Park and Return | 4.12.6.1 | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.6.1.5 | 564171 | Cause | Operator initiates park and return on outdated / invalid position data | 4.12.6.1.5.1 | 564172 | Measure | Prerequisites park and return and position memory | A park and return or position memory movement shall only be possible when the system can reconstruct the movement path (e.g. when the system is rebooted, when the rescue function has been activated, the function is not available).
271 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.1 | 564200 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 4.12.7.1.1.1 | 564201 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
272 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.2 | 564202 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.7.1.2.1 | 564203 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
273 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.2 | 564202 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.7.1.2.2 | 599812 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
274 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.3 | 564205 | Cause | Deviation of target position too large: External influences (slip, obstacles,â¦) | 4.12.7.1.3.1 | 564206 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
275 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.1 | 564208 | Measure | Reference - Movement preview for automatic movements | <XREF>563648</XREF>
276 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.2 | 564209 | Measure | Reference - Timeout for automatic movement selection | <XREF>564193</XREF>
277 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.3 | 564219 | Measure | Movement range limit position memory | TheÂ distance between starting and end point forÂ position memoryÂ movements of the wheel platform shall be limited toÂ 4 m. The maximum wheel platform rotation shall be limited to 135Â°.
278 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.4 | 564210 | Measure | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF>
279 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.5 | 564211 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
280 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.6 | 564212 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
281 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.7 | 564213 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
282 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.8 | 564214 | Measure | Reference - System speed limits | <XREF>563486</XREF>
283 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.9 | 564215 | Measure | Reference - Collision detection | <XREF>563490</XREF>
284 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.4 | 564207 | Cause | Operator has selected an incorrect target position from position storage | 4.12.7.1.4.10 | 564216 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
285 | 4.12 | 563438 | Chapter | Move main unit | 4.12.7 | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 4.12.7.1 | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.7.1.5 | 564217 | Cause | Operator initiates movement to stored position on outdated / invalid position data | 4.12.7.1.5.1 | 564218 | Measure | Reference - Prerequisites park and return and position memory | <XREF>564172</XREF>
286 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.1 | 564175 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 4.12.8.1.1.1 | 564176 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
287 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.2 | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.8.1.2.1 | 564178 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
288 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.2 | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.8.1.2.2 | 599813 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
289 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.2 | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.8.1.2.3 | 594826 | Measure | Reference - no movement takeover across controls | <XREF>569307</XREF>
290 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.2 | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.8.1.2.4 | 594827 | Measure | Reference - only one active control element | <XREF>563634</XREF>
291 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.3 | 564179 | Cause | Deviation of target position too large: External influences (slip, obstacles,â¦) | 4.12.8.1.3.1 | 564180 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
292 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.1 | 564182 | Measure | Reference - Orientation indication for remote movement controls | <XREF>563651</XREF>
293 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.2 | 564183 | Measure | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF>
294 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.3 | 564195 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
295 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.4 | 564184 | Measure | Reference - Set speed values | <XREF>562463</XREF>
296 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.5 | 564185 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
297 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.6 | 564186 | Measure | Reference - System speed limits | <XREF>563486</XREF>
298 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.7 | 564187 | Measure | Reference - Collision detection | <XREF>563490</XREF>
299 | 4.12 | 563438 | Chapter | Move main unit | 4.12.8 | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 4.12.8.1 | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.8.1.4 | 564181 | Cause | Operator has selected an incorrect movement function | 4.12.8.1.4.8 | 564188 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
300 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.1 | 564312 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 4.12.9.1.1.1 | 564313 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
301 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.2 | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.9.1.2.1 | 564315 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
302 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.2 | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.9.1.2.2 | 599814 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
303 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.2 | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.9.1.2.3 | 564316 | Measure | Reference - no movement takeover across controls | <XREF>569307</XREF>
304 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.2 | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.12.9.1.2.4 | 594828 | Measure | Reference - only one active control element | <XREF>563634</XREF>
305 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.3 | 564317 | Cause | Deviation of target position too large: External influences (slip, obstacles,â¦) | 4.12.9.1.3.1 | 564318 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
306 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.1 | 564320 | Measure | Reference - Movement preview for automatic movements | <XREF>563648</XREF>
307 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.2 | 564321 | Measure | Reference - Timeout for position movement selection | <XREF>564193</XREF>
308 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.3 | 569237 | Measure | Movement range limit click and go | The travel range for click and goÂ movements shall be limited toÂ 1 m.
309 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.4 | 564322 | Measure | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF>
310 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.5 | 564323 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
311 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.6 | 564324 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
312 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.7 | 564325 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
313 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.8 | 564326 | Measure | Reference - System speed limits | <XREF>563486</XREF>
314 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.9 | 564327 | Measure | Reference - Collision detection | <XREF>563490</XREF>
315 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.4 | 564319 | Cause | Operator has selected an incorrect target position on the image | 4.12.9.1.4.10 | 564328 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
316 | 4.12 | 563438 | Chapter | Move main unit | 4.12.9 | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 4.12.9.1 | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.12.9.1.5 | 564329 | Cause | Operator initiates click and go on outdated / invalid position / image data | 4.12.9.1.5.1 | 564190 | Measure | Prerequisites click and go | Click and go shall only be possible when a patient is registered, an image has been acquired, and when the system can reconstruct the movement path. The function shall consider system movements between image acquisition and function activation, and shall consider image rotation and flip. (I. e. when a previous patient is loaded and an image is loaded again, the feature is not available; after a system reboot or activation of rescue function, the feature is not available. )
317 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.1 | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 4.15.1.1.1.1 | 564129 | Measure | Provide rescue button | The system chassis shall be able to be movable manually after an intentional action by the operator (e.g. releasing the brakes via a separate button which is secured against unintended activation) in case the chassis is not movable. Â The movement shall be possible with reasonable force (max. Â 400 N (200 N is threshold for mobile equipment propulsion under normal conditions)).
318 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.1 | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 4.15.1.1.1.2 | 562449 | Measure | Override angular / orbital brakes manually | The orbital and angular axle shallÂ be designed in a way that the axle is movable manually e. g. Â in case of aÂ power failure.
319 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.1 | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 4.15.1.1.1.3 | 564130 | Measure | Caution - Use rescue function | CAUTIONRisk: The system cannot be removed from the patient table in case of an emergency situation, e.g. cardiopulmonary resuscitation. Cause: Motorized system movements are blocked. Remedy: Check if any emergency stop has been activated and deactivate it. Use the rescue button in order to move the system away from the patient. Note that the orbital and angular axles still can be moved manually.
320 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.2 | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 4.15.1.1.2.1 | 564132 | Measure | Reference - Mechanical design drive and wheels | <XREF>561959</XREF>
321 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.2 | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 4.15.1.1.2.2 | 564133 | Measure | Rotate chassis with blocked mecanum wheel | In case one of the mecanum wheels is permanently blocked, the system shall still be able to be moved / rotated 45 Â° with reasonable force (max. 400 N) after activating the rescue function.
322 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.2 | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 4.15.1.1.2.3 | 564140 | Measure | Reference - Provide rescue function | <XREF>564129</XREF>
323 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.2 | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 4.15.1.1.2.4 | 564134 | Measure | Reference - Override angular / orbital brakes manually | <XREF>562449</XREF>
324 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.2 | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 4.15.1.1.2.5 | 564135 | Measure | Reference - Caution - Use rescue function | <XREF>564130</XREF>
325 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.3 | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.15.1.1.3.1 | 564137 | Measure | Reference - Behaviour in collision | <XREF>563719</XREF>
326 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.3 | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.15.1.1.3.2 | 564139 | Measure | Reference - Collision detection | <XREF>563490</XREF>
327 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.3 | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.15.1.1.3.3 | 564141 | Measure | Reference - Provide rescue button | <XREF>564129</XREF>
328 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.3 | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.15.1.1.3.4 | 564142 | Measure | Reference - Override angular / orbital brakes manually | <XREF>562449</XREF>
329 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.1 | 564127 | Hazard | Patient rescue hindered (HS_140) | 4.15.1.1.3 | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 4.15.1.1.3.5 | 564143 | Measure | Reference - Caution - Use rescue function | <XREF>564130</XREF>
330 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.1 | 564150 | Cause | Technical failure in the rescue function activates the function and releases the mecanum brakes, system rolls | 4.15.1.2.1.1 | 564151 | Measure | Prevent unintentional activation of rescue function by design | The unintentional activation of the rescue function due to technical errors shall be prevented by redundant activation channels (e.g. two switches and safety relays). Only when both channels are activated, the function shall be activated. Inconsistent activation of one channel or unintended permanent activation of both channelsÂ shall be checked during startup. InÂ these casesÂ movement shall be inhibited.
331 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.1 | 564150 | Cause | Technical failure in the rescue function activates the function and releases the mecanum brakes, system rolls | 4.15.1.2.1.2 | 576703 | Measure | Check rescue functionality during planned maintenance | Work step during planned maintenance to ensure the correct functionality of the rescue function including correct deactivation of the function after releasing the button.
332 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.2 | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 4.15.1.2.2.1 | 564146 | Measure | Prevent unintentional activation of rescue function by user error | The button for initiating the rescue function shall be protected against unintended activation. The function shall require continuous activation by the operator and the activation signal shall be implemented redundantly. A consistency check of the rescue signals shall be performed at startup in order to detect a first failure in the rescue circuit. If the test is not passed the system shall be blocked.
333 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.2 | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 4.15.1.2.2.2 | 564147 | Measure | Inhibit motorized movements of chassis - rescue function | If the rescue button is activated during an ongoing movement, the movement shall be stopped by braking (regular braking ramp or quick-stop) in order to prevent spinning out of the drives. Once the movement is stopped the STO shall be activated and movement enable blocked for all motorized axes. Once the rescue button is disengaged the STO is deactivated and movements can be initiated again. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. The SFS functionality is applicable for the breaking functionality, not for recognition and display of of first errors. Â For design constraints for single fault safety, see <XREF>578311</XREF>. Â´
334 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.2 | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 4.15.1.2.2.3 | 564148 | Measure | Warning - Correct use of rescue function | WARNINGCause: Rescue button is operated incorrectly. Risk: Increased risk of collision, risk of injury to patient or operator, risk of damage to the system or other equipment. Remedy: The rescue button is only to be used in emergency situations when immediate access to the patient is required, e.g. for cardiopulmonary resuscitation. Do not operate other motorized movements when rescue button is activated. Do not operate the function during movements. Note that the collision detection and emergency stop will have no effect when this function is activated.
335 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.3 | 576759 | Cause | Rescue function is activated during activated emergency stop | 4.15.1.2.3.1 | 576760 | Measure | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF>
336 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.2 | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.15.1.2.4 | 580834 | Cause | Operator misuses the rescue function intentionally or unintentionally | 4.15.1.2.4.1 | 580835 | Measure | Reference - Inhibit motorized movements of chassis - rescue function | <XREF>564147</XREF>
337 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.3 | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.15.1.3.1 | 576758 | Cause | Unexpected arbitrary movement | 4.15.1.3.1.1 | 576762 | Measure | Reference - Prevent unintentional activation of rescue function by user error | Continuous activation <XREF>564146</XREF>
338 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.3 | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.15.1.3.2 | 576761 | Cause | Rescue function is activated during collision - collision detection ""non-functional"" | 4.15.1.3.2.1 | 576763 | Measure | Reference - Prevent unintentional activation of rescue function by user error | Continuous activation <XREF>564146</XREF>, as soon as the button is released, the brakes are activated again (other than non-motorized design with manual brake activation).
339 | 4.15 | 563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1 | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 4.15.1.3 | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.15.1.3.2 | 576761 | Cause | Rescue function is activated during collision - collision detection ""non-functional"" | 4.15.1.3.2.2 | 576764 | Measure | Reference - Warning - Correct use of rescue function | <XREF>564148</XREF>
340 | 4.18 | 569262 | Chapter | Acquire images | 4.18.1 | 563444 | Function | Acquire images - general | 4.18.1.1 | 580656 | Hazard | Unnecessary radiation, dose not as desired | 4.18.1.1.1 | 580654 | Cause | Radiation is switched off later than required in regulatory limits (IEC 60601-2-54:2018), normal conditions | 4.18.1.1.1.1 | 580655 | Measure | No additional measures required | See feature 505763.
341 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.1 | 568783 | Cause | Wireless control comes into reach and radiation button is activated or technical error in radiation release circuit | 4.18.2.1.1.1 | 562292 | Measure | Block radiation when controls connecting | The system shall check the status of both radiation release signal paths when controls are connected to the system. When an active radiation release signal path is detected, radiation from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths takes place.
342 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.2 | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 4.18.2.1.2.1 | 596992 | Measure | Labeling wireless hand control radiation | CAUTIONCause: Radiation released from aÂ wireless hand controlÂ in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy:Â As part of the system installation, the wirelessÂ hand controlÂ is labeled according to which system it belongs to. When utilizingÂ theÂ wireless hand control, check thatÂ it belongs to the systemÂ according to theÂ indicating labels.
343 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.2 | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 4.18.2.1.2.2 | 596634 | Measure | Optical indication of system and hand controls radiation | In Service DocumentationãCAUTIONCause: Radiation released from aÂ remote hand controlÂ in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy: Before performing any service activities with the remote hand control, check that the remote hand control belongs to the system and check that the indicating label is in place and properly visible.
344 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.2 | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 4.18.2.1.2.3 | 600718 | Measure | Reference pairing wireless hand control | <XREF>564224</XREF>
345 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.2 | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 4.18.2.1.2.4 | 600719 | Measure | Reference Optical indication of system and hand control pairs | <XREF>564225</XREF>
346 | 4.18 | 569262 | Chapter | Acquire images | 4.18.2 | 569260 | Function | Acquire images via wireless hand control | 4.18.2.1 | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.2.1.2 | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 4.18.2.1.2.5 | 600720 | Measure | Reference Label wireless hand control | <XREF>596993</XREF>
347 | 4.18 | 569262 | Chapter | Acquire images | 4.18.3 | 569279 | Function | Acquire images via wired footswitch | 4.18.3.1 | 569277 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 4.18.3.1.1 | 569278 | Cause | Footswitch gets damaged due to ingress of fluid | 4.18.3.1.1.1 | 568819 | Measure | IP protection class footswitch | The wireless and the wired footswitch shall have protection class IPx8 according to EN 60529 with closed battery compartment.
348 | 4.18 | 569262 | Chapter | Acquire images | 4.18.4 | 563447 | Function | Acquire Roadmap images | 4.18.4.1 | 591776 | Hazard | Unnecessary exposure to limited levels of ionizing radiation (HS_061) | 4.18.4.1.1 | 591777 | Cause | Position mismatch between subtraction and roadmap phases | 4.18.4.1.1.1 | 591778 | Measure | No system movement during or between subtraction and roadmap phases | Case 1: Movement during radiation in subtraction or roadmap mode:When a movement of any system axle is detected during any phase of acquiring an image in subtraction or roadmap mode,Â the workflow shall be aborted, radiation shall be stopped. Â Case 2: Movement between acquired peak opacification image within a subtraction workflow and roadmap acquisition:When a movement of any system axle is detected between an acquired peak opacification image within a subtractionÂ and a plannedÂ roadmap acquisition, the system shall discard this outdated peak opacification image (not allow roadmap), but instead, shall start the subtraction workflow automatically.
349 | 4.18 | 569262 | Chapter | Acquire images | 4.18.4 | 563447 | Function | Acquire Roadmap images | 4.18.4.1 | 591776 | Hazard | Unnecessary exposure to limited levels of ionizing radiation (HS_061) | 4.18.4.1.1 | 591777 | Cause | Position mismatch between subtraction and roadmap phases | 4.18.4.1.1.2 | 591779 | Measure | Warning - Prevent movements during roadmap workflow | WARNINGÂ Cause: Movements of the patient, of the system, or deformations of the vessel tree between acquiring theÂ mask imageÂ and using it as mask in the RoadmapÂ workflow. Â Possible consequences: Treatment at the wrong location. Remedy:-Â Prevent the movement of the patientâs relevant body region between acquiring the mask and using it as a mask in the RoadmapÂ workflows. - Check whether the device position in the image is matching with the vessel tree. - In case of doubt, visualize the actual vessel position by means of a short contrast injection and check it versus the vessel tree in the mask.
350 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.1 | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.18.5.1.1 | 564090 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 4.18.5.1.1.1 | 564091 | Measure | Reference - Braking force | <XREF>564086</XREF>
351 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.1 | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.18.5.1.1 | 564090 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 4.18.5.1.1.2 | 564092 | Measure | Reference - Automatic brake test (selftest) | <XREF>564087</XREF>
352 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.1 | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.18.5.1.2 | 564102 | Cause | Operator moves chassis (motorized) or other motorized axle during collision check or shortly after collision check | 4.18.5.1.2.1 | 564103 | Measure | Repeat collision check | Motorized movements other than orbital axle of the system during collision check or shortly after completion of the collision check shall be blocked. Note:Â If the user wants to perform any other motorized movement, the user has to actively abort the collision check / 3D scan workflow.
353 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.1 | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.18.5.1.3 | 564105 | Cause | Other motorized movement is triggered by technical fault during 3D scan | 4.18.5.1.3.1 | 564106 | Measure | Block other axles during 3D scan | During 3D scans, the system shall block all other motorized movements.
354 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.2 | 564097 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.5.2.1 | 564098 | Cause | Collision sensor fails or is activated during 3D scan, scan is aborted, restart possible | 4.18.5.2.1.1 | 564101 | Measure | Collision detection during 3D | Hint: Collision handling during 3D is equivalent to normal collision detection. An unwanted interruption of a 3D scan via wheel-based sensors is assessed to be highly improbable because during 3D the wheel-base is blocked and no operator staff is in proximity of the C-Arm. Therefore, it does not need to be addressed with further measures. Â <XREF>563490</XREF>
355 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.3 | 564093 | Hazard | Severely reduced image quality (HS_120) | 4.18.5.3.1 | 564094 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 4.18.5.3.1.1 | 564095 | Measure | Reference - Braking force | <XREF>564086</XREF>
356 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.3 | 564093 | Hazard | Severely reduced image quality (HS_120) | 4.18.5.3.1 | 564094 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 4.18.5.3.1.2 | 564096 | Measure | Reference - Automatic brake test (selftest) | <XREF>564087</XREF>
357 | 4.18 | 569262 | Chapter | Acquire images | 4.18.5 | 563448 | Function | Acquire 3D images | 4.18.5.4 | 578312 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 4.18.5.4.1 | 578313 | Cause | Movement stops during acquisition | 4.18.5.4.1.1 | 578315 | Measure | Stop radiation when movement is stopped | When the system movement stops for any reason (e.g. collision)Â during a 3D or a panorama scan, radiation shall be stopped.
358 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.1 | 564272 | Hazard | Severely reduced image quality (HS_120) | 4.18.6.1.1 | 564273 | Cause | Wrong stiching of images: Increased mechanical tolerances, compromised position resolution, vibrations (systematic) | 4.18.6.1.1.1 | 564274 | Measure | Trajectory accuracy | The requiredÂ trajectory accuracy shall be reached, which is necessary to provide panorama images with sufficient quality and geometric fidelity. The travel range for this application shall be restricted so that this value is reached. Â Note: It must be ensured, that the trajectory accuracy is sufficient to be able to perform image based corrections.
359 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.1 | 564272 | Hazard | Severely reduced image quality (HS_120) | 4.18.6.1.1 | 564273 | Cause | Wrong stiching of images: Increased mechanical tolerances, compromised position resolution, vibrations (systematic) | 4.18.6.1.1.2 | 564291 | Measure | Reduce vibrations during panorama | The system shall perform the movement for panorama image acquisition with a vibration free drive profile.
360 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.1 | 564272 | Hazard | Severely reduced image quality (HS_120) | 4.18.6.1.2 | 564275 | Cause | Wrong stiching of images: Source to object information is not known | 4.18.6.1.2.1 | 564276 | Measure | No additional measure required | The operator has a correction factor to change the source to object information before and after scan (reconstruction can be repeated). Default is isocenter.
361 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.1 | 564272 | Hazard | Severely reduced image quality (HS_120) | 4.18.6.1.3 | 564285 | Cause | Wrong synchronization at panorama imaging: Technical error | 4.18.6.1.3.1 | 564286 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
362 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.1 | 564272 | Hazard | Severely reduced image quality (HS_120) | 4.18.6.1.4 | 564288 | Cause | Wrong stiching of images: Software error | 4.18.6.1.4.1 | 564289 | Measure | Save all images panorama | All single 2D images of a panorama scan shall be stored. Â Note: Multiple reconstructions can be performed, as well as review of single slices.
363 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.1 | 564278 | Cause | Operator initates panorama imaging on any (outdated) image data | 4.18.6.2.1.1 | 564280 | Measure | Prerequisites panorama imaging | The acquisition of a panorama image shall only be possible when- a patient is registered, an image has been acquired on which the planning took place (path planning must have taken place and been completedÂ in standstill)- when the system has not been moved and no image rotation or flip took place between path planning / confirmation and movement activation- no collimation error (collimator must be able to reach the correct slot positions for the panorama slices)- enough storage capacity must be available for the maximum scan length- tube temperature must be sufficiently low to be able to completeÂ the maximum scan length
364 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.2 | 569290 | Cause | Vibrations during scan loosens connectors | 4.18.6.2.2.1 | 596956 | Measure | Reference - class_shock_vibration_operation | <XREF>568836</XREF>
365 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.3 | 564281 | Cause | Wrong synchronization at panorama imaging: Technical error | 4.18.6.2.3.1 | 564283 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
366 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.4 | 564282 | Cause | Wrong synchronization at panorama imaging: External influences (slip, obstacles,...) | 4.18.6.2.4.1 | 564290 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
367 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.5 | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 4.18.6.2.5.1 | 596954 | Measure | Reference - movement check after service/maintenance | <XREF>568847</XREF>
368 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.5 | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 4.18.6.2.5.2 | 596955 | Measure | Reference - check_movement_manufacturing | <XREF>558848</XREF>
369 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.5 | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 4.18.6.2.5.3 | 569288 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
370 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.5 | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 4.18.6.2.5.4 | 569289 | Measure | Reference - Mechanical design drive and wheels | <XREF>561959</XREF>
371 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.2 | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 4.18.6.2.5 | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 4.18.6.2.5.5 | 599815 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
372 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.1 | 564294 | Measure | Reference - Movement preview for automatic movements | <XREF>563648</XREF>
373 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.2 | 568854 | Measure | Restricted movement range for panorama imaging | The movement range for panorama imaging shall be restricted to 110 cm.
374 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.3 | 564301 | Measure | Reference - Timeout for automatic movement selection | <XREF>564193</XREF>
375 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.4 | 564295 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
376 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.5 | 564296 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
377 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.6 | 564297 | Measure | Reference - System speed limits | <XREF>563486</XREF>
378 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.7 | 564298 | Measure | Reference - Collision detection | <XREF>563490</XREF>
379 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.8 | 564299 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
380 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.1 | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsâ¦) | 4.18.6.3.1.9 | 564300 | Measure | Reference - Warning - Collisions during automatic movements without test run | <XREF>563650</XREF>
381 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.2 | 564302 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 4.18.6.3.2.1 | 564303 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
382 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.3 | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.18.6.3.3.1 | 564305 | Measure | Reference - Main Unit movement shock detection | <XREF>563684</XREF>
383 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.3 | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.18.6.3.3.2 | 599816 | Measure | Reference - Positioning accuracy | <XREF>564155</XREF>
384 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.3 | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.18.6.3.3.3 | 564306 | Measure | Reference - no movement takeover across controls | <XREF>569307</XREF>
385 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.3.3 | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 4.18.6.3.3.4 | 594829 | Measure | Reference - only one active control element | <XREF>569307</XREF>
386 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.4 | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 4.18.6.4.1 | 577223 | Measure | Reference - Set speed values | The system sets the movement speeds according to different modes and control, see <XREF>569146</XREF>, and in order to meet the overtravel limits <XREF>562465</XREF>ãand to not exceed the speed limits <XREF>563486</XREF> in normal conditions.
387 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.4 | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 4.18.6.4.2 | 577225 | Measure | Reference - Collision detection | <XREF>563490</XREF>
388 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.4 | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 4.18.6.4.3 | 577226 | Measure | Reference - System speed limits | <XREF>563486</XREF>
389 | 4.18 | 569262 | Chapter | Acquire images | 4.18.6 | 564271 | Function | Acquire Panorama images | 4.18.6.3 | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 4.18.6.4 | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 4.18.6.4.4 | 577224 | Measure | Reference - Emergency stop at every point of motorized movement activation | Multiple emergency stops are available at the system <XREF>562345</XREF>.
390 | 4.19 | 563461 | Chapter | Display acquired image(s) | 4.19.1 | 569299 | Function | Display acquired image(s) | 4.19.1.1 | 569300 | Hazard | Slightly reduced image quality (HS_121) | 4.19.1.1.1 | 569301 | Cause | Impaired displayed image quality due to dirt on the touch monitor | 4.19.1.1.1.1 | 569148 | Measure | Reference - Cleaning of the touch monitor | Cleaning of the monitor is covered in cleaning guide. Reference to Caution to follow the cleaning instruction <XREF>562240</XREF>. The monitor must withstand the procedure. RegardingÂ DIN 6868-157: 5 c) Before starting these tests, the viewing surface of the IMAGE DISPLAY DEVICE is to be cleaned in accordance with operating instructions, if needed.
391 | 4.25 | 563436 | Chapter | Save images | 4.25.1 | 569280 | Function | Save images | 4.25.1.1 | 569282 | Hazard | Image data is unavailable or lost (HS_130) | 4.25.1.1.1 | 569283 | Cause | Storage rate of images during acquisition is reduced | 4.25.1.1.1.1 | 594856 | Measure | Reference - full storage rate radiography | <XREF>578317</XREF>
392 | 4.27 | 563463 | Chapter | Perform 3D navigation | 4.27.1 | 569284 | Function | Perform 3D navigation | 4.27.1.1 | 569285 | Hazard | Contact with high level electrical energy (HS_002) | 4.27.1.1.1 | 569286 | Cause | Damaged 3rd party devices are connected to the system | 4.27.1.1.1.1 | 568831 | Measure | Electrical safety external interfaces | Galvanic separation (voltage withstand 4 kV) shall be provided for the following external interfaces: * HDMI * Network interface * Injector interface
393 | 4.27 | 563463 | Chapter | Perform 3D navigation | 4.27.1 | 569284 | Function | Perform 3D navigation | 4.27.1.1 | 569285 | Hazard | Contact with high level electrical energy (HS_002) | 4.27.1.1.1 | 569286 | Cause | Damaged 3rd party devices are connected to the system | 4.27.1.1.1.2 | 568833 | Measure | Warning - Electrical safety USB | WARNINGCause: Connect damaged component with external power supply to USB interface. Possible consequences: Electric shock, increased leakage current. Remedy: Only connect USB sticks, external hard drives without external power supply or other approved components to the system.
394 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.1 | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.36.1.1.1 | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 4.36.1.1.1.1 | 568824 | Measure | Caution - Close mecanum cover after cleaning | CAUTIONCause: Incorrect cleaning of wheels. Risk: Increased risk of crushing e.g. , ofÂ fingers, risk of damage to the system. Remedy: Clean the wheels according to the instructions. Disconnect the C-arm chassis from the trolley. Do not move the system while a person is cleaning wheels. Before opening the wheel covers, activate the lock movement function at the chassis handle panel. Store wireless hand control in drawer and wired hand control in holder. Â Close the wheel covers after cleaning activities are finished.
395 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.1 | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.36.1.1.1 | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 4.36.1.1.1.2 | 571852 | Measure | Reference - Lock all motorized movements | <XREF>564347</XREF>
396 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.1 | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.36.1.1.1 | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 4.36.1.1.1.3 | 580832 | Measure | Label - Hand injury mecanum | Labels for hand crushing shall be placed near each mecanum wheel. They shall be visible when the covers for cleaning are open.
397 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.1 | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.36.1.1.2 | 568823 | Cause | Unintended movement activation during cleaning activities of the mecanum wheels | 4.36.1.1.2.1 | 578106 | Measure | Reference - Warning - Close mecanum cover after cleaning | <XREF>568824</XREF>
398 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.1 | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 4.36.1.1.2 | 568823 | Cause | Unintended movement activation during cleaning activities of the mecanum wheels | 4.36.1.1.2.2 | 578107 | Measure | Reference - Lock all motorized movements | <XREF>564347</XREF>
399 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.2 | 568826 | Hazard | Contact with high level electrical energy (HS_002) | 4.36.1.2.1 | 568827 | Cause | Contact with electrical energy during mecanum wheel cleaning (e. g. ingress of cleaning agent) | 4.36.1.2.1.1 | 568828 | Measure | Protection against ingress of fluids - mecanum | The mecanum wheel electronics shall be protected against ingress of fluids with IPX4 according to IEC 60529. Â Hint: This applies to the electric drives. The mechanical rollers on the wheels themselves do not have to fulfill this standard.
400 | 4.36 | 563468 | Chapter | Clean the system | 4.36.1 | 568820 | Function | Clean the mecanum wheels | 4.36.1.2 | 568826 | Hazard | Contact with high level electrical energy (HS_002) | 4.36.1.2.1 | 568827 | Cause | Contact with electrical energy during mecanum wheel cleaning (e. g. ingress of cleaning agent) | 4.36.1.2.1.2 | 578318 | Measure | Reference - IEC 60601-1 Clause 8.7 and 8.8 | <XREF>562365</XREF>
401 | 4.37 | 569267 | Chapter | Transport the system | 4.37.1 | 569269 | Function | Transport the system | 4.37.1.1 | 569268 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 4.37.1.1.1 | 569270 | Cause | System gets damaged due to transport (e. g. vibrations, rough handling) | 4.37.1.1.1.1 | 568837 | Measure | No additional measure required | The classification for mechanical shock and vibrationÂ during transportÂ according to IEC 60721-3-2 is sufficient (2M4, except for free fall 2M5), see <XREF>561868</XREF>. No additional risk control measures are required.
402 | 4.37 | 569267 | Chapter | Transport the system | 4.37.1 | 569269 | Function | Transport the system | 4.37.1.2 | 569271 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 4.37.1.2.1 | 569272 | Cause | System gets damaged due to transport (e. g. vibrations, rough handling) | 4.37.1.2.1.1 | 569273 | Measure | Reference - Classification for mechanical shock and vibration - transport | The classification for mechanical shock and vibration during transport according to IEC 60721-3-2 is sufficient (2M4, except for free fall 2M5), see <XREF>561868</XREF>. No additional risk control measures are required.
403 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.1 | 561954 | Hazard | Pinching through breakage of supporting parts | 4.42.1.1.1 | 561955 | Cause | Breakage of welded joints or wheels | 4.42.1.1.1.1 | 561956 | Measure | Design of the welded joints | The welded joints or wheels must be designed in a way so that an overload can be carried in compliance with DIN EN 60601-1 (2006), chap. 9. 8. , table 21
404 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.2 | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 4.42.1.2.1 | 561961 | Cause | Brake failure | 4.42.1.2.1.1 | 561962 | Measure | User Manual - Brake failure | CAUTIONCause: Brake failurePossible consequences: Risk of crushingRemedy:Â Before beginning the examination, perform the daily function and safety checks.
405 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.2 | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 4.42.1.2.2 | 561963 | Cause | Braking power not sufficient | 4.42.1.2.2.1 | 561964 | Measure | Reference - Design of the brakes | <XREF>562436</XREF>
406 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.2 | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 4.42.1.2.2 | 561963 | Cause | Braking power not sufficient | 4.42.1.2.2.2 | 561965 | Measure | User Manual - Parking Trolley on inclined Plane (i.e. Ramp) | CAUTIONCause: Insufficient braking force. Possible consequences: Risk of crushing!Remedy:-Â When parking the monitor trolley on an inclined surface, face the front toward the incline and lock the brakes. -Â Make sure the monitor trolley is not parked on a wet floor or other places where the adhesion of the wheels is greatly reduced.
407 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.2 | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 4.42.1.2.2 | 561963 | Cause | Braking power not sufficient | 4.42.1.2.2.3 | 561966 | Measure | Warning Label - Parking Position Trolley | Place warning label (3868681) next to foot brake.
408 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.3 | 561967 | Hazard | Risk of injury through damaged components | 4.42.1.3.1 | 561968 | Cause | Crossing a step | 4.42.1.3.1.1 | 561969 | Measure | Reference - Rough handling test trolley | <XREF>561951</XREF>
409 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.4 | 561970 | Hazard | Crushing when adjusting heights of monitors | 4.42.1.4.1 | 561971 | Cause | Motor-driven movement | 4.42.1.4.1.1 | 561972 | Measure | Distance in case of fault according to IEC 60601-1 | Compliance with DIN EN 60601-1 (2006) Kap. 9. 2; Table 20, Acceptable gaps for finger > 25mmNote: the monitor movements are not considered by the emergency stop
410 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.4 | 561970 | Hazard | Crushing when adjusting heights of monitors | 4.42.1.4.1 | 561971 | Cause | Motor-driven movement | 4.42.1.4.1.2 | 561973 | Measure | Attention note in user manual | CAUTIONCause:Â Motorized movements. Possible consequences: Risk of crushingRemedy:Â Do not stand in the travel range of the monitor or place objects on the trolley housing.
411 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.5 | 561974 | Hazard | Electric shock | 4.42.1.5.1 | 561975 | Cause | Protective earth broken | 4.42.1.5.1.1 | 597040 | Measure | Reference Cable protection | <XREF>562104</XREF> and <XREF>562103</XREF>
412 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.1 | 561940 | Function | Move Trolley into OR | 4.42.1.6 | 561980 | Hazard | Injury through dropping parts | 4.42.1.6.1 | 561981 | Cause | During transportation not secured parts may drop | 4.42.1.6.1.1 | 561982 | Measure | Built-in parts have to be secured because of the risk for dropping | Built-in parts have to be secured because of the risk for dropping
413 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.1 | 561984 | Hazard | Delayed or no interventional decision | 4.42.2.1.1 | 561985 | Cause | Untrained users / Missing knowledge of the system | 4.42.2.1.1.1 | 561986 | Measure | User Manual - Trained Users | WARNINGCause: Operation of the system by non-trained usersPossible consequences: Misinterpretation of image information or of indications of the system, leading to inappropriate decisions in the clinical workflow, incorrect diagnosis or treatment, delay of the clinical procedure, or exposure to radiation without clinical benefit. Remedy: The system must only be used by persons with the necessary knowledge, after an appropriate application training.
414 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.1 | 561984 | Hazard | Delayed or no interventional decision | 4.42.2.1.2 | 561987 | Cause | Not intended use of system | 4.42.2.1.2.1 | 561988 | Measure | Reference - Operator Manual - Not intended use | <XREF>562190</XREF>
415 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.2 | 561989 | Hazard | Loss of functionality | 4.42.2.2.1 | 561990 | Cause | Complexity of the system (component failures), mains supply failure or voltage out of specification | 4.42.2.2.1.1 | 561991 | Measure | Reference to User Manual - Establish emergency procedures | <XREF>562205</XREF>
416 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.2 | 561989 | Hazard | Loss of functionality | 4.42.2.2.2 | 561992 | Cause | (Cooling) Liquid is stored in the trolley and leaks, damaging the electronics | 4.42.2.2.2.1 | 561993 | Measure | No fluid storage in trolley | Service DocumentsWARNINGCause: Fluids are stored (e.g. cooling liquid) in the monitor trolley and leaks, destroying electronics. Possible Consequences: Damage of components in the trolley, possibly delay or interruption of the procedure. Remedy: Do not store fluids like the Service Pack with cooling liquid in the monitor trolley. Â Note:This is relevant to systems with cooling option. Nevertheless, storing of liquids inside of trolley is prohibited for all systems.
417 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.2 | 561989 | Hazard | Loss of functionality | 4.42.2.2.3 | 561994 | Cause | Foreseeable misuse | 4.42.2.2.3.1 | 561995 | Measure | Reference - OperatorManual-Trained Users | <XREF>561986</XREF>
418 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.2 | 561989 | Hazard | Loss of functionality | 4.42.2.2.4 | 561996 | Cause | Some parts in system could lead to a hazard by wear | 4.42.2.2.4.1 | 561997 | Measure | Reference - OperatorManual-Abrasion of wear | <XREF>562201</XREF>
419 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.1 | 561999 | Cause | Overheating and Flammability Materials | 4.42.2.3.1.1 | 562000 | Measure | Reference - IEC 60601-1 Standards | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
420 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.1 | 561999 | Cause | Overheating and Flammability Materials | 4.42.2.3.1.2 | 562001 | Measure | Reference - IT equipment safety | Compliance withÂ IEC 62368-1, 04S452203 and 04S451630 in <XREF>561856</XREF>
421 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.1 | 561999 | Cause | Overheating and Flammability Materials | 4.42.2.3.1.3 | 562002 | Measure | Reference - User Manual - Switch off the system in the event of fire User Manual | <XREF>562218</XREF>
422 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.2 | 562003 | Cause | Auto-ignition or Melting of Material | 4.42.2.3.2.1 | 562004 | Measure | No measure required | Result of risk evaluation: Temperature inside of trolley will never reach such threshold to enable auto-ignition or any melting of material. Therefore, ho risk control measures are necessary. Note:System specific documents (e.g. protocols etc. , i.e. blue "Service-Folder") may be stored inside of trolley. ThisÂ wish come up as system specific documents needs to be on-site during product life time.
423 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.3 | 562005 | Cause | Explosion | 4.42.2.3.3.1 | 562006 | Measure | User Manual - Danger of explosion | WARNINGCause: The system is not designed for operation in explosion-endangered areas. It does not fulfill the requirements for AP/APG classification. Possible consequences: Danger of explosionRemedy: The system may not be used in locations where explosive atmospheres may occur.
424 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.3 | 561998 | Hazard | Hazards from fire or smoke | 4.42.2.3.4 | 562007 | Cause | Too high temperature of the detector | 4.42.2.3.4.1 | 562008 | Measure | Reference - System Warning - Temperature of the Detector | See <XREF>562226</XREF>.
425 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.4 | 562009 | Hazard | Chemical hazards | 4.42.2.4.1 | 562010 | Cause | Toxic materials and fluids | 4.42.2.4.1.1 | 562011 | Measure | Reference - Technical instruction for the disposal procedure | <XREF>562246</XREF>
426 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.4 | 562009 | Hazard | Chemical hazards | 4.42.2.4.1 | 562010 | Cause | Toxic materials and fluids | 4.42.2.4.1.2 | 562012 | Measure | Service Instruction - Toxic Fluid | CAUTIONCause: Cooling liquid (such asÂ Antifrogen N) is not handled correctly. Possible consequences: Skin or eye irritation, irritation of the breathing passages,Â toxic or allergic reactions caused by contact with hazardous substances. Remedy: Avoid skin contact, inhalation and swallowing, use personal protective equipment such as gloves and goggles.
427 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.5 | 562013 | Hazard | electromagnetic disturbances | 4.42.2.5.1 | 562014 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 4.42.2.5.1.1 | 562016 | Measure | User Manual - Degradation of essential performance by electromagnetic interferences | CAUTIONCause:Â Electromagnetic interference of the system. Possible consequences: Degradation of the image quality or accuracy of the radiation parameters!Remedy:Â - Remove the source of interference (for instance, a mobile phone) from the area. - Shut down the system and turn it on again. - Pay attention to error messages during restart.
428 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.5 | 562013 | Hazard | electromagnetic disturbances | 4.42.2.5.1 | 562014 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 4.42.2.5.1.2 | 562017 | Measure | Reference to generator measure for IEC 60601-1-2 | See HM_S1_mec_F1G1U1M1 andÂ HM_S1_emv_F1G1U2M1Â in <XREF>578770</XREF>.
429 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.6 | 562018 | Hazard | Hearing impairment of patients and/or staff | 4.42.2.6.1 | 562019 | Cause | Excessive audible noise | 4.42.2.6.1.1 | 562020 | Measure | No measure required | The acoustic energy / audible noise is limited via IEC60601-1:2005, chapter 9. 6. 2. 1. Therefore, no additional risk control measures are necessary. Â has to be smaller/equal 80 dB(A).
430 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.7 | 562021 | Hazard | Injuries due to monitor dropping | 4.42.2.7.1 | 562022 | Cause | Breakage of the monitor support | 4.42.2.7.1.1 | 562023 | Measure | Compliance with DIN EN 60601-1 | The monitor support must be designed in a way that an overload can be carried. Â Remark:Compliance with DIN EN 60601-1 (2006), Kap. 9. 8 Table 21
431 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.8 | 562024 | Hazard | Risk of injury through impaired parts | 4.42.2.8.1 | 562025 | Cause | Transport damages e.g. through collision | 4.42.2.8.1.1 | 562026 | Measure | Protection of monitor during transport | The monitor shall be robust against mechanicalÂ impacts from all sides:A cover shall protect the monitor on all sides, except the front side from impact of collisions in order to meet at least IEC 60601-1 cl. 15. 3. 3.
432 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.8 | 562024 | Hazard | Risk of injury through impaired parts | 4.42.2.8.1 | 562025 | Cause | Transport damages e.g. through collision | 4.42.2.8.1.2 | 596969 | Measure | Robust touch glass of monitor | The touch glass of the monitor shall be robust against mechanical impacts (e.g. at least meeting IEC 60601-1 cl. 15. 3. 3). If the touch glass is broken due to an impact, it shall not splinter or expellÂ shards.
433 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.1 | 562028 | Cause | High bleeder resistor | 4.42.2.9.1.1 | 562029 | Measure | Leakage current measurement | Leakage current measurement during final inspection
434 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.1 | 562028 | Cause | High bleeder resistor | 4.42.2.9.1.2 | 562030 | Measure | Work step in maintenance documents | Work stepÂ in maintenance documents: operate leakage current measurement after maintenance work
435 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.1 | 562028 | Cause | High bleeder resistor | 4.42.2.9.1.3 | 562031 | Measure | Attention note in service manual | Attention note in the relevantÂ service manual: Operate leakage current measurement after service.
436 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.2 | 562032 | Cause | Interruption of the protective earth | 4.42.2.9.2.1 | 562033 | Measure | Protective earth measurement | Protective earth measurement during final inspection
437 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.2 | 562032 | Cause | Interruption of the protective earth | 4.42.2.9.2.2 | 562034 | Measure | Attention note in service manual | Attention note in service manual: Perform protective earth measurement after service work.
438 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.2 | 562032 | Cause | Interruption of the protective earth | 4.42.2.9.2.3 | 562035 | Measure | Work step in service instructions | Work step in maintenance documents: Perform protective earth measurement after maintenance work.
439 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.3 | 562036 | Cause | Mechanically impaired cables/wiring | 4.42.2.9.3.1 | 562037 | Measure | Cable conduit for movable cables | Moving components such like cables, plugs, screws which are located in the proximity of isolators or wires may come into contact with voltages and increase the risk of fire or electric shocks. Mitigation is performed by fixing all moving wires, plugs connectors, etc. with means such like cable conduit, cable straps, plastic bridges or screws, glue, heat shrink tubes, etc. to avoid contacts caused by attrition. This measure is in accordance to IEC60601-1:2005 chap. 8. 10. Â Note:Movable components are affected by e.g. adjustment of monitor height orÂ horizontal movement of C-arm (but not limited to). In addition, this shall address service workflows as well (e.g. pulling components from its defined position in system during maintenance or replacement).
440 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.4 | 562038 | Cause | Static charge | 4.42.2.9.4.1 | 562039 | Measure | Create conductance on the trolley and main unit | In order to prevent electrostatic charging (e.g. when the system is moved over the floor), the system shall provideÂ a conductive connection to the ground, e.g. conductance band, conductive wheels.
441 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.4 | 562038 | Cause | Static charge | 4.42.2.9.4.2 | 562040 | Measure | Work step in maintenance documents | Work step in maintenance documents: Scanning of the conducting elements for damage and fouling
442 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.5 | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 4.42.2.9.5.1 | 562042 | Measure | Reference to compliance to IEC 60601-1
443 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.5 | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 4.42.2.9.5.2 | 562043 | Measure | Reference to Service Documentation - Trained service personnel
444 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.9 | 562027 | Hazard | Electric shock | 4.42.2.9.5 | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 4.42.2.9.5.3 | 562044 | Measure | Caution in Service Documentation - Discharge capacitor | Service DocumentationCAUTIONCause: Replacing the fuses F1 or F2 on D819 without discharging the capacitor C107Possible Consequences: Electrical shock due to capacitor dischargeRemedy: Discharge capacitor C107 on D819 by connecting a multimeter for 10 minutes to X34 and X35.
445 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.10 | 562045 | Hazard | Risk of injury due to bad surface finish | 4.42.2.10.1 | 562046 | Cause | Sharp edges | 4.42.2.10.1.1 | 562047 | Measure | Develop design without sharp edges | Develop design without sharp edges. (referring to IEC60601-1:2005 chapter 9. 3. )
446 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.10 | 562045 | Hazard | Risk of injury due to bad surface finish | 4.42.2.10.1 | 562046 | Cause | Sharp edges | 4.42.2.10.1.2 | 562048 | Measure | Reference to measures | HM_S1_mec_F4G2U1M1 in <XREF>578770</XREF>.
447 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.10 | 562045 | Hazard | Risk of injury due to bad surface finish | 4.42.2.10.2 | 562049 | Cause | Rough surface finish | 4.42.2.10.2.1 | 562050 | Measure | Specify smooth surface finishes | Specify smooth surface finishes(referring to IEC60601-1:2005 chapter 9. 3)
448 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.11 | 562052 | Hazard | Effect on people, e.g. tissue injury, reddened skin | 4.42.2.11.1 | 562053 | Cause | Electromagnetic radiation | 4.42.2.11.1.1 | 562054 | Measure | No measure required | The permissible electromagnetic emissions of the systemÂ and its components (e. g. Â WLAN) fulfill the limits specified inÂ IEC 60601-1-2. Therefore, no further risk control measures are necessary.
449 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.12 | 562055 | Hazard | Effect on people with implants e.g. pacemaker | 4.42.2.12.1 | 562056 | Cause | Interference into the implant\s electronics | 4.42.2.12.1.1 | 562058 | Measure | User Manual - Interference into Implants | CAUTIONCause: InterferencePossible consequences: Effect on people with implantsRemedy:Â Maintain a minimum distance in accordance with IEC 60601-1-2. If this is not possible, switch off the WLAN client.
450 | 4.42 | 561939 | Chapter | Setup Trolley | 4.42.2 | 561983 | Function | Startup Trolley | 4.42.2.12 | 562055 | Hazard | Effect on people with implants e.g. pacemaker | 4.42.2.12.1 | 562056 | Cause | Interference into the implant\s electronics | 4.42.2.12.1.2 | 576774 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria | <XREF>562015</XREF>
451 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.1 | 562065 | Hazard | Crushing due to system tipping over | 4.43.1.1.1 | 562066 | Cause | C-arm movement shifts over-all center of gravity | 4.43.1.1.1.1 | 562067 | Measure | Design of system | The system must be designed in a way that prevents tipping over on inclined plane (i.e. ramp) at an angle of 5Â°. Â Remark:Compliance with DIN EN 60601-1 tilt test
452 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.1 | 562065 | Hazard | Crushing due to system tipping over | 4.43.1.1.2 | 562068 | Cause | Weight put on in addition changes over-all center of gravity | 4.43.1.1.2.1 | 562069 | Measure | User Manual - Weight Balancing | WARNING:Cause: Change of overall center of gravityPossible consequences: Risk of crushingRemedy: Only products/components approved by Siemens Healthineers may be installed.
453 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.1 | 562065 | Hazard | Crushing due to system tipping over | 4.43.1.1.3 | 562070 | Cause | Transport of the device over an inclined plane (i.e. ramp) | 4.43.1.1.3.1 | 562071 | Measure | Design of system | The system are designed in a way that prevents tipping over on inclined plane (i.e. ramp) at an angle of 10Â°. Â Remark:Compliance with DIN EN 60601-1 tilt test for transportation
454 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.1 | 562065 | Hazard | Crushing due to system tipping over | 4.43.1.1.4 | 562072 | Cause | Crossing a step | 4.43.1.1.4.1 | 562073 | Measure | Reference - Rough handling test (main unit) | <XREF>562087</XREF>
455 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.2 | 562074 | Hazard | Crushing through breakage | 4.43.1.2.1 | 562075 | Cause | Breakage of welded joints | 4.43.1.2.1.1 | 562076 | Measure | Reference - Design of the welded joints | <XREF>561956</XREF>
456 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.3 | 562077 | Hazard | Crushing through breakage of the wheels | 4.43.1.3.1 | 562078 | Cause | Carrying capacity of the wheels too small | 4.43.1.3.1.1 | 562079 | Measure | Reference - Design of the wheels | <XREF>561959</XREF>
457 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.4 | 562080 | Hazard | Crushing due to autonomous movements on inclined plane (i.e. ramp) | 4.43.1.4.1 | 562081 | Cause | Brake failure | 4.43.1.4.1.1 | 562082 | Measure | Reference - self test of the brakes | <XREF>564087</XREF>
458 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.4 | 562080 | Hazard | Crushing due to autonomous movements on inclined plane (i.e. ramp) | 4.43.1.4.2 | 562083 | Cause | Braking power not sufficient | 4.43.1.4.2.1 | 562084 | Measure | Reference - Specify sufficient braking power | <XREF>562436</XREF>
459 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.5 | 562085 | Hazard | Risk of injury through damaged components | 4.43.1.5.1 | 562086 | Cause | Moving the system over a doorstep, against a doorframe, wall or other equipment | 4.43.1.5.1.1 | 562087 | Measure | Rough handling test (main unit) | The chassis shall be designed to comply with IEC 60601-1Â cl. 15. 3. 5. Â Scenarios to be tested:Â - Transport position of the system reached and fast drive move activatedÂ for a) Ascending step shock, b) Descending step shock, c) Door frame shock and forward directionÂ (Note: During a) and c) the collision avoidance is activated and slows down the system to operating room speed. For b) fast drive mode speed is applied. )Â - Transport position of the system andÂ operating roomÂ speed for a), b), c) in forward and lateral direction (Note: During a) and c), this is expected to repeat the first test withÂ a) and c) with collision avoidance activated. b) is expected to be less critical than the first test with b) with transport speed).
460 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.5 | 562085 | Hazard | Risk of injury through damaged components | 4.43.1.5.1 | 562086 | Cause | Moving the system over a doorstep, against a doorframe, wall or other equipment | 4.43.1.5.1.2 | 594861 | Measure | Reference - steep ramp | <XREF>578343</XREF>
461 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.6 | 562090 | Hazard | Injury through damaged control element | 4.43.1.6.1 | 562091 | Cause | Control element falls on the floor during transportation | 4.43.1.6.1.1 | 562092 | Measure | Drop test | The portable control elements except Hand-Helds with masses <= 10 kg have to be designed so that they can be dropped from a height of 5 cm without having damages that cause unacceptable risks. Hand-Helds have to be designed to withstand drop from heights at which the ME EQUIPMENT is used (as specified in the ACCOMPANYING DOCUMENTS), or from a height of 1 m, whichever is greater, onto a 50 mm +/-5 mm thick hardwood board (hardwood > 600 kg/m3) lying flat on a concrete or a similar rigid base. Test criterion:Execute drop tests according to standard IEC 60601-1 at heights of 5 cm for portable mobile control panels or heights of >= 1m for Hand-Helds. Possible emerged damages have to be examined for hazards (e.g. access to electrical and moveable parts has to be prevented).
462 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.7 | 562093 | Hazard | Biological and Chemical Hazards | 4.43.1.7.1 | 562094 | Cause | Detector is damaged, service personnel comes into contact with CsI | 4.43.1.7.1.1 | 562096 | Measure | Supply Chain Documentation - Toxic Materials | Make note in the relevantÂ factory documents. How these parts are disposed in an environmentally friendly manner have to be documented in the disposal instructions. Cause: Toxic materialsPossible consequences: Toxic contamination in case of detector panel is broken by mechanical shockRemedy: Handle detector with care! If the detector is broken do not use it, inspect if any particles are spilled. Leaked particles have to be collected and kept in a sealed container. Spilled particles have to be returned to Siemens Healthineers.
463 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.7 | 562093 | Hazard | Biological and Chemical Hazards | 4.43.1.7.1 | 562094 | Cause | Detector is damaged, service personnel comes into contact with CsI | 4.43.1.7.1.2 | 562097 | Measure | Service Instruction - Toxic Materials | Make note in the relevantÂ service documents. How these parts are disposed in an environmentally friendly manner have to be documented in the disposal instructions. Â CAUTIONÂ Cause: Toxic material is accessible when the image intensifier / flat detector is damaged due to mechanical impact. Possible consequences: Skin or eye irritation, irritation of the breathing passages,Â toxic or allergic reactions caused by contact with hazardous substances. Remedy: Handle the image intensifier / flat detector with care. If it is damaged, do not use it. Inspect if any particles are spilled. Avoid skin contactÂ and inhalation of dust, use personal protective equipment if required. Spilled particles have to be collected and kept in a sealed container andÂ returned to Siemens HealthineersÂ for disposal.
464 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.7 | 562093 | Hazard | Biological and Chemical Hazards | 4.43.1.7.2 | 562098 | Cause | Detector not disposed correctly | 4.43.1.7.2.1 | 562099 | Measure | Service Instruction - Toxic Materials - Disposal | See <XREF>407319</XREF>.
465 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.7 | 562093 | Hazard | Biological and Chemical Hazards | 4.43.1.7.3 | 569205 | Cause | Detector is damaged, operator comes into contact with CsI | 4.43.1.7.3.1 | 562095 | Measure | User Manual - Detector Cover Broken | CAUTIONCause: Flat detector cover brokenPossible consequences: Skin or eye irritation, irritation of the breathing passages,Â toxic or allergic reactions caused by contact with hazardous substances. Remedy: If the flat detector cover is broken, do not touch the detector,Â do not use the system at all. Please call Siemens Healthineers service.
466 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.1 | 562064 | Function | Move Main Unit in OR | 4.43.1.7 | 562093 | Hazard | Biological and Chemical Hazards | 4.43.1.7.3 | 569205 | Cause | Detector is damaged, operator comes into contact with CsI | 4.43.1.7.3.2 | 569206 | Measure | No additional measure required | 2 covers around hazardous substance provided.
467 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.1 | 562102 | Cause | Damage of the cable | 4.43.2.1.1.1 | 562103 | Measure | Certified mains power cable | The mains power supply cable shall be in accordance with IEC 60601-1 cl. 8. 11. 3. Country specific certificates may be used when equivalent.
468 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.1 | 562102 | Cause | Damage of the cable | 4.43.2.1.1.2 | 569245 | Measure | Robust trolley-chassis cable | Robust trolley-chassis cableÂ The system shall prevent damage to trolley-chassis cable from extensive tension caused by movements of the Main Unit. Note, a collision detection system is implemented which detects tension at the TC cable. Â The trolley chassis cable shall tolerate forces and bending up to the limit when tension detection is activated. <XREF>563740</XREF>
469 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.1 | 562102 | Cause | Damage of the cable | 4.43.2.1.1.3 | 562104 | Measure | Cable anchorage | All power cables have to be protected against cable damage to prevent electrical shock. In addition, it has to ensured, that system's connection to power line remains active as long as possible (a connection via socket outlet and plug is not sufficient). Â All power cables need to be secured with cable anchorage / strain-relief, e.g. by means of a strain-relief sleeve. Pull forces for cable retention and cable anchorage shall be in accordance to IEC62444-2010 (Table 2).
470 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.1 | 562102 | Cause | Damage of the cable | 4.43.2.1.1.4 | 562105 | Measure | User Manual - Damage of Cable | WARNING:Cause: Cable damagePossible consequences: Risk of electrical shockRemedy: Check the power cable. Do not use the device if the power cable is damaged.
471 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.2 | 562106 | Cause | Interruption of the protective earth | 4.43.2.1.2.1 | 562107 | Measure | Reference - Protective earth measurement | <XREF>562033</XREF>
472 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.2 | 562106 | Cause | Interruption of the protective earth | 4.43.2.1.2.2 | 562108 | Measure | Reference - Attention note in service instructions | <XREF>562034</XREF>
473 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.3 | 562109 | Cause | Excess current or short-circuit | 4.43.2.1.3.1 | 562110 | Measure | Fuse for the supply line | Fuses shall be provided in each supply lead
474 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.3 | 562109 | Cause | Excess current or short-circuit | 4.43.2.1.3.2 | 562111 | Measure | Reference - Warning label | <XREF>563226</XREF>
475 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.4 | 562112 | Cause | Static charge | 4.43.2.1.4.1 | 562113 | Measure | Reference - Create conductance on the trolley and main unit | <XREF>562039</XREF>
476 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.1 | 562101 | Hazard | Electric shock | 4.43.2.1.4 | 562112 | Cause | Static charge | 4.43.2.1.4.2 | 562114 | Measure | Reference - Work step in maintenance documents | <XREF>562040</XREF>
477 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.2 | 562115 | Hazard | Risk to fall | 4.43.2.2.1 | 562116 | Cause | Tripping over cables | 4.43.2.2.1.1 | 562117 | Measure | User Manual - Installing Cables | CAUTIONCause: Tripping over cablePossible consequences: Risk of fallingRemedy:Â When routing cables, make sure they run on the floor properly and without loops.
478 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.1 | 562119 | Cause | Overcharge of rechargeable batteries | 4.43.2.3.1.1 | 562120 | Measure | Provide protection from overcharge HICTouchCon | Provide protection from overcharge
479 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.1 | 562119 | Cause | Overcharge of rechargeable batteries | 4.43.2.3.1.2 | 562121 | Measure | Provide protection from overcharge HICSwitch | Provide protection from overcharge
480 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.1 | 562119 | Cause | Overcharge of rechargeable batteries | 4.43.2.3.1.3 | 562122 | Measure | Use approved batteries | Use approved batteries
481 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.2 | 562123 | Cause | Battery Gas Emission during Charge | 4.43.2.3.2.1 | 562124 | Measure | Batteries with Temperature Sensors | Use batteries with temperature sensors to enable interrupt of charging device. Interruption ofÂ charging willÂ prevent (toxic) gas emission.
482 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.3 | 562125 | Cause | Overcharge of the capacitors | 4.43.2.3.3.1 | 562126 | Measure | Mounted under covers | Capacitors are mounted under a dedicated cover.
483 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.3 | 562125 | Cause | Overcharge of the capacitors | 4.43.2.3.3.2 | 562127 | Measure | Reference - Protection of the energy storage | <XREF>562523</XREF>
484 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.3 | 562118 | Hazard | Excessive temperature, fire, explosion | 4.43.2.3.3 | 562125 | Cause | Overcharge of the capacitors | 4.43.2.3.3.3 | 562128 | Measure | Reference - Rupture joint by Overcharge and pressure relief valve | <XREF>562525</XREF><XREF>562524</XREF>
485 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.4 | 562129 | Hazard | Foreseeable missuse | 4.43.2.4.1 | 562130 | Cause | Misplugging of cables | 4.43.2.4.1.1 | 562131 | Measure | Design considering misplugging | The design of the external accessible connectors must consider misplugging of cables. Compare also chapter 16. 9 of IEC60601-1:2005.
486 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.1 | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 4.43.2.5.1.1 | 562134 | Measure | Check radiation release signal status at startup | The system shall check the status of both radiation release signal paths at startup of every control. When an active radiation release signal path is detected, radiation from the affected control (sharing one enable hardware line) shall be blocked and an error message shall be displayed, until a deactivation and regular re-activation of both radiation release paths take place.
487 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.1 | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 4.43.2.5.1.2 | 562135 | Measure | Emergency stop radiation | IfÂ an emergency stop <XREF>562345</XREF> is activated, radiationÂ shall be stopped, but not inhibited. When the operator activates radiation again, the system shall release radiation again. Â Note: TheÂ SysCB monitors the status of the emergency stop and provides this information to theÂ XCU.
488 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.1 | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 4.43.2.5.1.3 | 562136 | Measure | User Manual - Emergency Stop Button | CAUTIONCause: Because of a technical error the acquisition is not stopped or there is an unintended release of radiation. Possible consequences: Risk of unnecessary radiation exposureRemedy:Â In the event of unwanted radiation, press the nearest EMERGENCY STOP button.
489 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.1 | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 4.43.2.5.1.4 | 562137 | Measure | User Manual - Shutdown switch | CAUTIONCause: Because of a technical error the acquisition is not stopped or there is an unintended release of radiation. Possible consequences: Risk of unnecessary radiation exposureRemedy: Press the OFF button on the monitor trolley, even if the EMERGENCY STOP button has already been pressed. Radiation terminates immediately.
490 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.1 | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 4.43.2.5.1.5 | 562138 | Measure | Reference to radiation indication signal | See HM_S1_hw_F1G1U1M3 inÂ <XREF>578770</XREF>.
491 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.2 | 562139 | Cause | Radiation indicator fails | 4.43.2.5.2.1 | 568841 | Measure | Monitoring of radiation indicators | The radiation indicator circuit shall be monitored (e.g. current in indicator circuit when radiation is emitted). When an error is detected, the system shall show an appropriate error message (no blocking of radiation). Â Note, radiation indicators are provided on top of the trolley, on top of theÂ CCP and on top ofÂ the RCP, if connected to the system.
492 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.2 | 562139 | Cause | Radiation indicator fails | 4.43.2.5.2.2 | 562140 | Measure | Plug with stablized connection | Utilize plug with stabilized mechanical connection against disconnection.
493 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.2 | 562139 | Cause | Radiation indicator fails | 4.43.2.5.2.3 | 562141 | Measure | User Manual - Radiation Indicator | CAUTIONCause: Failure of radiation indicator. Possible consequences: Danger of unintended radiation!Remedy: If a radiation indicator fails contact Siemens Healthineers Customer Service.
494 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.5 | 562132 | Hazard | Unintentional Radiation | 4.43.2.5.3 | 562142 | Cause | Radiation indicator LED fails | 4.43.2.5.3.1 | 562143 | Measure | No measure required | Note:Lifetime of LED (approx. 100,000 hours) will exceed lifetime of C-arm system. In addition, radiation indicator is equipped with multiple LEDs,Â so indication of X-ray is possible while failure ofÂ a single LED. Therefore, no measure is required.
495 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.1 | 562145 | Cause | Prevent startup | 4.43.2.6.1.1 | 562146 | Measure | Reference - User Manual - Wear and tear | <XREF>562201</XREF>
496 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.2 | 562147 | Cause | Data loss because the System is not safely turned off | 4.43.2.6.2.1 | 562148 | Measure | User Manual - ShutDown during X-Ray | CAUTIONCause: Loss of image dataPossible consequences:Â Risk of unnecessary radiation exposure!Remedy: Do not shut down the system while radiation is being released.
497 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.3 | 562149 | Cause | System not in a proper condition e.g. error in a component | 4.43.2.6.3.1 | 562150 | Measure | Display system status | System status is visible to indicate the presence of an error.
498 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.4 | 562151 | Cause | User loads non-approved SW or add non approved HW to the system | 4.43.2.6.4.1 | 562152 | Measure | User Manual - Use of non-approved parts | WARNING:Cause: Use of unreleased or altered software or hardware components. Possible consequences: Risk of malfunction that could endanger the patient or product. Remedy:- Only use Siemens Healthineers-authorized software and hardware components. -Â Only perform repairs, or have them performed by an authorized third party, after having received our express written approval.
499 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.5 | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 4.43.2.6.5.1 | 562154 | Measure | User Manual - Impermissible or faulty manipulation from network | CAUTIONCause: Unauthorized or incorrect alterations/changes of the software or connection of the system to a network. Possible consequences: Risk of unauthorized access. Remedy: Make sure that all necessary precautions are taken with regard to the applicable security level if a functionality is added or the configuration of the condition on delivery is changed.
500 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.5 | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 4.43.2.6.5.2 | 562155 | Measure | Reference to measures against unauthorized installation of non-approved software | <XREF>563266</XREF><XREF>563267</XREF>
501 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.5 | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 4.43.2.6.5.3 | 569133 | Measure | Imaging system access restrictions | The system user does not have access to the system level. This includes the access to the underlaying operating system and the belonging functions like file menu, system properties or start menu.
502 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.5 | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 4.43.2.6.5.4 | 569135 | Measure | Disable autorun for removable media | The autorun feature for removable media must be disabled on the FLC imaging system.
503 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.5 | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 4.43.2.6.5.5 | 569136 | Measure | Imaging system access restrictions for internet and BIOS | Define a testcase to verify that the access to the internet is not possible for the clinical user. The service user shall be able to access internet via a special service environment. The access to BIOS is password-protected. .
504 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.6 | 562156 | Cause | Incorrect treatment | 4.43.2.6.6.1 | 562157 | Measure | Reference - User Manual - Trained Users | <XREF>561986</XREF>
505 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.7 | 562158 | Cause | Harddisk or file system defect | 4.43.2.6.7.1 | 562159 | Measure | User manual - No Archive Medium | CAUTIONCause: Loss of image dataPossible Consequences: Risk of unnecessary X-ray exposureRemedy:- The hard disk of the imaging system is not suitable for long-term archiving of patient and image data. There are statutory requirements governing the archiving period, data availability, and data security (data integrity), as well as recommendations concerning fire protection or water damage for the archiving of image data. The operator of the archive is responsible for observing these regulations. - Because of advances in technical development and statutory requirements, the storage of image data and access to it cannot be realized with a single storage and media technology alone. Therefore, the migration of data may be necessary under the responsibility of the operator of the digital archive.
506 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.8 | 562160 | Cause | Electric power failure | 4.43.2.6.8.1 | 562161 | Measure | Reference - The imaging system performs a controlled shutdown | <XREF>563249</XREF>
507 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.9 | 562162 | Cause | Insufficient charge (power) of UPS | 4.43.2.6.9.1 | 562163 | Measure | No measure required | Utilization of UPS has changed: intended use is not keeping system operational but ensuring a controlled shut down. Hazard as well as hazardous situation/Cause is not applicable.
508 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.2 | 562100 | Function | Connect Trolley and Main Unit | 4.43.2.6 | 562144 | Hazard | Delayed or no interventional decision | 4.43.2.6.10 | 562164 | Cause | Charging capacity of the Uninterruptible Power Supply worn out | 4.43.2.6.10.1 | 562165 | Measure | No measure required | Utilization of UPS has changed: intended use is not keeping system operational but ensuring a controlled shut down. Hazard as well as hazardous situation/Cause is not applicable.
509 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.3 | 562166 | Function | Performing system relevant check | 4.43.3.1 | 562167 | Hazard | No hazard on system level | 4.43.3.1.1 | 562168 | Cause | None | 4.43.3.1.1.1 | 562169 | Measure | None
510 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.4 | 562170 | Function | Performing quality control | 4.43.4.1 | 562171 | Hazard | No hazard on system level | 4.43.4.1.1 | 562172 | Cause | None | 4.43.4.1.1.1 | 562173 | Measure | None
511 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.5 | 562174 | Function | Calibration | 4.43.5.1 | 562175 | Hazard | Unnecessary radiation | 4.43.5.1.1 | 562176 | Cause | Incorrect parameters | 4.43.5.1.1.1 | 562177 | Measure | Automatic procedure to perform FD calibration correctly | Automatic FD calibration procedure has to be initiated manually.
512 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.5 | 562174 | Function | Calibration | 4.43.5.1 | 562175 | Hazard | Unnecessary radiation | 4.43.5.1.2 | 562178 | Cause | Operating error | 4.43.5.1.2.1 | 562179 | Measure | Admin Guide - Calibration Procedure | CAUTIONCause: Operator error. Possible Consequences: Risk of unnecessary radiation exposure!Â Remedy: Follow the calibration procedure instructions.
513 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.5 | 562174 | Function | Calibration | 4.43.5.2 | 562180 | Hazard | Image artifacts - Wrong interventional decision possible | 4.43.5.2.1 | 562181 | Cause | Bad calibration | 4.43.5.2.1.1 | 562182 | Measure | Reference - Automatic procedure to perform calibration correctly | <XREF>562177</XREF>
514 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.5 | 562174 | Function | Calibration | 4.43.5.2 | 562180 | Hazard | Image artifacts - Wrong interventional decision possible | 4.43.5.2.1 | 562181 | Cause | Bad calibration | 4.43.5.2.1.2 | 562183 | Measure | Reference - User Manual - Calibration Procedure | <XREF>562179</XREF>
515 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.5 | 562174 | Function | Calibration | 4.43.5.3 | 562184 | Hazard | Unwanted radiation | 4.43.5.3.1 | 562185 | Cause | Person in room during calibration | 4.43.5.3.1.1 | 562186 | Measure | Service training | Product specific service training (for service calibration)
516 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.1 | 562188 | Hazard | Danger for patient or user | 4.43.6.1.1 | 562189 | Cause | Wrong Intended Use | 4.43.6.1.1.1 | 562190 | Measure | User Manual - Intended use only | CAUTIONCause: Improper use of the system beyond its intended use. Possible consequences:Â Risk to the patient and user!Remedy: The system may only be used as defined by its intended use for clinical applications.
517 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.1 | 562188 | Hazard | Danger for patient or user | 4.43.6.1.2 | 562191 | Cause | Crash of subsystems | 4.43.6.1.2.1 | 562192 | Measure | Architectural Design of Subsystems and Components | see following sections in System Architecture (<XREF>561860</XREF>): * <xref>286376</xref> , see spacing * <xref>286377</xref> * <xref>353392</xref> * <xref>353414</xref>
518 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.1 | 562188 | Hazard | Danger for patient or user | 4.43.6.1.2 | 562191 | Cause | Crash of subsystems | 4.43.6.1.2.2 | 562193 | Measure | User Manual - Reset procedure for a freezed imaging SW | CAUTIONCause: Imaging system is frozen or restart is not possible. Possible consequences: Risk of delay in the clinical procedureRemedy: Restart the system as presented in the following reset procedure.
519 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.1 | 562188 | Hazard | Danger for patient or user | 4.43.6.1.2 | 562191 | Cause | Crash of subsystems | 4.43.6.1.2.3 | 562194 | Measure | Block keyboard / mouse inputs during radiation | Functionality that would cause dialogs or other UI elements to cover the live image or that would cause performance bottlenecks for live image processing and display are disabled during radiation.
520 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.2 | 562195 | Hazard | Loss of functionality | 4.43.6.2.1 | 562196 | Cause | Complexity of the system or line voltage below the specified threshold | 4.43.6.2.1.1 | 562197 | Measure | Reference to User Manual - Establish emergency procedures | <XREF>562205</XREF>.
521 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.2 | 562195 | Hazard | Loss of functionality | 4.43.6.2.2 | 562198 | Cause | Foreseeable misuse | 4.43.6.2.2.1 | 562199 | Measure | Reference - User Manual - Trained Users | <XREF>561986</XREF>
522 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.2 | 562195 | Hazard | Loss of functionality | 4.43.6.2.3 | 562200 | Cause | Some parts in system could lead to a hazard by wear | 4.43.6.2.3.1 | 562201 | Measure | User Manual - Wear and tear | CAUTIONCause: Wear and tear. Possible consequences:Â Risk of injury to the patient, operating personnel, and other persons. Permanent damage to the system. Remedy:Â Follow the maintenance guidelines to maintain the safety and functionality of the system.
523 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.2 | 562195 | Hazard | Loss of functionality | 4.43.6.2.4 | 562202 | Cause | Operational hazard due to missing knowledge of the system | 4.43.6.2.4.1 | 562203 | Measure | Provide training for technician | Provide training for technician. Service Documentation:WARNING:Cause: Missing training or not following service instructions can result in failure. Risk: Incorrect service can lead to damage for patient, personnel and system. Remedy: Service activities are only permitted by trained personnel and according to the appropriate service instructions.
524 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.3 | 562206 | Hazard | Danger to electric high voltage or line voltage | 4.43.6.3.1 | 562207 | Cause | Problems with protective ground | 4.43.6.3.1.1 | 562208 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
525 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.3 | 562206 | Hazard | Danger to electric high voltage or line voltage | 4.43.6.3.1 | 562207 | Cause | Problems with protective ground | 4.43.6.3.1.2 | 562209 | Measure | Start up instruction | Start up instructionProtective earth measurements are described inÂ start upÂ instructions and done during system installation.
526 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.3 | 562206 | Hazard | Danger to electric high voltage or line voltage | 4.43.6.3.2 | 562210 | Cause | Contact with Parts Carrying Voltage | 4.43.6.3.2.1 | 562211 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
527 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.3 | 562206 | Hazard | Danger to electric high voltage or line voltage | 4.43.6.3.2 | 562210 | Cause | Contact with Parts Carrying Voltage | 4.43.6.3.2.2 | 562212 | Measure | Warning Label - Voltage is present | Attach label at the Main UnitÂ with safety sign: Label 3 of IEC 60601-1 Table D. 2,Â dangerous voltage presentÂ Additionally it shall contian theÂ followingÂ inscription:WARNING: Voltage is present after shutdown.
528 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.3 | 562206 | Hazard | Danger to electric high voltage or line voltage | 4.43.6.3.3 | 562213 | Cause | Isolation defect | 4.43.6.3.3.1 | 562214 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
529 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.4 | 562215 | Hazard | Hazards from fire or smoke | 4.43.6.4.1 | 562216 | Cause | Fire inside the system | 4.43.6.4.1.1 | 562217 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
530 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.4 | 562215 | Hazard | Hazards from fire or smoke | 4.43.6.4.1 | 562216 | Cause | Fire inside the system | 4.43.6.4.1.2 | 562218 | Measure | User Manual - Switch off the system in the event of fire | WARNING:Cause: Fire in or near the system. Possible consequences: Injury to patient and personnel, and damage to device. Risk of gas poisoning due to burning plastic. Remedy:- In case of fire switch off the system. - Be aware of and inform patients of escape routes. - Be aware of where fire extinguishers are located and know how to use them. - In case of fire switch off the system.
531 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.1 | 562221 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
532 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.2 | 562222 | Measure | Surface temperature applied parts | For all surfaces which can be touched by the patient (applied parts according to IEC 60601-1 cl. 4. 6, <XREF>561933</XREF>), the surface temperaturesÂ under normal and single fault conditions according to IEC 60601-1, table 24 (see <XREF>434357</XREF>) Â for the following contact duration shall be applied:Affected components and areas:- Cover of the detector incl. FD handle and panel:Â 10 min <= t- CoverÂ of the monoblock, excluded monoblock backside: 1 min <= t <10 min- Wired and wireless hand control: 10 min <= t- Remote control panel on the rear: t < 1 min (relatively heavy, therefore it is not expected to be placed on the patient directly, besides shortly putting it down, when mounted at the table, the patient cannot come into contact with the rear side)- Remote control panel on the top and the sides: 1 min <= t < 10 min- C-shaped mechanics of the system: 1 min <= t < 10 min
533 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.3 | 562225 | Measure | Temperature sensor self test | During operating of the system there will be a self test of the temperature sensing circuit. In the case of a defect x-ray radiation will be prevented.
534 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.4 | 562226 | Measure | System Warning - Temperature of the Detector | The temperature inside the detector is checked quasi-continously (i.e. short interval) by FLC. According to the detector type an appropiate warning for the user is displayed, when first warning temperature is reached. When the second warning temperature is reached, an additional warning message is displayed to inform the user. Temperature for first warning message:- 52Â°C for FDmediumHTemperature for 2nd warning message:- 54Â°C for FDmediumHThe detector will not be switched off at any time, during an intervention.
535 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.5 | 562228 | Measure | User Manual - Hot tube | CAUTIONCause: Under extreme operating conditions (continuous operation close to the maximum tube voltage), the X-ray tube can become very hot. This can shorten the service life of the system components. Â Possible consequences: Risk of burns and prolongation of the operation!Â Remedy:- Avoid touching the X-ray tube housing. - Change your parameter settings, e.g. by reducing the dose level. - We also recommend keeping a second system ready for continuous radiation lasting longer than 30 minutes. - Please make sure that the system has cooled down before performing critical interventional applications.
536 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.1 | 562220 | Cause | Parts too warm | 4.43.6.5.1.6 | 562229 | Measure | ServiceDoc - Hot tube housing | Service DocumentsCAUTIONCause: In extreme operating mode (continuous operation), Â the temperature of the X-ray tube assembly can heat up. Possible consequences: Danger of burns. Remedy: Avoid contact with the tube housing
537 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.2 | 562230 | Cause | Leakage of hot oil from tube assembly | 4.43.6.5.2.1 | 562231 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
538 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.5 | 562219 | Hazard | Thermal energy and high temperature | 4.43.6.5.2 | 562230 | Cause | Leakage of hot oil from tube assembly | 4.43.6.5.2.2 | 562232 | Measure | Pressure switch in tube | Pressure switch in tube is monitored during exposure. Â If excess pressure is detected, the switch is opened and the radiation is switched off.
539 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.7 | 562237 | Hazard | Patient is infected by a virus due to a bad disinfection, system is damaged | 4.43.6.7.1 | 562238 | Cause | Incorrect cleaning of system after each patient, liquids or blood enter the system | 4.43.6.7.1.1 | 562239 | Measure | Reference - Spillage of Liquids on System | <XREF>562238</XREF>
540 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.7 | 562237 | Hazard | Patient is infected by a virus due to a bad disinfection, system is damaged | 4.43.6.7.1 | 562238 | Cause | Incorrect cleaning of system after each patient, liquids or blood enter the system | 4.43.6.7.1.2 | 562240 | Measure | User Manual - Cleaning and draping | WARNINGCause: Fluid or blood seeps into the system, use ofÂ inappropriate cleaning agents. Possible consequences: Risk of damage to the system, risk of infection or electrical hazard. Remedy: Use sterile or sterilizable covers at least in C-arm area (C-arm, cover of the flat detector and X-ray monobloc). Do not let cleaning liquids seep into the openings of the system (e.g. air openings, gaps between covers). Observe the cleaning and disinfection instructions. Only use the recommended substances for cleaning and disinfection.
541 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.8 | 562241 | Hazard | Environmental hazard | 4.43.6.8.1 | 562242 | Cause | Explosion | 4.43.6.8.1.1 | 562243 | Measure | Reference - User Manual - Danger of explosion | <XREF>562006</XREF>
542 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.9 | 562244 | Hazard | Chemical hazards | 4.43.6.9.1 | 562245 | Cause | Toxic materials and fluids | 4.43.6.9.1.1 | 562246 | Measure | Technical instruction for the disposal procedure | A technical instruction for the disposal procedureÂ shall exist, which describes which parts must be removed prior to scrapping and how these parts are disposed in an environment friendly manner.
543 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.10 | 562247 | Hazard | Electro magnetic disturbances | 4.43.6.10.1 | 562248 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 4.43.6.10.1.1 | 562249 | Measure | Reference - EMC compliance - emissions, immunity and pass / fail criteria | <XREF>562015</XREF>
544 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.11 | 562250 | Hazard | Danger by releasing of stored mechanical energy | 4.43.6.11.1 | 562251 | Cause | Rotation of c-arm when power fails | 4.43.6.11.1.1 | 562252 | Measure | User Manual - Rotation of c-arm | CAUTIONCause: C-arm rotation. Possible consequences: Risk of crushingRemedy: Pay particular attention to crushing risks between moving system parts and the corresponding guide openings.
545 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.11 | 562250 | Hazard | Danger by releasing of stored mechanical energy | 4.43.6.11.2 | 562253 | Cause | Cable c-arm falling down | 4.43.6.11.2.1 | 562254 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
546 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.11 | 562250 | Hazard | Danger by releasing of stored mechanical energy | 4.43.6.11.3 | 562255 | Cause | Collimator comes loose | 4.43.6.11.3.1 | 562256 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
547 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.1 | 562260 | Cause | Unintentional activation of control elements | 4.43.6.12.1.1 | 578758 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
548 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.2 | 562262 | Cause | Manual or motorized movements of any parts | 4.43.6.12.2.1 | 562263 | Measure | User Manual - Risk of Crushing | CAUTIONCause: Manual or motorized movementsPossible consequences: Risk of crushing and collision for persons and objects in the vicinity of moving parts. Remedy:- During manual and motorized movements of the device, make sure you take into account any persons or objects in the way of movement, and control the movement to avoid collisions. - Make sure that everyone is outside the hazardous zone.
549 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.2 | 562262 | Cause | Manual or motorized movements of any parts | 4.43.6.12.2.2 | 562264 | Measure | User Manual - Danger zones | CAUTIONCause:Â The points marked on the figure indicate hazardous locations around the system. Possible consequences: Risk of injury to the patient and personnel due to crushing or collision. Remedy: Be careful around the hazardous locations indicated.
550 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.2 | 562262 | Cause | Manual or motorized movements of any parts | 4.43.6.12.2.3 | 562265 | Measure | Limit motorized orbital operating radius | After saving the software end positions for motorized orbital movement, it shall be checked that the orbital movement stops so that the remaining gap to the end stop is >= 25 mm.
551 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.2 | 562262 | Cause | Manual or motorized movements of any parts | 4.43.6.12.2.4 | 562266 | Measure | Warning Label - Risk of crushing hand | WarningÂ labels "risk of crushing - hand" shall be placed on the angulation cover (upper and lower side), for orbital movement.
552 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.3 | 562267 | Cause | Motorized movements of lifting column | 4.43.6.12.3.1 | 562268 | Measure | Reference - User Manual - Risk of crushing | <XREF>562263</XREF>
553 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.3 | 562267 | Cause | Motorized movements of lifting column | 4.43.6.12.3.2 | 562269 | Measure | Reference - User Manual - Danger Zones | <XREF>562264</XREF>
554 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.3 | 562267 | Cause | Motorized movements of lifting column | 4.43.6.12.3.3 | 562270 | Measure | Reference - User Manual - Trained Users | <XREF>561986</XREF>
555 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.3 | 562267 | Cause | Motorized movements of lifting column | 4.43.6.12.3.4 | 562271 | Measure | Warning Label - Risk of crushing foot | WarningÂ labels "risk of crushing - foot" shall be placed on the cover of the transport skid on the front side, as well as onÂ and the left and right side.
556 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.12 | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 4.43.6.12.3 | 562267 | Cause | Motorized movements of lifting column | 4.43.6.12.3.5 | 562272 | Measure | Warning Label - Risk of crushing hand | A warningÂ label ""risk of crushing - hand"" shall be placed on the upper main unit cover to indicate crushing between theÂ vertical movement unit.
557 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.13 | 562273 | Hazard | Injuries because of falling parts | 4.43.6.13.1 | 562274 | Cause | Wrong calculated max load, unforeseeable load. | 4.43.6.13.1.1 | 562275 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
558 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.14 | 562276 | Hazard | Injury or crushing due to emergency stop out of order | 4.43.6.14.1 | 562277 | Cause | Defect in button or cable harness | 4.43.6.14.1.1 | 562279 | Measure | Check emergency stop | Check emergency stop during periodic maintenance.
559 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.15 | 562280 | Hazard | Operator fails to actuate emergency stop | 4.43.6.15.1 | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 4.43.6.15.1.1 | 578759 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
560 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.15 | 562280 | Hazard | Operator fails to actuate emergency stop | 4.43.6.15.1 | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 4.43.6.15.1.2 | 578760 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
561 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.15 | 562280 | Hazard | Operator fails to actuate emergency stop | 4.43.6.15.1 | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 4.43.6.15.1.3 | 562284 | Measure | User Manual - Emergency stop location | CAUTIONCaution: Position of EMERGENCY STOP button not known. Possible consequences: In an emergency, the EMERGENCY STOP button will not be actuated in a timely manner. Remedy: Make sure you know where the EMERGENCY STOP buttons are located. EMERGENCY STOP buttons are in the following locations (see the figure).
562 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.16 | 562285 | Hazard | Malfunction of the system | 4.43.6.16.1 | 562286 | Cause | Inappropriate accessory | 4.43.6.16.1.1 | 562287 | Measure | User Manual - Inappropriate accessory | CAUTIONCause: Inappropriate accessories. Possible consequences: The use of accessories that do not comply with the safety requirements of this equipment can reduce the safety of the entire system. Remedy: Use only original Siemens Healthineers accessories or accessories approved by Siemens Healthineers.
563 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.6 | 562187 | Function | System Operation | 4.43.6.16 | 562285 | Hazard | Malfunction of the system | 4.43.6.16.2 | 562288 | Cause | Any defect parts in the system | 4.43.6.16.2.1 | 562289 | Measure | User Manual - Defect of system components | CAUTIONCause: System component error. Possible consequences: An error on one system component can reduce the safety of the entire system. Remedy:- Terminate system operation. - Contact Siemens Healthineers Customer Service.
564 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.7 | 562295 | Function | Stop button | 4.43.7.1 | 562296 | Hazard | Emergency stop out of order | 4.43.7.1.1 | 562297 | Cause | Defect in button or cable harness | 4.43.7.1.1.1 | 562298 | Measure | Reference - Check emergency stop - maintenance | <XREF>562279</XREF>
565 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.8 | 562299 | Function | Clean system | 4.43.8.1 | 562300 | Hazard | Patient collision | 4.43.8.1.1 | 562301 | Cause | Swapped wireless footswitch | 4.43.8.1.1.1 | 562302 | Measure | Reference to measures regarding footswitch | See <XREF>563088</XREF>, <XREF>563089</XREF>, <XREF>563090</XREF>, <XREF>563091</XREF>.
566 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.8 | 562299 | Function | Clean system | 4.43.8.2 | 562303 | Hazard | Unwanted radiation | 4.43.8.2.1 | 562304 | Cause | Swapped wireless footswitch | 4.43.8.2.1.1 | 562305 | Measure | Reference to measures regarding footswitch | See <XREF>563088</XREF>, <XREF>563089</XREF>, <XREF>563090</XREF>, <XREF>563091</XREF>.
567 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.9 | 562306 | Function | Store patient data | 4.43.9.1 | 562307 | Hazard | Unnecessary minor exposition to radiation | 4.43.9.1.1 | 562308 | Cause | Image store does not keep pace with acquisition under error conditions | 4.43.9.1.1.1 | 562309 | Measure | Warning Message | When the image storage does not keep pace with the acquisition (e.g. , due to buffer issues), radiography imaging shall be stopped and a message shall be displayed to the operator. Fluoroscopy imaging shall not be stopped or inhibited.
568 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.9 | 562306 | Function | Store patient data | 4.43.9.1 | 562307 | Hazard | Unnecessary minor exposition to radiation | 4.43.9.1.2 | 562310 | Cause | Storage Limit Reached (License) | 4.43.9.1.2.1 | 591790 | Measure | Reference - Storage capacity exhausted | <XREF>562612</XREF>
569 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.9 | 562306 | Function | Store patient data | 4.43.9.1 | 562307 | Hazard | Unnecessary minor exposition to radiation | 4.43.9.1.2 | 562310 | Cause | Storage Limit Reached (License) | 4.43.9.1.2.2 | 591791 | Measure | Reference - Display storage space | <XREF>562613</XREF>
570 | 4.43 | 562063 | Chapter | Setup Main Unit | 4.43.9 | 562306 | Function | Store patient data | 4.43.9.2 | 562312 | Hazard | Wrong interventional decision | 4.43.9.2.1 | 562313 | Cause | Image quality cannot be improved after acquisition | 4.43.9.2.1.1 | 562314 | Measure | Store raw data | The imaging system shall store acquired 2D images immediately to the hard disk, before any user manipulation is done. The operator shall always be able to restore the original image visualization.
571 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.1 | 562316 | Function | HIS/RIS | 4.44.1.1 | 562317 | Hazard | Non-harmonized configuration | 4.44.1.1.1 | 562318 | Cause | Incorrect Configuration of HIS/RIS | 4.44.1.1.1.1 | 562319 | Measure | User Manual - HIS/RIS Configuration | CAUTIONCause: Non-harmonized configuration of HIS/RISPossible consequences: Impossibility of CommunicationRemedy: Do not use configurationÂ "patient ID only" within HIS/RIS as Imaging System (FLC) identifies a patient by quadruple (patient name, ID, sex, date-of-birth).
572 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.1 | 562323 | Measure | Use of reliable data transfer standards | For transfer of dicom images (e.g. to a pacs or a share)Â reliable data transfer standards and protocols shall be used. Error messages in case of network problems in particular by transferring DICOM images to specific destinations(e.g. Pacs, Share,. . ).
573 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.2 | 562324 | Measure | Prompt by local deleting of data | Confirmation by user when deleting studies where nothing has been archived or printed. Message: ""Do you want unarchived studies to be deleted?""The user shall have the possibility to protect studies against deletion. Â Note:Information on archiving status will be set image by image (during transfer). As soon as one image out ofÂ study was sent successfullyÂ the study as a whole will be considered as archived.
574 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.3 | 562325 | Measure | Data integrity of xray image transfer via NXS | For transfer of xray imagesÂ via NXS the data integrity of these images shall beÂ considered by using a check sum algorithm and a ID/counter of transfered images.
575 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.4 | 562326 | Measure | NXS protocol with heartbeat | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection.
576 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.5 | 562327 | Measure | NXS protocol with heartbeat for extapp | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection.
577 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.6 | 562328 | Measure | Additional information of xray image transfer via NXS | For transfer of xray imagesÂ via NXSÂ some additional information (like patient quadrupel)Â shall beÂ incorporatedÂ to identify clearly which xray image belong to which patient and the properties of the xray image.
578 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.7 | 562329 | Measure | Additional information of xray image transfer via NXS for extapp | For transfer of xray imagesÂ via NXSÂ some additioinal information (like patient quadrupel)Â shall beÂ incorporatedÂ to identify clearly which xray image belong to which patient and the properties of the xray image.
579 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.8 | 568849 | Measure | Image compression threshold | For static images, the grey scale value of a pixel may vary by +/- 2% of the original grey scale value. For live images, theÂ image quality may be further reduced, keeping the frame rate. TheÂ static app UI contentÂ shall achieve the above mentioned threshold.
580 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.2 | 562320 | Function | Transfer exam | 4.44.2.1 | 562321 | Hazard | Delayed or no interventional decision | 4.44.2.1.1 | 562322 | Cause | Network disturbance | 4.44.2.1.1.9 | 569294 | Measure | Reference - Image display delay time - external | <XREF>568835</XREF>
581 | 4.44 | 562315 | Chapter | Connect to Network | 4.44.3 | 562330 | Function | Connectivity | 4.44.3.1 | 562331 | Hazard | No hazard on system level | 4.44.3.1.1 | 562332 | Cause | None | 4.44.3.1.1.1 | 562333 | Measure | None
582 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.1 | 562336 | Hazard | Impaired vision through laser-beam | 4.45.1.1.1 | 562337 | Cause | People looking into laser-beam | 4.45.1.1.1.1 | 562338 | Measure | Usage of laser class 1 | Usage of laser class 1 (according to IEC 60825-1:2014) in monoblock, detector and horizontal laser, and LIDAR scanner.
583 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.1 | 562347 | Cause | Unintended movement activation when wired control element is connected to the system | 4.45.1.2.1.1 | 578761 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
584 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.1 | 562347 | Cause | Unintended movement activation when wired control element is connected to the system | 4.45.1.2.1.2 | 578766 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
585 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.2 | 562351 | Cause | Movement without intentional activation on remote console | 4.45.1.2.2.1 | 578762 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
586 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.2 | 562351 | Cause | Movement without intentional activation on remote console | 4.45.1.2.2.2 | 578767 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
587 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.3 | 562354 | Cause | Movement does not stop | 4.45.1.2.3.1 | 578768 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
588 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.3 | 562354 | Cause | Movement does not stop | 4.45.1.2.3.2 | 578320 | Measure | Reference - Continuous activation of movements with redundant paths | <XREF>562377</XREF>
589 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.2 | 562342 | Hazard | Collision with patient, operator or equipment | 4.45.1.2.3 | 562354 | Cause | Movement does not stop | 4.45.1.2.3.3 | 578763 | Measure | Reference - Warning - General risk of collisions | <XREF>563493</XREF>
590 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.3 | 562358 | Hazard | Electrical shock | 4.45.1.3.1 | 562359 | Cause | Protective earth issues | 4.45.1.3.1.1 | 562360 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
591 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.3 | 562358 | Hazard | Electrical shock | 4.45.1.3.1 | 562359 | Cause | Protective earth issues | 4.45.1.3.1.2 | 562361 | Measure | Reference - Start up instruction | <XREF>562209</XREF>
592 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.3 | 562358 | Hazard | Electrical shock | 4.45.1.3.2 | 562362 | Cause | Contact with parts carrying voltage. | 4.45.1.3.2.1 | 562363 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
593 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.3 | 562358 | Hazard | Electrical shock | 4.45.1.3.3 | 562364 | Cause | Defective insulation | 4.45.1.3.3.1 | 562365 | Measure | IEC 60601-1 Clause 8.7 and 8.8 | The system shall comply withÂ IEC 60601-1 cl. Â 8. 7 and 8. 8. Voltages which are accessible without tool shall not exceed 42,4 V peak a. c. and 60 V d. c. Touch current and patient leakage currents shall comply with cl. 8. 7The operator shall be protected by 2 MOOP. The patient shall be protected by 2 MOPP. Applied parts of type B according to cl. 4. 6 are:- Cover of the detector incl. FD handle and panel- CoverÂ of the monoblock, excluded monoblock backside- Wired and wireless hand control- Remote control panel- C-shaped mechanics of the systemNo parts are identified which are classified as type BF or CF.
594 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.3 | 562358 | Hazard | Electrical shock | 4.45.1.3.3 | 562364 | Cause | Defective insulation | 4.45.1.3.3.2 | 562366 | Measure | IEC 60601-1 Clause 9.8 | Parts of the system that are suspended or supported shall be developed using the tensile safety factors listed in IEC60601-1:2005, Table 21. These parts shall fulfill the requirements of section 9. 8 of the standard.
595 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.4 | 562367 | Hazard | Burns | 4.45.1.4.1 | 562368 | Cause | Hot surfaces | 4.45.1.4.1.1 | 562369 | Measure | Reference - Compliance with IEC 60601-1 | ompliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
596 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.5 | 562370 | Hazard | Injuries because of falling parts | 4.45.1.5.1 | 562371 | Cause | Inadequate layout or construction of mechanical system parts. | 4.45.1.5.1.1 | 562372 | Measure | Reference - Compliance with IEC 60601-1 Standards | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
597 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.1 | 562335 | Function | Position Patient | 4.45.1.5 | 562370 | Hazard | Injuries because of falling parts | 4.45.1.5.1 | 562371 | Cause | Inadequate layout or construction of mechanical system parts. | 4.45.1.5.1.2 | 562373 | Measure | Reference to mounting of the monoblock | See HM_S1_mec_F4G1U1M1 in <XREF>578770</XREF>.
598 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.1 | 562376 | Cause | Unintentional activation of movement | 4.45.2.1.1.1 | 562380 | Measure | No additional measure required - Reference to set values motorized speed | <XREF>563486</XREF>
599 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.2 | 562381 | Cause | Malfunction of the control element | 4.45.2.1.2.1 | 578094 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
600 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.1 | 562384 | Measure | Stop vertical lift | WhenÂ the system is moved downwards, the movement shall be stopped at aÂ safety distanceÂ of min. 120 mm between the floor and the C-arm parts (either detector or monoblock rotated downwards). Â Note: The crushing distance for feet between C-arm and front side of the wheel cover of the chassis is allowed to be belowÂ 120Â mm due to the crushing label on the cover. The crushing distance for toes of 50 mm is met.
601 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.2 | 562385 | Measure | Acoustic signal | After reaching the safety distance inÂ <XREF>562384</XREF>, an acoustic alarm signal is triggered, when the C-arm is moved further downwards in direction of crushing hazard. When a distance of min. Â 25 mm between horizontalÂ unit and the top of the chassisis reached, downwards movement shall be stopped and prevented. Â Note: Crushing distance for toes between floor and C-arm or C-arm and front side of the wheel cover of the chassis is allowed to be below 50 mm due to acoustic signal andÂ crushing label.
602 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.3 | 562386 | Measure | Reference - User Manual - Guide Openings | <XREF>562444</XREF>
603 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.4 | 562387 | Measure | Reference - User Manual - Risk of Crushing | <XREF>562263</XREF>
604 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.5 | 562388 | Measure | User Manual - Check Audio Signal | CAUTIONCause: Audible warning for safety distance not functionalPossible consequences: Risk of crushingRemedy: Check for proper loudspeaker function before beginning the examination. Note when booting the system whether an audible signal is heard.
605 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.6 | 562389 | Measure | Supply Chain Documentation - Assembly Crushing Zone Switch | Ensure location of crushing zone switch during final inspection. Measure the safety distances defined after theÂ intermediate switch off, see <XREF>562384</XREF> and after the final switch off, seeÂ <XREF>562385</XREF>.
606 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.3 | 562383 | Cause | Falling below the safety distance during vertical movement | 4.45.2.1.3.7 | 562390 | Measure | User Manual - Meaning of Audio Signals | CAUTIONCause: Occurrence of audible signals. Possible consequences: Hazards possible, for example due to crushing risks, extended radiation release, and high dose rate. Remedy:- Please note that continued system operation may lead to one of these hazards. - Use the system with extreme caution to avoid these hazards.
607 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.4 | 562391 | Cause | Movement speed too high | 4.45.2.1.4.1 | 578771 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
608 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.5 | 562394 | Cause | Overtravel (stopping distance) too long | 4.45.2.1.5.1 | 578326 | Measure | Reference - Maximum overtravel values | <XREF>562465</XREF>
609 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.1 | 562375 | Hazard | Crushing through motor-driven movement | 4.45.2.1.5 | 562394 | Cause | Overtravel (stopping distance) too long | 4.45.2.1.5.2 | 562396 | Measure | Force limit orbital / angular motorized movement | The system shall stop motorized movement if the holding force exeeds 200 N for angular and orbital movements.
610 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.1 | 562399 | Cause | Weight causes drive to lower autonomously | 4.45.2.2.1.1 | 562400 | Measure | Provide brake or self-locking of the drive | The drive of the vertical movement must be designed in a way that prevents an autonoums movement, e.g. to provide a brake or self-locking of the drive. No autonomous movement is given, as long as a lowering within one hour does not exceed 2 mm.
611 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.2 | 562401 | Cause | C-arm not weight balanced | 4.45.2.2.2.1 | 562402 | Measure | Design C-arm weight balanced | Design of C-arm shall be weight balanced to prevent autonomous movement withÂ brakes open. Â Note:AÂ remaining movement may be observable. Range of remaining movement depends on angular/orbital position and shall not exceed 20 degree. Nonetheless,Â remaining movementÂ shall be uncritical for the user.
612 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.3 | 562403 | Cause | Additional weight changes weight balance | 4.45.2.2.3.1 | 562404 | Measure | User Manual - Unbalanced C-arm | CAUTIONCause: Changing the weight distributionPossible consequences: Risk of crushingRemedy: Adding weight to the detector or single tank side means a loss in counterweight and can result in unintended movement of the C-arm.
613 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.4 | 562405 | Cause | Weakening of magnetic force | 4.45.2.2.4.1 | 562406 | Measure | Service Instruction - periodic maintenance | Demagnetization may appear, if one of following actions will affect on OR table holder1) external magnetic field (magnetic hysteresis, coercivity field Hc),2) strong heating (threshold: Curie temperature), and/or3) strong hits/shocksÂ There is no external magnetic field available. Temperature thresholds to demagnetize are 310-340Â° Celsius (Neodym orÂ Neodym-Iron-Boron combinations, NdFeB) 354 degree Celsius (Nickel, Ni), and 450Â° Celsius (Ferrite, i.e. Fe2O3 or Fe3O4 combinations), and 770Â° Celsius (Iron, Fe). These environmental conditions will not appear in lifetime. Remaining issue are (3) strong hits/shocks. These seems to be very unlikely that this kind of force will be applied to the OR table holder, but can be assumed as improbable, but not impossible. Â Therefore, magnetic force of OR table holder shall be checked during periodic maintenance: The magnets must fix the RCP so that pulling them with up to 150N will not loosen the RCP from the OR table holder. "or "The magnets shall fix the RCP so that, when pulling on it with at least 150N, the RCP will not loosen from the OR holder.
614 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.5 | 562407 | Cause | Angular / orbital brake not functional | 4.45.2.2.5.1 | 562408 | Measure | Reference to Caution Operator Manual - Check brakes | <XREF>561962</XREF>
615 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.2 | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 4.45.2.2.5 | 562407 | Cause | Angular / orbital brake not functional | 4.45.2.2.5.2 | 562409 | Measure | Reference to weight balanced design | <XREF>562402</XREF>.
616 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.1 | 562414 | Cause | Carrying capacity of the unit too small | 4.45.2.3.1.1 | 562415 | Measure | Design of the movement unit | The moving positioning units must be designed so that a overload can be carried in compliance with DIN EN 60601-1 (2006), Kap. 9. 8 , Table 21.
617 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.2 | 562416 | Cause | Motor-driven movement and collision with fixed parts | 4.45.2.3.2.1 | 562417 | Measure | Design of the movement unit collision | The moving positioning units must be designed so that a overload due to a collision can be handled. Â Remark:The component parts shall be carrying a overload in compliance with DIN EN 60601-1 (2006). Kap. 9. 8, Table 21.
618 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.3 | 562418 | Cause | Breakage or wear of the spindle nut | 4.45.2.3.3.1 | 562419 | Measure | Additional safety mechanism - lift column | There shall be aÂ mechanical safety mechanism for the lift column which prevents the C-arm from crashing down in case of a spindle failure. The mechanical safety mechanism shall be designed toÂ fulfill a tensileÂ safety factor >= 4. The overtravel (stopping distance from the time of engagement to the time of no further movement) shall be max. Â 5 mm. Note: IEC 60601-1, cl. 9. 8. 4 Systems with MECHANICAL PROTECTIVE DEVICES
619 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.3 | 562418 | Cause | Breakage or wear of the spindle nut | 4.45.2.3.3.2 | 562420 | Measure | Activation of safety mechanism blocks lift drive | After the activation of the mechanical safety mechanismÂ <XREF>562419</XREF> of the vertical column, the operation of the lift driveÂ shall not be possible anymore. Note: The drive is not capable of moving again due to the activated mechanism, and an motor overcurrent message is displayed. The column can be manually screwed down with a tool (ratch) only which can be inserted in a whole at the bottom of the chassis (moving it down is necessary for transporting it back to the factory). Repair is only done in factory, not in the field.
620 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.3 | 562418 | Cause | Breakage or wear of the spindle nut | 4.45.2.3.3.3 | 562421 | Measure | Labeling of the safety equipment | A labeling, safety sign ISO 7010-W001, has to be placed next to the mechanical safety equipment or in a way so the connection to it is obvious.
621 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.3 | 562413 | Hazard | Crushing through breakage | 4.45.2.3.3 | 562418 | Cause | Breakage or wear of the spindle nut | 4.45.2.3.3.4 | 562422 | Measure | User Manual - Spindle Nut | WARNING:Cause: Spindle nut broken or worn. Possible consequences: Risk of crushingRemedy: Vertical lift is blocked if the mechanical safeguard was triggered. Contact Siemens Healthineers Customer Service to replace the mechanical safeguard.
622 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.4 | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 4.45.2.4.1 | 562424 | Cause | Overtravel / movement range exceeds range | 4.45.2.4.1.1 | 562425 | Measure | End stop | Every movement axle shall be equipped with mechanical end stops at the end of the movement range.
623 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.4 | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 4.45.2.4.2 | 562426 | Cause | Automatic 3D-scan | 4.45.2.4.2.1 | 578319 | Measure | Reference - Continuous activation of movements with redundant signals | <XREF>562377</XREF>
624 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.4 | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 4.45.2.4.3 | 562428 | Cause | Weight of the device | 4.45.2.4.3.1 | 562429 | Measure | User Manual - Fast Braking | CAUTIONCause: Weight of the devicePossible consequences: Risk of crushingRemedy:Â Fast braking can be impacted by the weight of the device. Pay attention to your speed as well as the floor covering, slope, and uneven areas.
625 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.4 | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 4.45.2.4.4 | 562430 | Cause | Mounting RCP on OR table holder | 4.45.2.4.4.1 | 562431 | Measure | Design without crushing zones | The OR table holder shall have a flat design (e.g. Â without beadings), soÂ size of emerging gaps between OR table holder and mounted RCP are above threshold (> 25 mm), inÂ accordance to finger gaps for adults in IEC60601-1:2005 clause 9. 2. 2. 2 (Table 20). In caseÂ of Â a non-flat design of OR table holder a protection for all possible finger positions during mounting of RCP shall be available.
626 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.5 | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 4.45.2.5.1 | 562433 | Cause | Brake failure | 4.45.2.5.1.1 | 562434 | Measure | User Manual - Daily Inspection | CAUTIONCause: Brake failurePossible Consequences: Crushing due to arbitrary movementsRemedy:Â Before beginning the examination, perform the daily function and safety checks.
627 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.5 | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 4.45.2.5.2 | 562435 | Cause | Braking power not sufficient | 4.45.2.5.2.1 | 562436 | Measure | Specify sufficient braking power | Specify sufficient braking power. This is in compliance with IEC 60601-1:2005;Â clauseÂ 9. 4. 3 (Instability).
628 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.5 | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 4.45.2.5.2 | 562435 | Cause | Braking power not sufficient | 4.45.2.5.2.2 | 562437 | Measure | User Manual - Parking Chassis on inclined Plane (i.e. Ramp) | CAUTIONCause: Insufficient braking force. Possible consequences: Risk of crushing!Remedy: When parking the C-arm system on an inclined surface drive the vertical colum completely down. Verify that the brakes are locked and that the main unit does not slide away.
629 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.6 | 562439 | Hazard | Risk of crushing during movement execution | 4.45.2.6.1 | 562440 | Cause | Distance between operator rail and other components too small | 4.45.2.6.1.1 | 562441 | Measure | Compliance with distances according to IEC 60601-1 | Compliance with DIN EN 60601-1 (2006) Kap. 9. 2 Table 20, Acceptable gaps for finger > 25mm
630 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.6 | 562439 | Hazard | Risk of crushing during movement execution | 4.45.2.6.1 | 562440 | Cause | Distance between operator rail and other components too small | 4.45.2.6.1.2 | 562442 | Measure | User Manual - Distance Operator Rail | CAUTIONCause: The distance between the handles and other components is too small. Possible consequences: Risk of crushing!Remedy: Be careful around the hazardous locations indicated.
631 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.6 | 562439 | Hazard | Risk of crushing during movement execution | 4.45.2.6.1 | 562440 | Cause | Distance between operator rail and other components too small | 4.45.2.6.1.3 | 562443 | Measure | User Manual - Appropriate Distance | CAUTIONCause: The distance between the patient or other persons and the system is too small. Possible consequences: Risk of crushing!Remedy:Â Use the system only if the patient and other persons can be observed during movements.
632 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.6 | 562439 | Hazard | Risk of crushing during movement execution | 4.45.2.6.1 | 562440 | Cause | Distance between operator rail and other components too small | 4.45.2.6.1.4 | 562444 | Measure | User Manual - Guide Openings | CAUTIONCause:Â C-arm rotation. Possible consequences: Risk of crushingRemedy: Pay particular attention to crushing risks between moving system parts and the corresponding guide openings.
633 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.6 | 562439 | Hazard | Risk of crushing during movement execution | 4.45.2.6.2 | 562445 | Cause | Activation of movement without intervisibility | 4.45.2.6.2.1 | 562446 | Measure | User Manual - Observe Patient/Persons during Movement | CAUTIONCause: Activation of system movements without visual contactPossible Consequences: Risk of crushingRemedy: Use the system only if the patient and other persons can be observed during movements.
634 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.8 | 562450 | Hazard | Injuries through OR-instruments | 4.45.2.8.1 | 562451 | Cause | C-arm movements during operations | 4.45.2.8.1.1 | 562452 | Measure | User Manual - Collision with OR Instruments | WARNINGCause: C-arm movements during surgeryPossible consequences: Risk of injuriesRemedy:Â Due to the maneuverability of the C-arm, collisions with other surgical devices are possible if the device is not used as intended and in accordance with the instructions in this manual.
635 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.9 | 562453 | Hazard | Misjudgment due to poor image quality | 4.45.2.9.1 | 562454 | Cause | calibration incorrect | 4.45.2.9.1.1 | 562455 | Measure | Calibration cycles | Calibration cycles are monitored, when exceeding, notification is sent to operator. Â Note:DEXI adjustment is not affected by mentioned calibration cycles.
636 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.10 | 562456 | Hazard | Injury through dropping parts | 4.45.2.10.1 | 562457 | Cause | Collision with other systems | 4.45.2.10.1.1 | 562458 | Measure | User Manual - Collisions due to misuse | CAUTIONCause: Collision with other systemsPossible consequences: Risk of injuryRemedy: Due to the maneuverability of the C-arm, collisions with other systems are possible if the device is not used as intended and in accordance with the instructions in this manual.
637 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.11 | 562468 | Hazard | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) | 4.45.2.11.1 | 562469 | Cause | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) | 4.45.2.11.1.1 | 562470 | Measure | No measure required | No measure required.
638 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.2 | 562374 | Function | Position System | 4.45.2.12 | 562471 | Hazard | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) | 4.45.2.12.1 | 562472 | Cause | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) | 4.45.2.12.1.1 | 562473 | Measure | No measure required | No measure required.
639 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.3 | 562474 | Function | Wrap c-arm with Sterile Cover | 4.45.3.1 | 562475 | Hazard | Risk of infection | 4.45.3.1.1 | 562476 | Cause | Improperly installed sterile drape | 4.45.3.1.1.1 | 562477 | Measure | Reference to User Manual - Cleaning and draping | <XREF>562240</XREF>.
640 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.4 | 562478 | Function | Wrap Foot Switch with Covers | 4.45.4.1 | 562479 | Hazard | Risk of infection | 4.45.4.1.1 | 562480 | Cause | Improperly installed sterile drape | 4.45.4.1.1.1 | 562481 | Measure | Reference to User Manual - Cleaning and draping | <XREF>562240</XREF>.
641 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.5 | 562482 | Function | Wrap RCP with Covers | 4.45.5.1 | 562483 | Hazard | Risk of infection | 4.45.5.1.1 | 562484 | Cause | Damage of sterile drape | 4.45.5.1.1.1 | 562485 | Measure | Avoid sharp edges | The control panel (RCP) as well as its holder, e.g. cartÂ needs to be designed without sharp edges which could damage a sterile cover.
642 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.5 | 562482 | Function | Wrap RCP with Covers | 4.45.5.1 | 562483 | Hazard | Risk of infection | 4.45.5.1.2 | 562486 | Cause | Inappropriate positioning of RCP | 4.45.5.1.2.1 | 562487 | Measure | User Manual - Positioning of RCP | CAUTIONCause: Improper positioning of the remote control unit. Possible consequences: Risk of contaminationRemedy: Do not place the remote control unit on the patient.
643 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.5 | 562482 | Function | Wrap RCP with Covers | 4.45.5.1 | 562483 | Hazard | Risk of infection | 4.45.5.1.3 | 562488 | Cause | Self-loosening of OR table holder | 4.45.5.1.3.1 | 562489 | Measure | Fixation on OR table rail system | The OR table holder shall be fixed on OR table rail system. FixationÂ needs to be designed to prevent unintended self-loosening. Â Note:As drapery of patient is trapped between OR table holder and RCP, an unintended self-loosening of OR table holder will undrape patient.
644 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.1 | 562492 | Cause | Increasing temperature due to extended use | 4.45.6.1.1.1 | 562493 | Measure | Provide Monobloc Housing | Monobloc needs to be covered with housing to prohibit direct contact of user/patients to parts ofÂ monoblock.
645 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.1 | 562492 | Cause | Increasing temperature due to extended use | 4.45.6.1.1.2 | 562494 | Measure | Reference - Pressure switch in tube | <XREF>562232</XREF>
646 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.1 | 562492 | Cause | Increasing temperature due to extended use | 4.45.6.1.1.3 | 562495 | Measure | Temperature monitoring monoblock | The temperature of the monoblock shall be monitored. X-ray shall be inhibited when the maximum operating temperature is reached.
647 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.2 | 562497 | Cause | Oil leakage out of leaky monoblock | 4.45.6.1.2.1 | 562499 | Measure | Pressure Compensation Membrane | A pressure compensation membrane for oil expansion has to be integrated in monoblock.
648 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.2 | 562497 | Cause | Oil leakage out of leaky monoblock | 4.45.6.1.2.2 | 578321 | Measure | Reference - Temperature monitoring monoblock | <XREF>562495</XREF>
649 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.2 | 562497 | Cause | Oil leakage out of leaky monoblock | 4.45.6.1.2.3 | 562500 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
650 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.2 | 562497 | Cause | Oil leakage out of leaky monoblock | 4.45.6.1.2.4 | 562501 | Measure | Reference - Pressure switch in tube | <XREF>562232</XREF>
651 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.3 | 562502 | Cause | Error of cooling system | 4.45.6.1.3.1 | 562503 | Measure | Reference - Temperature Control | <XREF>562495</XREF>
652 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.1 | 562491 | Hazard | Risk of burnings | 4.45.6.1.4 | 562504 | Cause | Longer operation heats up detector | 4.45.6.1.4.1 | 562505 | Measure | Provide contact protection/access cover for detector | No direct contact possible due to access cover of the detector.
653 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.2 | 562507 | Hazard | Risk of infection | 4.45.6.2.1 | 562508 | Cause | Oil leakage out of leaky monoblock | 4.45.6.2.1.1 | 562509 | Measure | Stability test | The monobloc must be designed in a way that it withstands excessive energy inputs certain malfunctionsÂ Remark:Test of X-ray tube compliance IEC 60601-2-28 (2010) Kap. 9. 7
654 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.2 | 562507 | Hazard | Risk of infection | 4.45.6.2.1 | 562508 | Cause | Oil leakage out of leaky monoblock | 4.45.6.2.1.2 | 562510 | Measure | Reference - Pressure switch in tube | <XREF>562232</XREF>
655 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.2 | 562507 | Hazard | Risk of infection | 4.45.6.2.1 | 562508 | Cause | Oil leakage out of leaky monoblock | 4.45.6.2.1.3 | 562511 | Measure | Reference - Temperature control | <XREF>562495</XREF>
656 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.2 | 562507 | Hazard | Risk of infection | 4.45.6.2.1 | 562508 | Cause | Oil leakage out of leaky monoblock | 4.45.6.2.1.4 | 562512 | Measure | Reference - Check leakage of monobloc | <XREF>563289</XREF>
657 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.2 | 562507 | Hazard | Risk of infection | 4.45.6.2.2 | 562513 | Cause | Cooling circuit leaky, cooling medium drips into OR-field | 4.45.6.2.2.1 | 562514 | Measure | Reference to Cooling Unit Leakage Check | SeeÂ <XREF>563289</XREF> and <XREF>563364</XREF>.
658 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.3 | 562515 | Hazard | Injury through expelled parts | 4.45.6.3.1 | 562516 | Cause | Breakage rotating anode | 4.45.6.3.1.1 | 562517 | Measure | Reference - Stability test | <XREF>562509</XREF>
659 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.3 | 562515 | Hazard | Injury through expelled parts | 4.45.6.3.2 | 562518 | Cause | Implosion | 4.45.6.3.2.1 | 562519 | Measure | Reference to stability test and constructive measures | See <XREF>562509</XREF> andÂ HM_S1_mec_F4G3U1M1 in <XREF>578770</XREF>.
660 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.1 | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 4.45.6.4.1.1 | 562522 | Measure | Reference - Mounted under covers | <XREF>562126</XREF>
661 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.1 | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 4.45.6.4.1.2 | 562523 | Measure | Protection of the energy storage | Protection of the energy storage by a switch off of the short circuit in <= 20 ms. [HM_S1_esu_F6G3U1M1 in <XREF>578770</XREF>]
662 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.1 | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 4.45.6.4.1.3 | 562524 | Measure | Rupture joint by Overcharge | A rupture joint which breaks at an excessive overcharge must be available at the capacitors. [HM_S1_esu_F6G3U2M1Â in <XREF>578770</XREF>]
663 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.1 | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 4.45.6.4.1.4 | 562525 | Measure | Pressure relief valve | Pressure relief valve at the capacitors prevents an excessive pressurization. [HM_S1_esu_F6G3U2M1Â in <XREF>578770</XREF>]
664 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.1 | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 4.45.6.4.1.5 | 562526 | Measure | Reference to monitoring of the intermediate circuit voltage | See HM_S1_hw_F6G3U2M2 for ESU and Elko,Â HM_S1_esu_F6G3U2M3 only for ESU in <XREF>578770</XREF>.
665 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.2 | 562527 | Cause | Short-circuit of electrical components | 4.45.6.4.2.1 | 562528 | Measure | Protection against excessive current and voltage | The system shall provide a protection against excessive current and voltage. Alternatively, a test shall show that a short circuit does not result in a hazardous situation.
666 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.2 | 562527 | Cause | Short-circuit of electrical components | 4.45.6.4.2.2 | 562529 | Measure | Protection against excessive current and voltage - PowerUnit | The PowerUnit shall provide a protection against excessive current and voltage.
667 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.4 | 562520 | Hazard | Excessive temperature, fire, explosion | 4.45.6.4.2 | 562527 | Cause | Short-circuit of electrical components | 4.45.6.4.2.3 | 562530 | Measure | Reference to generator measures to UL listed materials and fuses in intermediate circuit | See HM_S1_ul_F2G1U1M1 andÂ HM_S1_hw_F2G1U1M2 in <XREF>578770</XREF>.
668 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.5 | 562531 | Hazard | Implosion | 4.45.6.5.1 | 562532 | Cause | Collision/impact against component | 4.45.6.5.1.1 | 562533 | Measure | Mechanical strength of the X-ray generation | Mechanical strength of the monobloc has to be verified by the manufacturer
669 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.5 | 562531 | Hazard | Implosion | 4.45.6.5.1 | 562532 | Cause | Collision/impact against component | 4.45.6.5.1.2 | 562534 | Measure | Reference - Stability test | <XREF>562509</XREF>
670 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.5 | 562531 | Hazard | Implosion | 4.45.6.5.2 | 562535 | Cause | Breakage rotating anode | 4.45.6.5.2.1 | 562536 | Measure | Housing has to stay integral | Housing has to stay integral, ifÂ there is a breakage of the rotating anode.
671 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.5 | 562531 | Hazard | Implosion | 4.45.6.5.2 | 562535 | Cause | Breakage rotating anode | 4.45.6.5.2.2 | 562537 | Measure | Reference - Stability test | <XREF>562509</XREF>
672 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.5 | 562531 | Hazard | Implosion | 4.45.6.5.2 | 562535 | Cause | Breakage rotating anode | 4.45.6.5.2.3 | 562538 | Measure | Reference to constructive measures | See HM_S1_mec_F4G3U1M1 in <XREF>578770</XREF>.
673 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.6 | 562539 | Hazard | Risk of injury due to bad surface finish | 4.45.6.6.1 | 562540 | Cause | Sharp edges | 4.45.6.6.1.1 | 562541 | Measure | Develop design without sharp edges | <XREF>562047</XREF>
674 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.6 | 562539 | Hazard | Risk of injury due to bad surface finish | 4.45.6.6.2 | 562542 | Cause | Rough surface finish | 4.45.6.6.2.1 | 562543 | Measure | Reference - Specify smooth surface finishes | <XREF>562050</XREF>
675 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.7 | 562544 | Hazard | Injury due to damaged housing | 4.45.6.7.1 | 562545 | Cause | Transport damages e.g. collision | 4.45.6.7.1.1 | 562546 | Measure | User Manual - Professional Repair | CAUTIONCause: Transport damagePossible consequences: Risk of injuryRemedy: Carry out professional repairs only, otherwise consequential damage to the product cannot be ruled out!
676 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.8 | 562547 | Hazard | Electric shock | 4.45.6.8.1 | 562548 | Cause | Intrusion of liquids | 4.45.6.8.1.1 | 562549 | Measure | Spillage of Liquids on System | Spillage of liquids on systemÂ shall not result in a hazardous situation. The system shall be designed in accordance to IEC 60601-1 clause 11. 6. 3 (compliance to spillage test). Â The spillage test shall be performed with following instructions:Â Spillage test, scenario 1: * Scenario: capacity of drapes/ max. suction performance exhausted * liquid is seeping continuously from operating area downwards on the system * Quantity: 500 ml; duration: 30 min * Effected parts of system: FD/I. I. , monoblock, C-arm (without Chassis)Â Spillage test, scenario 2: * Scenario: splatter of blood * liquid is splashing from operating area up- and/or sidewards to the system * Quantity: 20 ml; duration: approx. 1 s (i.e. a bolus) * Effected parts of system: FD/I. I. , monoblock, C-arm (without Chassis)Â No liquids need to be spilled on trolley (incl. trolley control panel, TCP). Â Hints:Spilling of FD is only allowed with coverage ofÂ drape (see <XREF>562552</XREF>). Remote control panel (RCP) will be used with drapes only.
677 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.8 | 562547 | Hazard | Electric shock | 4.45.6.8.1 | 562548 | Cause | Intrusion of liquids | 4.45.6.8.1.2 | 562552 | Measure | User Manual - Draping of FD | WARNING:Cause: Fluid penetration and contamination. Possible Consequences: Damage to the flat detectorRemedy: Make sure that the detector is always covered by a disposable cover during the examination.
678 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.1 | 562554 | Cause | No radiation protection behind the entering side of radiation | 4.45.6.9.1.1 | 562555 | Measure | Compliance with DIN EN 60601-1-3 | Compliance with DIN EN 60601-1-3 (2008); Kap. 11 residual radiation
679 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.2 | 562556 | Cause | Increased leakage radiation monoblock | 4.45.6.9.2.1 | 562557 | Measure | Reference to constructive measures and tests to comply with leakage radiation requirements | See HM_S1_mec_F5G1U3M1, HM_S1_manu_F5G1U3M2, HM_S1_mec_F5G1U3M3 inÂ <XREF>578770</XREF>.
680 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.3 | 562558 | Cause | Stay in endangered area | 4.45.6.9.3.1 | 562559 | Measure | User Manual - Unwanted Radiation | CAUTIONCause: Remaining in hazardous area. Possible consequences: Danger of inadvertent radiation!Remedy: During radiation release, make sure you are in a location subject to as little radiation as possible.
681 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.4 | 562560 | Cause | Accidental release of radiation | 4.45.6.9.4.1 | 562561 | Measure | Design of X-ray release controls | The X-ray release controls shall be designed to preventÂ accidental activation. For the standard footswitch it is acceptable that putting the switch upside down does not activate a function.
682 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.4 | 562560 | Cause | Accidental release of radiation | 4.45.6.9.4.2 | 562562 | Measure | Display Radiation On | A light signal interface will be provided by the generator when X-ray is on.
683 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.5 | 562564 | Cause | Malfunction of exposure release circuit. | 4.45.6.9.5.1 | 562565 | Measure | Hardware signal | InÂ addition to SW aÂ separate hardware signal shall be checked to be set ""inactive"" during start-up of the x-ray generator.
684 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.5 | 562564 | Cause | Malfunction of exposure release circuit. | 4.45.6.9.5.2 | 580167 | Measure | Quasi-continuous activation of radiation with redundant activation paths | Radiation shall be possible only when two redundant signals provided by two independent paths are present within certain time constraints (quasi-continuous activation), preventingÂ unintended X-ray activation. When only one of two signals is present, radiation shall be stopped and inhibited for the respective control. ãNote: Safety architecture C + P, see <XREF>561860</XREF> and <XREF>543971</XREF>. Â For design constraints for single fault safety, see <XREF>578311</XREF>.
685 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.5 | 562564 | Cause | Malfunction of exposure release circuit. | 4.45.6.9.5.3 | 562567 | Measure | Reference to generator safety circuit and test at start-up | See HM_S1_hw_F1G1U1M2, HM_S1_sw_F3G2U1M2 in <XREF>578770</XREF>.
686 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.5 | 562564 | Cause | Malfunction of exposure release circuit. | 4.45.6.9.5.4 | 578773 | Measure | Reference - Emergency stop radiation | <XREF>562135</XREF>.
687 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.6 | 562568 | Cause | Loss of connection between FD and imaging system | 4.45.6.9.6.1 | 562569 | Measure | Reference - Imaging system hardware test at startup | <XREF>562637</XREF>
688 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.7 | 562570 | Cause | High radiation dose due to excessive procedure duration and / or heavily inadequate radiation settings without clinical benefi | 4.45.6.9.7.1 | 562571 | Measure | Reference to Warning - Trained Users | <XREF>561986</XREF>.
689 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.8 | 562572 | Cause | Bad image quality of single image - make single image without having done flluoro before | 4.45.6.9.8.1 | 562573 | Measure | No additional measures required | No additional measures are required. The operators are trained on system behavior and correct operation. The description is included in the operator manual (""advice to Fluoro exposure with > 2fps before acquiring a single image in order to set the exposure parameters and to achieve good image quality).
690 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.9 | 562574 | Cause | Failed subtraction / roadmap due to operator error or configuration error (licensing,...) | 4.45.6.9.9.1 | 562575 | Measure | No additional measures required | No additional measures are required. The operators are trained on system behavior and correct operation / workflow.
691 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.9 | 562553 | Hazard | Unwanted radiation | 4.45.6.9.10 | 562576 | Cause | Radiation is not switched off or switched off too late (e. g. pulses) due to HW / SW errors | 4.45.6.9.10.1 | 562577 | Measure | Reference to measures for radiation switch off | <XREF>562717</XREF> and HM_S1_hw_F5G1U1M2 (HW radiation timer), HM_S1_hw_F5G1U2M1 (watchdog), HM_S1_hw_F5G1U2M4 (mAs switch off by HW), HM_S1_sw_F5G1U3M1 (CAN monitoring) in <XREF>578770</XREF>.
692 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.1 | 562579 | Cause | Image orientation mixed up due to image rotation | 4.45.6.10.1.1 | 562580 | Measure | User Manual - Image Orientation | CAUTIONCause: Inverted image orientation. Possible consequences: Risk of an incorrect decision during surgery. Remedy: Prior to releasing radiation, the user must ensure the correct patient position with respect to the X-ray beam.
693 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.2 | 562581 | Cause | Image orientation mixed up due to flipped image | 4.45.6.10.2.1 | 562582 | Measure | User Manual - Label each side | CAUTIONCause: The image orientation may be inverted, e.g. if 2D images are displayed on external viewers. Possible consequences: Risk of an incorrect therapeutic decision. Remedy: The operator must apply due diligence to correctly identify the 2 D image orientation.
694 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.3 | 562583 | Cause | Image orientation mixed up | 4.45.6.10.3.1 | 562584 | Measure | User Manual - Image Orientation | CAUTIONCause: Inverted image orientation. Possible Consequences:Â´Incorrect intervention decision!Remedy:Â It is recommended to mark the patient position using lead letters.
695 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.4 | 562586 | Cause | Wrong image orientation applied by imaging system | 4.45.6.10.4.1 | 562587 | Measure | Encapsulation for image flip and rotate | Encapsulation and separate software tests of image flip/rotate.
696 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.5 | 562588 | Cause | Wrong calibration display | 4.45.6.10.5.1 | 562589 | Measure | Update calibration | After each calibration the length of the lines in the image is updated. If no calibration has been performed, the length of the lines is displayed in pixels.
697 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.5 | 562588 | Cause | Wrong calibration display | 4.45.6.10.5.2 | 562590 | Measure | User Manual - 2D measurements | CAUTIONCause: Wrong calibration value is entered or length or angle information are not interpreted correctlyPossible Consequences: Insufficient or incorrect basis for diagnostic or therapeutic decisionsRemedy: Note that the measuring results of the functions Distance and Angle depend on various factors, e.g. the position of the calibration object. Note that the acquired images represent two-dimensional projections of real objects. Verify the results before diagnostic or therapeutic decisions are taken.
698 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.6 | 562591 | Cause | Wrong interpretation of results after calibration | 4.45.6.10.6.1 | 562592 | Measure | Reference to Caution for 2D measurements | <XREF>562590</XREF>
699 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.7 | 562593 | Cause | Wrong mask image for subtraction | 4.45.6.10.7.1 | 562594 | Measure | Correct mask calculation | The FLC software triggers the start of the mask calculation phase depending on the system state (subtraction or roadmap).
700 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.8 | 562595 | Cause | Movements (either extrinsic or intrinsic) or image manipulations (e.g. AAA line) | 4.45.6.10.8.1 | 562596 | Measure | User Manual - Mismatch of AAA Line | WARNINGCause: An annotated image with graphic overlay lines is displayed. Following system and/or table or patient movements and/or image modifications, subsequent images may not match the overlay line displayed due to projective changes. Â Possible Consequences:Â Incorrect interpretation of the display and resulting incorrect decision during operation!Â Remedy: The graphic overlay line is displayed until the user actively turns it off. This is done independently of any changes (for example, movements, operating mode changes, image modifications). The user must ensure that the displayed overlay line matches the current, underlying image (for example, display of anatomical structures via the label function and/or use of contrast agent).
701 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.9 | 562599 | Cause | Image orientation reset due to application change | 4.45.6.10.9.1 | 562600 | Measure | User Manual - Image orientation reset | CAUTIONCause: The image orientation is reset when changing the program. Possible consequences: Risk of an incorrect decision during surgery!Remedy: Be aware that the image orientation is reset when changing the program. All previous settings are lost.
702 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.10 | 562601 | Cause | LIH is confused with live image | 4.45.6.10.10.1 | 562602 | Measure | Clearance of LIH during radiation | When the user starts radiation, FLC clears the LIH image.
703 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.11 | 562603 | Cause | Live image covered by messages or UI elements | 4.45.6.10.11.1 | 562604 | Measure | Keep live imaging in foreground | The displayed live image is not covered by dialogs or other UI elements.
704 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.12 | 562605 | Cause | Incorrect K-wire display | 4.45.6.10.12.1 | 562606 | Measure | User Manual - Incorrect K-wire display | CAUTIONCause: Time gap between display of K-wire including overlay grafic and real moved K-wire. False positive or false negative detection of K-wires. Possible consequences: Risk of confusion and wrong interventional decisionRemedy: Duty of care for operator, that he/she shall be aware of this possible mismatch.
705 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.13 | 562607 | Cause | overlay hide K-wire on image | 4.45.6.10.13.1 | 562608 | Measure | User Manual - overlay hide K-wire on image | CAUTIONCause: Overlay itself may hide K-wire in live X-ray image (depends on thickness of K-wire and overlay configuration). Possible consequences: Risk of wrong interventional decisionRemedy: Target Pointer can be switched-off (toggle ON/OFF) to allow free view on X-ray live image.
706 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.10 | 562578 | Hazard | Wrong interventional decision | 4.45.6.10.13 | 562607 | Cause | overlay hide K-wire on image | 4.45.6.10.13.2 | 562609 | Measure | Visibility of top of K-wire | Visibility of the top of K-Wire happens via colored transparent lines overlaying the image.
707 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.11 | 562610 | Hazard | Loss of images | 4.45.6.11.1 | 562611 | Cause | Insufficient image storage space | 4.45.6.11.1.1 | 562612 | Measure | Storage capacity exhausted | When the image storage limit is reached, the system shall inhibit further radiography. Note: If the limit is reached during an acquisition, this acquisition will be continued until finished and stored due to a small buffer. ãFluoroscopy shall not be stopped or inhibited, and a message shall be displayed to the operator. LIH functionality shall not be impaired in any case.
708 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.11 | 562610 | Hazard | Loss of images | 4.45.6.11.1 | 562611 | Cause | Insufficient image storage space | 4.45.6.11.1.2 | 562613 | Measure | Display storage space | On monitorÂ the actual storage capacity is shown and updated when an image is stored. The image system displays the storage space still available as relative part of the assignable capacity. A message shall be shown to the user whenÂ the available storage capacity is equal or below 10%. Â Note: The accuracy of the displayed storage capacity should be +/- 5% of the used split ration for the icons.
709 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.11 | 562610 | Hazard | Loss of images | 4.45.6.11.2 | 562614 | Cause | Images / series lost in imaging system | 4.45.6.11.2.1 | 569132 | Measure | Immediate image storage on hard disk | Acquired images are always stored immediately on hard disk.
710 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.11 | 562610 | Hazard | Loss of images | 4.45.6.11.2 | 562614 | Cause | Images / series lost in imaging system | 4.45.6.11.2.2 | 578317 | Measure | Full storage rate radiography | For radiography, the storage rate shall not be changeable by the customer. Every frame shall be stored. Replay of scenes shall apply the same behaviour as live (fps). Â Note, the default storage rate shall be set to 100%.
711 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.11 | 562610 | Hazard | Loss of images | 4.45.6.11.3 | 562616 | Cause | Reducing the storage rate or being confused about replay | 4.45.6.11.3.1 | 562617 | Measure | Reference - Immediate image storage on hard disk. | <XREF>569132</XREF>.
712 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.12 | 562618 | Hazard | Unwanted Injection | 4.45.6.12.1 | 562619 | Cause | Unwanted radiation release | 4.45.6.12.1.1 | 562620 | Measure | No measure required | no measure required due to risk classification
713 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.12 | 562618 | Hazard | Unwanted Injection | 4.45.6.12.2 | 562621 | Cause | Triggersignal failure | 4.45.6.12.2.1 | 562622 | Measure | No measure required | no measure required due to risk classification
714 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.13 | 580664 | Hazard | Unintended injection | 4.45.6.13.1 | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 4.45.6.13.1.1 | 580666 | Measure | Reference - User Manual - Injector Readiness | <XREF>562625</XREF>
715 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.13 | 580664 | Hazard | Unintended injection | 4.45.6.13.1 | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 4.45.6.13.1.2 | 581162 | Measure | Injector trigger in dedicated modes and conditions only | The system shall trigger the injector only in the following modes andÂ conditions:- Subtraction- Roadmap: Only when the operator has not performed a subtraction before (then automatically a subtraction scene is performed when roadmap is selected), or when no valid peak opacification image is available (e.g. after changing magnification)- Digital Cine ModeÂ Note: That means when a subtraction has been performed before roadmap is selected, then the injector shall not be triggered when roadmap is performed.
716 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.13 | 580664 | Hazard | Unintended injection | 4.45.6.13.1 | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 4.45.6.13.1.3 | 581195 | Measure | Display when injector is armed | The system shall display when the injector is armed, when the conditions under <XREF>581162</XREF> are fulfilled (that means that upon next X-ray release, the injector is triggered).
717 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.14 | 562623 | Hazard | Radiation without injection | 4.45.6.14.1 | 562624 | Cause | Acquire image, but injector not prepared correctly | 4.45.6.14.1.1 | 562625 | Measure | User Manual - Injector Readiness | CAUTIONCause: Radiation release without without contrast agent injection or unintended injection. Possible consequences: RiskÂ of unnecessary radiation exposure or unintended contrast agent injection. Remedy: Make sure that the injector is ready to administer contrast agent before releasing radiation. Note that every X-ray mode triggers the injector when it is in ready state.
718 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.14 | 562623 | Hazard | Radiation without injection | 4.45.6.14.2 | 581193 | Cause | Injector interface is damaged due to motorized platform movement with plugged injector connection | 4.45.6.14.2.1 | 581194 | Measure | Restricted platform movement range with plugged injector | When the injector is plugged, motorized movements of the mecanum platform shall beÂ restricted to a distance of < 50 cm. When it is reached, the system shall stop motorized movements and display an appropriate error message.
719 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.14 | 562623 | Hazard | Radiation without injection | 4.45.6.14.2 | 581193 | Cause | Injector interface is damaged due to motorized platform movement with plugged injector connection | 4.45.6.14.2.2 | 595189 | Measure | Unplug injector cable during movement | CAUTIONCause: A strain might be applied to the injector cable when plugged during movements. Possible consequence: The injector cable can get damaged or unplugged during a chassis movement of the system. Remedy: Unplug the injector cable if you perform longer movements that might put tension on the cable.
720 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.15 | 562626 | Hazard | Radiation interrupted | 4.45.6.15.1 | 562627 | Cause | Exceeding maximum delay time | 4.45.6.15.1.1 | 568834 | Measure | Image display delay time | An image display delay time of max. 200 ms is determined to beÂ practicable based on the image processing capabilities of the system andÂ to ensure sufficient reaction time for the operator for all imaging modes except Single Image. Â Note, the achieved image display delay time will then be below the average human reaction time of ~ 300 ms according to Sigmund Exner. For Single Image, which includes DR (Digital Radiography), DR-SUM (Digital Radiography with several pulses) and 3D Scan,Â as well as DCM and Subtraction, the image display delay time might be up to 4 s.
721 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.15 | 562626 | Hazard | Radiation interrupted | 4.45.6.15.1 | 562627 | Cause | Exceeding maximum delay time | 4.45.6.15.1.2 | 568835 | Measure | Image display delay time - external | If an external device is utilized to display images acquired with the CIOS system via theÂ NXS interface,Â the image display delay time shall be evaluated as part of the vendor risk management.
722 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.15 | 562626 | Hazard | Radiation interrupted | 4.45.6.15.1 | 562627 | Cause | Exceeding maximum delay time | 4.45.6.15.1.3 | 596894 | Measure | Caution in user manual to image display delay | CAUTIONÂ Cause:Â Radiography modesÂ are misused for real-time imaging on purpose by the operator. Â Possible consequences: With Radiography modes the image display delay timeÂ may be longer than in Radioscopy modes and may exceed limits for accepted delays for real-time imaging. Â Remedy: For real-time imaging do only select Radioscopy modes such as Fluoro andÂ Road. Do not select Radiography modes (e.g. single image including DR, DR-Sum, 3D Scan, DCM, Subtraction)Â for real-time imaging.
723 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.16 | 562629 | Hazard | Restricted imaging functionality due to low-level arcing | 4.45.6.16.1 | 562630 | Cause | Arcing in the X-ray tube | 4.45.6.16.1.1 | 562631 | Measure | No measure required | No additional risk control measures required.
724 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.17 | 562632 | Hazard | Restricted imaging functionality due to medium-level arcing | 4.45.6.17.1 | 562633 | Cause | Arcing in the X-ray tube | 4.45.6.17.1.1 | 562634 | Measure | No measure required | No additional risk control measures required.
725 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.18 | 562635 | Hazard | Image quality is reduced | 4.45.6.18.1 | 562636 | Cause | FLC hardware is defective | 4.45.6.18.1.1 | 562637 | Measure | Imaging system hardware test at startup | The imaging system hardware shall be tested at startup, including the connection to the detector. In case an error is detected, the imaging system shall not report ready state (see <XREF>578746</XREF>).
726 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.18 | 562635 | Hazard | Image quality is reduced | 4.45.6.18.1 | 562636 | Cause | FLC hardware is defective | 4.45.6.18.1.2 | 599844 | Measure | User Manual - High-Energy-Isotopes | CautionÂ Cause: Flat detector hardware is exposed to high-energy isotopes at low proximity, e.g. by means of Cobalt-60Â radiation sources during BrachytherapyPossible consequences: Flat detector hardware damaged, leading to low image quality or image artifactsRemedy: Keep the detector system and electronics in a safe distance (> 3 m) from high-energy isotopes (> 1 MeV), e.g. , during the insertion or removal process of the radiation probe inÂ Brachytherapy.
727 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.18 | 562635 | Hazard | Image quality is reduced | 4.45.6.18.2 | 562638 | Cause | Numerical calculation error or parameterization issue in imaging pipeline | 4.45.6.18.2.1 | 562639 | Measure | Consistency check in imaging pipeline | Consistency check in imaging pipeline and notification of host in case of error.
728 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.18 | 562635 | Hazard | Image quality is reduced | 4.45.6.18.3 | 562640 | Cause | Detector heats up out of calibration range | 4.45.6.18.3.1 | 562641 | Measure | Reference - System Warning - Temperature of the Detector | <XREF>562226</XREF>
729 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.6 | 562490 | Function | Acquire Image | 4.45.6.18 | 562635 | Hazard | Image quality is reduced | 4.45.6.18.3 | 562640 | Cause | Detector heats up out of calibration range | 4.45.6.18.3.2 | 562642 | Measure | Reference to Caution in Operator Manual - Detector temperature | <XREF>563301</XREF>.
730 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.7 | 562643 | Function | Patient browser | 4.45.7.1 | 562644 | Hazard | No hazard on system level | 4.45.7.1.1 | 562645 | Cause | None | 4.45.7.1.1.1 | 562646 | Measure | None
731 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.8 | 562647 | Function | Patient registration | 4.45.8.1 | 562648 | Hazard | Misinterpretation of patient data | 4.45.8.1.1 | 562649 | Cause | Patient-ID already in use | 4.45.8.1.1.1 | 562650 | Measure | User Manual - Patient ID | CAUTIONCause: Patient ID already exists. Possible Consequences: Misinterpretation of data of the examined patient if the selection is based solely on the patient ID. Remedy:-Always use four-way identification [patient name, sex, date of birth, patient ID]. - Avoid mixing up patient data by using four-way identification. Selecting patient data based on the patient ID alone is insufficient. Ensure the proper display of parameters in the Worklist and in the Patient list.
732 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.8 | 562647 | Function | Patient registration | 4.45.8.2 | 562651 | Hazard | Emergency login not possible | 4.45.8.2.1 | 562652 | Cause | Authenticated login, access not possible | 4.45.8.2.1.1 | 562653 | Measure | Emergency login without password protection | Emergency user login is possible without a password. Note: Due to PSS requirements, disk encryption is configurable: It is an option andÂ must be activated by serviceÂ on request. Consequences and boundary conditions are described in operator manual, admin guide and MDS2 statement. With these conditions, it is acceptable to restrict the effectivity of above mentioned measure.
733 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.9 | 562654 | Function | APS management | 4.45.9.1 | 562655 | Hazard | No hazard on system level | 4.45.9.1.1 | 562656 | Cause | None | 4.45.9.1.1.1 | 562657 | Measure | None
734 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.1 | 562661 | Measure | Reliable data transfer standards and protocols | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication)
735 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.2 | 562662 | Measure | NGS protocol with speed detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a link speed detection. Its also valid when NGS is used in the context of OpenApps. Incase the detected speed is below the threshold limit, the user shall be provided with anÂ audible or visualÂ indication.
736 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.3 | 562663 | Measure | NGS protocol with link loss detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a heartbeat and link loss detection. Its also valid for OpenApps. If the link loss is detected, it shall be indicated to the user and an automatic release of mouse events shall occur.
737 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.4 | 562664 | Measure | NGS protocol with latency detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a latency detection. This means measuringÂ the time between sending a frame tillÂ receiving this frame and in case of exceeding a configurable limit to indicate it with an audible or visual indicator. Â VxFUTURE:Â Â when using NGS in the OpenApps context, the latency detection shall be configurable specific to the app (i.e. either deactivated or threshold value adjusted based on the Risk Managmenent of the individual partner apps).
738 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.5 | 562665 | Measure | NGS protocol with update rate detection | Use of reliable data transfer standards and protocols (i.e. TCP/IP protocol for Ethernet-based communication) with a update rate detection between two frames. In case ofÂ exceeding a configurable timeÂ an audible or visual indication shall be done. Â Â VxFUTURE:Â This can be configured individually for partner apps (either deactivated or a threshold value configured based on the respective partner app's Risk Manamagement proposal)Â in case of OpenApps.
739 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.6 | 562666 | Measure | NGS protocol with data integrity | For transfer of frames via NGS the data integrity of these frames shall beÂ considered by using a check sum algorithm (e. g CRC32). In case ofÂ an incorrect checksum is detectedÂ an audible or visual indication shall be done. Â Â When NGS is used in context of OpenApps, the temporal integrity of an image series that builds on one another must also be monitored using a checksum. Incase of a missing image or incorrect CRC, an audible or visual indication shall be available.
740 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.7 | 569295 | Measure | Reference - Image display delay time | <XREF>568834</XREF>
741 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.8 | 569296 | Measure | Reference - Image display delay time - external | <XREF>568835</XREF>
742 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.1 | 562660 | Cause | Communication problems | 4.45.10.1.1.9 | 569238 | Measure | Caution - Reduced image quality streaming | CAUTIONãCause: When streaming via external systems, image quality (e.g. , resolution, frame rate or flickering) might be affected by connector type or external monitor hardware and settings. ãPossible Consequences: Reduced image qualityRemedy: In order to avoid loss of image quality ensure that the connection cable specifications and display parameters of the connected external monitors fulfill the required specifications as described in the chapter Technical SpecificationsÂ ------Â Not part of Caution: Note, the required specifications for the external monitor are medical grade, 4K, color display, 3840x2160 pixels, 60Hz, and 4:4:4 Chroma Subsampling. The cable requirements are a short HDMI2. 0 (or better) video cable (<3m) between CIOS Move and external monitor or fiber optic cable.
743 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.1 | 562659 | Hazard | Delayed or no interventional decision | 4.45.10.1.2 | 562667 | Cause | Lossy organprogram | 4.45.10.1.2.1 | 562668 | Measure | Reference - User Manual - Wrong exposure parameters | <XREF>562722</XREF>
744 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.2 | 562669 | Hazard | Unnecessary radiation | 4.45.10.2.1 | 562670 | Cause | Communication network disturbance -incorrect parameters | 4.45.10.2.1.1 | 562671 | Measure | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF>
745 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.2 | 562669 | Hazard | Unnecessary radiation | 4.45.10.2.2 | 562672 | Cause | Operating error | 4.45.10.2.2.1 | 562673 | Measure | Reference - User Manual - Trained Users | <XREF>561986</XREF>
746 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.2 | 562669 | Hazard | Unnecessary radiation | 4.45.10.2.2 | 562672 | Cause | Operating error | 4.45.10.2.2.2 | 568855 | Measure | Indication of loading factors | The system indicates the loading factors as per standard requirements, see 06S456591, and following, 06S434271, and following in <XREF>561879</XREF>.
747 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.3 | 562674 | Hazard | Dose not as desired | 4.45.10.3.1 | 562675 | Cause | Incorrect parameters due to communication/network disturbance | 4.45.10.3.1.1 | 562676 | Measure | Reference - Ready for radiation | <XREF>578746</XREF>
748 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.10 | 562658 | Function | Administration, selection and changing of applications | 4.45.10.3 | 562674 | Hazard | Dose not as desired | 4.45.10.3.1 | 562675 | Cause | Incorrect parameters due to communication/network disturbance | 4.45.10.3.1.2 | 562677 | Measure | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF>
749 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.11 | 562678 | Function | Display of patient relevant data on user interface | 4.45.11.1 | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 4.45.11.1.1 | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 4.45.11.1.1.1 | 562681 | Measure | Patient Data Display | The patient name and an additional configurable information entity (eitherÂ patient ID or request ID or accession number or date of birth) of the selected studyÂ shall beÂ displayed. Â Note: In case of limited number of characters the additional information shall be shown at least partially so thatÂ different patients with same names can be distinguished.
750 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.11 | 562678 | Function | Display of patient relevant data on user interface | 4.45.11.1 | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 4.45.11.1.1 | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 4.45.11.1.1.2 | 562682 | Measure | Reference to measures for patient or image data refresh and data consistency | <XREF>563323</XREF>
751 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.11 | 562678 | Function | Display of patient relevant data on user interface | 4.45.11.1 | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 4.45.11.1.1 | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 4.45.11.1.1.3 | 569126 | Measure | Refresh display after patient or image data correction | When patient or image data has been corrected, the imaging system refreshes the display.
752 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.11 | 562678 | Function | Display of patient relevant data on user interface | 4.45.11.1 | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 4.45.11.1.1 | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 4.45.11.1.1.4 | 569127 | Measure | Imaging system data consistency | When the image is processed for transfer, a consistency-check is done between the study UID in the job list and the study UID in the image header. For data exchange between FLC and RIS the DICOM3 standard is used.
753 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.11 | 562678 | Function | Display of patient relevant data on user interface | 4.45.11.1 | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 4.45.11.1.1 | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 4.45.11.1.1.5 | 569128 | Measure | Transfer images of current patient only | Only images of the selected patient can be transferred to the hardcopy camera
754 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.12 | 562683 | Function | Process patient exam list | 4.45.12.1 | 562684 | Hazard | No hazard on system level | 4.45.12.1.1 | 562685 | Cause | None | 4.45.12.1.1.1 | 562686 | Measure | None
755 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.13 | 562687 | Function | Presetting of collimator position | 4.45.13.1 | 562688 | Hazard | Unnecessary radiation | 4.45.13.1.1 | 562689 | Cause | Communication network disturbance -incorrect parameters | 4.45.13.1.1.1 | 562690 | Measure | Reference - Reliable data transfer standards and protocols | <XREF>562323</XREF>
756 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.13 | 562687 | Function | Presetting of collimator position | 4.45.13.1 | 562688 | Hazard | Unnecessary radiation | 4.45.13.1.2 | 562691 | Cause | Use Presettings of Collimator Position for Automatic Virtual Shutter | 4.45.13.1.2.1 | 562692 | Measure | No measure required | Note:Risk rating is in accordance to <XREF>562695</XREF> as this feature will effect measure <XREF>562696</XREF>. Nonetheless, no measure is required, as this feature can be switched off easily (via toggle button). Risk assessment was necessary due to feature "Automatic Virtual Shutter" (a. k. a. "set black", 04S407692 in <XREF>561854</XREF>).
757 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.14 | 562693 | Function | Exposure field | 4.45.14.1 | 562694 | Hazard | Unnecessary radiation | 4.45.14.1.1 | 562695 | Cause | Position of collimator leaves out of control due to stepping errors | 4.45.14.1.1.1 | 562696 | Measure | Perform Calibration of Collimator Leaves | Accuracy of collimator leaves position shall remain unchanged during examination. It is possible to perform a search run to calibrate position of collimatorÂ leaves. Â In case a search run will be implemented, it has to be ensured, that no search run will be triggered and/or performed during acquisition (i.e. whileÂ releasing X-ray).
758 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.15 | 562697 | Function | Wireless Remote control for movement functions | 4.45.15.1 | 562698 | Hazard | Injury by crushing | 4.45.15.1.1 | 562707 | Cause | Other wireless devices act as remote control | 4.45.15.1.1.1 | 562708 | Measure | Remote control and system are paired | Remote control and system are paired. No other wireless devices sending on the same frequency can interfere.
759 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.15 | 562697 | Function | Wireless Remote control for movement functions | 4.45.15.2 | 562709 | Hazard | Interfere with other devices | 4.45.15.2.1 | 562710 | Cause | Electromagnetic field disturbs life-supporting devices | 4.45.15.2.1.1 | 562711 | Measure | Reference to Operator Manual - Interference with life-supporting devices or critical equipment | <XREF>562061</XREF>.
760 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.15 | 562697 | Function | Wireless Remote control for movement functions | 4.45.15.2 | 562709 | Hazard | Interfere with other devices | 4.45.15.2.1 | 562710 | Cause | Electromagnetic field disturbs life-supporting devices | 4.45.15.2.1.2 | 562712 | Measure | Reference - EMC compliance - emissions, immunity and pass / fail criteria | <XREF>562015</XREF>
761 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.1 | 562715 | Cause | Dose controlling | 4.45.16.1.1.1 | 562716 | Measure | IQAP measurement | IQAP measurement after each dose adjustment.
762 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.1 | 562715 | Cause | Dose controlling | 4.45.16.1.1.2 | 562717 | Measure | Switch off radiation | Switch off radiation after max time.
763 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.1 | 562719 | Measure | Admin Guide - System calibration | CAUTIONCause: Incorrect calibration parameters degrade image quality. Possible consequences: Risk of unnecessary radiation exposure due to incorrect calibration parameters!Remedy: Calibrate the system after 24 months to preserve image quality.
764 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.2 | 569225 | Measure | Reference - Ready for radiation | <XREF>578746</XREF>
765 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.3 | 562721 | Measure | Display clinical procedure | The selected clinical procedureÂ shall beÂ shown on the monitor during examination.
766 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.4 | 562722 | Measure | User Manual - Wrong exposure parameters | CAUTIONCause: Incorrect exposure parametersPossible consequences: Risk of unnecessary X-ray exposure due to incorrect exposure parameters. Remedy: Ensure the exposure parameters are correct.
767 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.5 | 562723 | Measure | User Manual - Reduced Image Quality due to Filtering | <XREF>562727</XREF>
768 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.2 | 562718 | Cause | Incorrect parameters | 4.45.16.1.2.6 | 562724 | Measure | Encapsulation of brightness value calculation | Encapsulation and separate software tests for brightness value calculation.
769 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.3 | 562725 | Cause | Reduced image quality | 4.45.16.1.3.1 | 562726 | Measure | User Manual - System calibration | <XREF>562719</XREF>
770 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.3 | 562725 | Cause | Reduced image quality | 4.45.16.1.3.2 | 562727 | Measure | User Manual - Reduced Image Quality due to Filtering | CAUTIONCause: Processing images by changing the image display. Â Possible consequences: Loss of image quality by artifacts. Â Â Remedy:Â Use care when changing the image display settings.
771 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.1 | 562729 | Measure | Reference to measures to generator self-tests during start-up / initialisation and filtration test | See HM_S1_sw_F5G2U1M1 (RAM test at start-up), HM_S1_sw_F5G2U1M2 (integrity checkÂ program code),HM_S1_sw_F5G2U1M3 (integrity check of EEPROM at start-up), HM_S1_sw_F5G2U2M1 (self-test of measuring circuit kV / mA at initialisation), HM_S1_hw_F5G2U2M2 (kV min / max monitoring), HM_S1_manu_F5G2U3M1 (HVL test) in <XREF>578770</XREF>.
772 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.2 | 578746 | Measure | Ready for radiation | If a relevant component is not ready for radiation, the system shall block any radiation release request. Â Â Note:Relevant components are imaging system and X-Ray generator.
773 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.3 | 562731 | Measure | Reference to measure to ensure changing application is possible by operator | <XREF>562721</XREF>.
774 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.4 | 562732 | Measure | Reference - User Manual - Wrong exposure parameters | <XREF>562722</XREF>
775 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.5 | 562733 | Measure | Reference - User Manual - Reduced Image Quality due to Filtering | <XREF>562727</XREF>
776 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.16 | 562713 | Function | Dose control | 4.45.16.1 | 562714 | Hazard | Unnecessary radiation, dose not as desired | 4.45.16.1.4 | 562728 | Cause | Incorrect parameters due to error in generator or tube | 4.45.16.1.4.6 | 562734 | Measure | Reference - Encapsulation of brightness value calculation | <XREF>562724</XREF>
777 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.17 | 562735 | Function | Grid selection | 4.45.17.1 | 562736 | Hazard | Unnecessary radiation | 4.45.17.1.1 | 562737 | Cause | Operating error | 4.45.17.1.1.1 | 562738 | Measure | User Manual - Incorrect Grid Status | CAUTIONCause: Incorrect grid status: inserted, removed. Possible consequences: X-ray exposure without clinical benefit. Remedy:- Select the application according to your clinical procedure. - Make sure that the current grid status meets your image requirements. The imaging system indicates whether a grid has been inserted or removed. - If necessary, insert or remove the grid manually.
778 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.17 | 562735 | Function | Grid selection | 4.45.17.2 | 562739 | Hazard | Crushing between the housings of the detector case. | 4.45.17.2.1 | 562740 | Cause | Changing the grid | 4.45.17.2.1.1 | 562741 | Measure | User Manual - Changing the grid | CAUTIONCause: Changing the grid. Possible consequences: Crushing hazard between theÂ detector housing parts. Remedy: Note the hazardous zones/locations between the detector housing parts.
779 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.18 | 562742 | Function | Store fluoro | 4.45.18.1 | 562743 | Hazard | Loss of images | 4.45.18.1.1 | 562744 | Cause | Insufficient image storage space | 4.45.18.1.1.1 | 562745 | Measure | Reference - Display storage space | <XREF>562613</XREF>
780 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.19 | 562746 | Function | Position FOV | 4.45.19.1 | 562747 | Hazard | Dose not as desired | 4.45.19.1.1 | 562748 | Cause | Incorrect collimator setting | 4.45.19.1.1.1 | 562749 | Measure | Adjustment of the collimator plate | Adjustment of the collimator plate is designed so that the edge is visible in the fluoro image.
781 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.19 | 562746 | Function | Position FOV | 4.45.19.1 | 562747 | Hazard | Dose not as desired | 4.45.19.1.1 | 562748 | Cause | Incorrect collimator setting | 4.45.19.1.1.2 | 562750 | Measure | User Manual - Collimator Check | CAUTIONCause: Wrong collimator setting. Possible Consequences: Dose not as required!Remedy: The collimators need to be set so that at least one edge of the leaves is visible in the fluoroscopy image. This must be checked daily.
782 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.19 | 562746 | Function | Position FOV | 4.45.19.1 | 562747 | Hazard | Dose not as desired | 4.45.19.1.1 | 562748 | Cause | Incorrect collimator setting | 4.45.19.1.1.3 | 562751 | Measure | Adjustment of the collimator plate | Regular check of the collimator plate adjustment per the Maintenance Plan by prescribed check.
783 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.19 | 562746 | Function | Position FOV | 4.45.19.1 | 562747 | Hazard | Dose not as desired | 4.45.19.1.1 | 562748 | Cause | Incorrect collimator setting | 4.45.19.1.1.4 | 562752 | Measure | Reference - Perform calibration of collimator leaves | <XREF>562696</XREF>
784 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.19 | 562746 | Function | Position FOV | 4.45.19.2 | 562753 | Hazard | Eye injury (visual disturbance) | 4.45.19.2.1 | 562754 | Cause | Unintended look into the laser beam for a lengthy time | 4.45.19.2.1.1 | 562755 | Measure | Usage of laser class 1 | <XREF>562338</XREF>
785 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.20 | 562756 | Function | Zoom format change | 4.45.20.1 | 562757 | Hazard | Dose not as desired | 4.45.20.1.1 | 562758 | Cause | Error in FD zoom change | 4.45.20.1.1.1 | 562759 | Measure | Adjustment of the collimator plate | <XREF>562749</XREF>
786 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.21 | 562760 | Function | Exchange grid | 4.45.21.1 | 562761 | Hazard | Dose not as desired | 4.45.21.1.1 | 562762 | Cause | Incorrect grid status, inserted, removed | 4.45.21.1.1.1 | 562763 | Measure | Detect grid status | It has to be detectedÂ at theÂ system if grid is inserted or not.
787 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.21 | 562760 | Function | Exchange grid | 4.45.21.1 | 562761 | Hazard | Dose not as desired | 4.45.21.1.1 | 562762 | Cause | Incorrect grid status, inserted, removed | 4.45.21.1.1.2 | 562764 | Measure | User Manual - Incorrect Grid Status | <XREF>562738</XREF>
788 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.1 | 562766 | Hazard | Infection: A patient is infected by virus/bacteria due to bad disinfection of the system | 4.45.22.1.1 | 562767 | Cause | Insufficient cleaning of system between patients | 4.45.22.1.1.1 | 562768 | Measure | Reducing openings and gaps | Covers are designed that openings and gaps are reduced.
789 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.1 | 562766 | Hazard | Infection: A patient is infected by virus/bacteria due to bad disinfection of the system | 4.45.22.1.1 | 562767 | Cause | Insufficient cleaning of system between patients | 4.45.22.1.1.2 | 562769 | Measure | User Manual - Cleaning of Wheels | <XREF>563241</XREF>
790 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.2 | 562770 | Hazard | Delayed or no treatment | 4.45.22.2.1 | 562771 | Cause | Blood or other liquids damage system | 4.45.22.2.1.1 | 562772 | Measure | Reference - Spillage of Liquids on System | <XREF>562238</XREF>
791 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.2 | 562770 | Hazard | Delayed or no treatment | 4.45.22.2.2 | 562773 | Cause | Cleaning and disinfection substances damage system | 4.45.22.2.2.1 | 562774 | Measure | User Manual - Insufficient cleaning | <XREF>562240</XREF>
792 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.3 | 562775 | Hazard | Undesired radiation | 4.45.22.3.1 | 562776 | Cause | Cleaning person accidentally steps on footswitch | 4.45.22.3.1.1 | 562777 | Measure | User Manual - System switch-off | CAUTIONCause: Inadvertent actuation of the foot pedal. Possible consequences: Risk of radiationRemedy: Switch off system before cleaning.
793 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.22 | 562765 | Function | Release patient | 4.45.22.3 | 562775 | Hazard | Undesired radiation | 4.45.22.3.1 | 562776 | Cause | Cleaning person accidentally steps on footswitch | 4.45.22.3.1.2 | 562778 | Measure | EmergencyStop | see <XREF>562345</XREF>
794 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.23 | 562779 | Function | Select patient from patient list | 4.45.23.1 | 562780 | Hazard | No hazard on system level | 4.45.23.1.1 | 562781 | Cause | None | 4.45.23.1.1.1 | 562782 | Measure | None
795 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.24 | 562783 | Function | Image display | 4.45.24.1 | 562784 | Hazard | Incorrect treatment basis | 4.45.24.1.1 | 562785 | Cause | Confusing of, mix-up of images/series from different patients. | 4.45.24.1.1.1 | 562786 | Measure | Always Display Patient Data | The system shall show patient name during examination. Â Hint:Please note, this measure neither includes exported images nor print-outs (hardcopy) as hazard and cause is related to time period of examination in accordance to intended use (IFU). Anonymizing of image data sets is not in scope of this hazard, either.
796 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.24 | 562783 | Function | Image display | 4.45.24.1 | 562784 | Hazard | Incorrect treatment basis | 4.45.24.1.1 | 562785 | Cause | Confusing of, mix-up of images/series from different patients. | 4.45.24.1.1.2 | 562787 | Measure | Correct image assignment | FLC assigns acquired 2D images to the patient who was registered at the time the images have been acquired.
797 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.24 | 562783 | Function | Image display | 4.45.24.1 | 562784 | Hazard | Incorrect treatment basis | 4.45.24.1.2 | 562788 | Cause | Images are sorted or numbered incorrectly | 4.45.24.1.2.1 | 562789 | Measure | Image number storage in header | Storage of the image number in the image header of the 2D images.
798 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.25 | 562790 | Function | Close Patient | 4.45.25.1 | 562791 | Hazard | No hazard on system level | 4.45.25.1.1 | 562792 | Cause | None | 4.45.25.1.1.1 | 562793 | Measure | None
799 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.1 | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.1.1 | 562796 | Cause | Erroneous release of stored energy (ESU) | 4.45.26.1.1.1 | 562797 | Measure | Electrical separation of accessible parts not protected by MOOP | The system shall safely separate theÂ trolley-chassis cable pins (X10) and the X-ray Generation subsystemÂ as soon as the trolley and the main unit are disconnected. Â Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure.
800 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.1 | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.1.1 | 562796 | Cause | Erroneous release of stored energy (ESU) | 4.45.26.1.1.2 | 562798 | Measure | Check of electrical separation functionality during manufacturing - generator | The functionality of the measure implemented with <XREF>562797</XREF> shall be tested with 100%Â coverage during manufacturing. Â Â Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure.
801 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.1 | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.1.1 | 562796 | Cause | Erroneous release of stored energy (ESU) | 4.45.26.1.1.3 | 562799 | Measure | Check of electrical separation functionality during regular maintenance | Service Documentation:WARNINGCause: Undetected error in electrical separation of X10 pins. Possible consequences: Electrical shock when trolley-chassis cable is disconnected and X10 pins are touched. Remedy: Perform a check of the contactors according to description during maintenance. Â Note for ELKO systems: Measure valid for forward production ELKO systems as of AT_417077, see 06H573400, regulatory electrical safety requirements are fulfilled already without measure.
802 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.2 | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.2.1 | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 4.45.26.2.1.1 | 573395 | Measure | Reference - Electrical separation of accessible parts not protected by MOOP | <XREF>562797</XREF>
803 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.2 | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.2.1 | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 4.45.26.2.1.2 | 573396 | Measure | Reference - Check of electrical separation functionality during manufacturing - generator | <XREF>562798</XREF>
804 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.26 | 562794 | Function | Disconnect Trolley and Main Unit | 4.45.26.2 | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 4.45.26.2.1 | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 4.45.26.2.1.3 | 573397 | Measure | Reference - Check of electrical separation functionality during regular maintenance | <XREF>562799</XREF>
805 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.27 | 562800 | Function | Quick Clean | 4.45.27.1 | 562801 | Hazard | Contamination through improper cleaning | 4.45.27.1.1 | 562802 | Cause | Usage of improper cleaning agents | 4.45.27.1.1.1 | 562803 | Measure | User Manual - Improper Cleaning Agents | CAUTIONCause: Use of unsuitable cleaning agents. Possible consequences: Risk of contamination. Remedy: Only use the recommended cleaning agents and disinfectants.
806 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.28 | 562804 | Function | Retrieve and modify patient | 4.45.28.1 | 562805 | Hazard | No hazard on system level | 4.45.28.1.1 | 562806 | Cause | None | 4.45.28.1.1.1 | 562807 | Measure | None
807 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.1 | 562809 | Hazard | Loss of image data | 4.45.29.1.1 | 562810 | Cause | Unintentional deletion of images/series | 4.45.29.1.1.1 | 562811 | Measure | Reference - Prompt by local deleting of data | <XREF>562324</XREF>
808 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.1 | 562809 | Hazard | Loss of image data | 4.45.29.1.1 | 562810 | Cause | Unintentional deletion of images/series | 4.45.29.1.1.2 | 562812 | Measure | Security dialog for any deletion by the user | <XREF>562818</XREF>
809 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.1 | 562809 | Hazard | Loss of image data | 4.45.29.1.2 | 562813 | Cause | Destruction of the patient directory and loss of patient data and images | 4.45.29.1.2.1 | 562814 | Measure | File system of imaging system | Usage of a standard file system (NTFS).
810 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.1 | 562809 | Hazard | Loss of image data | 4.45.29.1.2 | 562813 | Cause | Destruction of the patient directory and loss of patient data and images | 4.45.29.1.2.2 | 569130 | Measure | Imaging system recovery | Availability of a recovery procedure to reconstruct the patient directory via the study instance UID which is unique.
811 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.2 | 562815 | Hazard | Delayed or no interventional decision | 4.45.29.2.1 | 562816 | Cause | Unintended local deleting of patient data | 4.45.29.2.1.1 | 562817 | Measure | Reference - Prompt by local deleting of data | <XREF>562324</XREF>
812 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.2 | 562815 | Hazard | Delayed or no interventional decision | 4.45.29.2.1 | 562816 | Cause | Unintended local deleting of patient data | 4.45.29.2.1.2 | 562818 | Measure | Security dialog for any deletion by the user | To reduce probability of unintentional loss of data, the deletion of images/series is performed by a security dialog. If non archived images/series are selected for deletion the user has to confirm this in addition.
813 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.29 | 562808 | Function | Image delete | 4.45.29.2 | 562815 | Hazard | Delayed or no interventional decision | 4.45.29.2.1 | 562816 | Cause | Unintended local deleting of patient data | 4.45.29.2.1.3 | 562819 | Measure | Storage commitment archive flag | Storage commitment: The archive flag ('A') is set only after a storage commitment has been received. Note: This is valid in case the DICOM node is capable of sending a storage commitment. The flag is set on study and image level .
814 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.30 | 562820 | Function | Calibrate detector | 4.45.30.1 | 562821 | Hazard | Image artifacts | 4.45.30.1.1 | 562822 | Cause | Bad calibration | 4.45.30.1.1.1 | 562823 | Measure | Automatic procedure to perform calibration correctly | <XREF>562177</XREF>
815 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.30 | 562820 | Function | Calibrate detector | 4.45.30.1 | 562821 | Hazard | Image artifacts | 4.45.30.1.1 | 562822 | Cause | Bad calibration | 4.45.30.1.1.2 | 562824 | Measure | Reference - User Manual - Calibration Procedure | <XREF>562179</XREF>
816 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.30 | 562820 | Function | Calibrate detector | 4.45.30.1 | 562821 | Hazard | Image artifacts | 4.45.30.1.1 | 562822 | Cause | Bad calibration | 4.45.30.1.1.3 | 562825 | Measure | Reference - User Manual - System calibration | <XREF>562719</XREF>
817 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.30 | 562820 | Function | Calibrate detector | 4.45.30.2 | 562826 | Hazard | Unwanted radiation exposure | 4.45.30.2.1 | 562827 | Cause | Patient or other person in exam room during calibration | 4.45.30.2.1.1 | 562828 | Measure | Reference - Service training | <XREF>562186</XREF>
818 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.31 | 562829 | Function | Stop button | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.1 | 562832 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
819 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.32 | 562833 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.2 | 562834 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
820 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.33 | 562835 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.3 | 562836 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
821 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.34 | 562837 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.4 | 562838 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
822 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.35 | 562839 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.5 | 562840 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
823 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.36 | 562841 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.6 | 562842 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
824 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.37 | 562843 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough |  | 562844 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
825 | 4.45 | 562334 | Chapter | Perform Clinical Procedure |  | 562845 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough |  | 562846 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
826 | 4.45 | 562334 | Chapter | Perform Clinical Procedure |  | 562847 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough |  | 562848 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
827 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.40 | 562849 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.10 | 562850 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
828 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.41 | 562851 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.11 | 562852 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
829 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.42 | 562853 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.12 | 562854 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
830 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.43 | 562855 | Function | Native language support | 4.45.31.1 | 562830 | Hazard | Unit stops too late in a critical situation | 4.45.31.1.1 | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 4.45.31.1.1.13 | 562856 | Measure | User Manual - Emergency Stop Location | <XREF>562284</XREF>
831 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.43 | 562855 | Function | Native language support | 4.45.36.1 | 562852 | Hazard | Delayed or no interventional decision | 4.45.36.1.1 | 562853 | Cause | Network disturbance | 4.45.36.1.1.1 | 562854 | Measure | Reference - Use of reliable data transfer standards | <XREF>562323</XREF>
832 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.43 | 562855 | Function | Native language support | 4.45.36.1 | 562852 | Hazard | Delayed or no interventional decision | 4.45.36.1.1 | 562853 | Cause | Network disturbance | 4.45.36.1.1.2 | 562855 | Measure | Reference - Prompt by local deleting of data | <XREF>562324</XREF>
833 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.37 | 562856 | Function | Virus Protection | 4.45.37.1 | 562857 | Hazard | Delayed or no interventional decision | 4.45.37.1.1 | 562858 | Cause | System gehts infected by malware | 4.45.37.1.1.1 | 562865 | Measure | Intrusion detection solution | An intrusion detection solution based on whitelisting is provided by the system.
834 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.37 | 562856 | Function | Virus Protection | 4.45.37.1 | 562857 | Hazard | Delayed or no interventional decision | 4.45.37.1.1 | 562858 | Cause | System gehts infected by malware | 4.45.37.1.1.2 | 562866 | Measure | Reference - Provide SVPs regularly | SVPs are regularly provided according to SOP 409 <XREF>578747</XREF>.
835 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.38 | 562867 | Function | Help functionality | 4.45.38.1 | 562868 | Hazard | No hazard on system level | 4.45.38.1.1 | 562869 | Cause | None | 4.45.38.1.1.1 | 562870 | Measure | None
836 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.39 | 562871 | Function | Measure Area Dose Product | 4.45.39.1 | 562872 | Hazard | Unnecessary dose | 4.45.39.1.1 | 562873 | Cause | Wrong interpretation of applied dose, dose reporting affected | 4.45.39.1.1.1 | 562874 | Measure | Maintenance of the dose area product meter | Request for maintenance of the dose area product meter on a regular base.
837 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.40 | 562875 | Function | IHE Profiles | 4.45.40.1 | 562876 | Hazard | No hazard on system level | 4.45.40.1.1 | 562877 | Cause | None | 4.45.40.1.1.1 | 562878 | Measure | None
838 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.41 | 562879 | Function | Data Model | 4.45.41.1 | 562880 | Hazard | No hazard on system level | 4.45.41.1.1 | 562881 | Cause | None | 4.45.41.1.1.1 | 562882 | Measure | None
839 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.42 | 562883 | Function | Worklist | 4.45.42.1 | 562884 | Hazard | Patient data mixed up | 4.45.42.1.1 | 562885 | Cause | Mixing of data caused by RIS data transfer | 4.45.42.1.1.1 | 562886 | Measure | Reference - Imaging system data consistency | <XREF>569127</XREF>
840 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.42 | 562883 | Function | Worklist | 4.45.42.1 | 562884 | Hazard | Patient data mixed up | 4.45.42.1.1 | 562885 | Cause | Mixing of data caused by RIS data transfer | 4.45.42.1.1.2 | 569129 | Measure | Check worklist data | Plausibility and consistency check of worklist data. Worklistdata are marked with a flag.
841 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.43 | 562887 | Function | Storage | 4.45.43.1 | 562888 | Hazard | No hazard on system level | 4.45.43.1.1 | 562889 | Cause | None | 4.45.43.1.1.1 | 562890 | Measure | None
842 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.44 | 562891 | Function | Archive to offline media | 4.45.44.1 | 562892 | Hazard | No hazard on system level | 4.45.44.1.1 | 562893 | Cause | None | 4.45.44.1.1.1 | 562894 | Measure | None
843 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.45 | 562895 | Function | Query/retrieve | 4.45.45.1 | 562896 | Hazard | No hazard on system level | 4.45.45.1.1 | 562897 | Cause | None | 4.45.45.1.1.1 | 562898 | Measure | None
844 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.46 | 562899 | Function | MPPS | 4.45.46.1 | 562900 | Hazard | No hazard on system level | 4.45.46.1.1 | 562901 | Cause | None | 4.45.46.1.1.1 | 562902 | Measure | None
845 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.47 | 562903 | Function | DICOM dose reporting | 4.45.47.1 | 562904 | Hazard | Wrong dose reported by MPPS | 4.45.47.1.1 | 562905 | Cause | Misleading Information | 4.45.47.1.1.1 | 562906 | Measure | No measure required | Note:This has no impact on interventional decision, but inappropriate actions may be taken based on the dose reading, impact on subsequent radiation events (underreporting. . . ).
846 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.48 | 562907 | Function | Standard display function | 4.45.48.1 | 562908 | Hazard | No hazard on system level | 4.45.48.1.1 | 562909 | Cause | None | 4.45.48.1.1.1 | 562910 | Measure | None
847 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.49 | 562911 | Function | Verification | 4.45.49.1 | 562912 | Hazard | No hazard on system level | 4.45.49.1.1 | 562913 | Cause | None | 4.45.49.1.1.1 | 562914 | Measure | None
848 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.50 | 562915 | Function | Help functionality | 4.45.50.1 | 562916 | Hazard | No hazard on system level | 4.45.50.1.1 | 562917 | Cause | None | 4.45.50.1.1.1 | 562918 | Measure | None
849 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.51 | 562919 | Function | General requirements | 4.45.51.1 | 562920 | Hazard | No hazard on system level | 4.45.51.1.1 | 562921 | Cause | None | 4.45.51.1.1.1 | 562922 | Measure | None
850 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.1 | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 4.45.52.1.1 | 562925 | Cause | Loss of 2D images throughout the 3D scan | 4.45.52.1.1.1 | 562926 | Measure | Save all 2D images data set | Save for all 2D images the corresponding data set of acquisition angles
851 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.1 | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 4.45.52.1.1 | 562925 | Cause | Loss of 2D images throughout the 3D scan | 4.45.52.1.1.2 | 562927 | Measure | User Manual - Reduced accuracy of 3D reconstruction | CAUTIONCause:Â Insufficient geometric accuracy of 3D reconstruction due to an incomplete and/or angulated 3D scan and/or 3D scan started away from 0-positions. Possible consequences: Wrong interventional decision or treatment!Remedy:- Make sure that a complete 3D scan of the entireÂ orbital range is performed. Note, for partial scans it is only required to ensure that the scan is performed within the selected start- and endposition. - Make sure that the C-arm is in 0-position for the angulation axis.
852 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.1 | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 4.45.52.1.2 | 562928 | Cause | C-arm cable influences c-arm geometry | 4.45.52.1.2.1 | 562929 | Measure | No external C-arm cable | To reduce mechanical and movement influence for a 3D scan there shall be noÂ external C-arm cable.
853 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.2 | 562930 | Hazard | Patient risk for biological hazard | 4.45.52.2.1 | 562931 | Cause | C-arm cable moves into sterile area | 4.45.52.2.1.1 | 562932 | Measure | Reference - C-arm cable avoids sterile area for all positions | <XREF>562929</XREF>
854 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.3 | 562933 | Hazard | Risk of injury during motorized orbital movement | 4.45.52.3.1 | 562934 | Cause | Collision of C-arm with people and goods | 4.45.52.3.1.1 | 562935 | Measure | Reference to workflow | CAUTIONCause: Change of c-arm position after manual test drivePossible consequences: Risk of crushing!Remedy: After the manual test drive the position of the C-arm must not be changed.
855 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.3 | 562933 | Hazard | Risk of injury during motorized orbital movement | 4.45.52.3.1 | 562934 | Cause | Collision of C-arm with people and goods | 4.45.52.3.1.2 | 562936 | Measure | Motorized operation only after manual test drive | Release for motorized 3D scan after manual test drive
856 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.3 | 562933 | Hazard | Risk of injury during motorized orbital movement | 4.45.52.3.1 | 562934 | Cause | Collision of C-arm with people and goods | 4.45.52.3.1.3 | 562937 | Measure | Reference - Continuous activation | <XREF>562377</XREF>
857 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.3 | 562933 | Hazard | Risk of injury during motorized orbital movement | 4.45.52.3.1 | 562934 | Cause | Collision of C-arm with people and goods | 4.45.52.3.1.4 | 562938 | Measure | Reference - Emergency stop at every point of motorized movement activation | <XREF>562345</XREF>
858 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.3 | 562933 | Hazard | Risk of injury during motorized orbital movement | 4.45.52.3.1 | 562934 | Cause | Collision of C-arm with people and goods | 4.45.52.3.1.5 | 562939 | Measure | Reference to set max. speed 3D scan | See <XREF>563486</XREF>.
859 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.4 | 562940 | Hazard | Patient cannot be salvaged | 4.45.52.4.1 | 562941 | Cause | Orbital drive is failing | 4.45.52.4.1.1 | 562942 | Measure | C-arm can be moved manually | <XREF>562449</XREF>
860 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.5 | 562943 | Hazard | Treatment on the wrong side | 4.45.52.5.1 | 562944 | Cause | Image orientation confused | 4.45.52.5.1.1 | 562945 | Measure | Reference to User Manual - Correct Position of Patient | <XREF>562959</XREF>
861 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.52 | 562923 | Function | 3D Scan and acquisistion | 4.45.52.6 | 562946 | Hazard | 3D scan is unavailable | 4.45.52.6.1 | 562947 | Cause | Orbital drive fails or other error | 4.45.52.6.1.1 | 562948 | Measure | No measure required
862 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.1 | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 4.45.53.1.1 | 562951 | Cause | Artifacts that might be mistaken for anatomical structures | 4.45.53.1.1.1 | 562952 | Measure | Artifacts Image Quality test for MPR | Service DocumentsCAUTIONCause: Artifacts in Multi Planar Reconstruction (MPR) 3D imaging, image quality testPossible Consequences: Artifacts could lead to wrong decisions or treatments. Remedy:Â Perform 3D relevant calibration (Tune UP/Projection Geometry + IZeroCalibration)and IQ test during the regular maintenance interval.
863 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.1 | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 4.45.53.1.1 | 562951 | Cause | Artifacts that might be mistaken for anatomical structures | 4.45.53.1.1.2 | 562953 | Measure | Artifacts Maintenance recommendation in the user manual | CAUTIONCause: Missing regular maintenancePossible consequences: Insufficient image quality!Remedy: It is recommended to maintain the system regularly to achieve the best possible image quality.
864 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.1 | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 4.45.53.1.2 | 562954 | Cause | Alignment error | 4.45.53.1.2.1 | 562955 | Measure | Alignment error Image Quality test for MPR | Image Quality test in service workflow
865 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.1 | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 4.45.53.1.2 | 562954 | Cause | Alignment error | 4.45.53.1.2.2 | 562956 | Measure | Reference - Artifacts Maintenance recommendation in the user manual | <XREF>562953</XREF>
866 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.2 | 562957 | Hazard | Wrong orientation of the reconstructed 3D volume | 4.45.53.2.1 | 562958 | Cause | Wrong input of patient orientation | 4.45.53.2.1.1 | 562959 | Measure | User Manual - Verify Patient Position and Orientation | CAUTIONCause: Wrong input of patient orientation. Possible consequences: Risk of injury to the patient!Remedy: Please verify the correct patient to system relation and patient orientation.
867 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.2 | 562957 | Hazard | Wrong orientation of the reconstructed 3D volume | 4.45.53.2.1 | 562958 | Cause | Wrong input of patient orientation | 4.45.53.2.1.2 | 562961 | Measure | Orientation of Image and P-matrices | Encapsulation and separate software tests for volume orientation calculation according to patient position. In the MPR and VRT view, images shall be displayed with an orientation indicator according to the 3D volume information.
868 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.3 | 562962 | Hazard | 3D Imaging incorrect | 4.45.53.3.1 | 562963 | Cause | Acquired set of 2D image assigned to wrong P-matrix | 4.45.53.3.1.1 | 562964 | Measure | Assure correct assignment | Assure an correct assignment of the 2D images to the P-Matrix throughout the imaging application.
869 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.53 | 562949 | Function | 3D Imaging | 4.45.53.3 | 562962 | Hazard | 3D Imaging incorrect | 4.45.53.3.2 | 562965 | Cause | Wrong assignment of P-matrices to scan protocol | 4.45.53.3.2.1 | 562966 | Measure | Correct assignment of scan config | Assure correct assignment of scan config to the correct P-Matrix
870 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.1 | 562968 | Hazard | Functional Hazards 3D Imaging insufficient | 4.45.54.1.1 | 562969 | Cause | Artifact level too high | 4.45.54.1.1.1 | 562970 | Measure | Reference - Artifacts Maintenance recommendation in the user manual | <XREF>562953</XREF>
871 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.1 | 562972 | Cause | Patient moves | 4.45.54.2.1.1 | 562973 | Measure | User Manual - patient movement during 3D | CAUTIONCause: Patient movement during the 3D scanPossible consequences: Insufficient image quality!Remedy: Patient movement has to be avoided during the 3D scan.
872 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.2 | 562974 | Cause | System/C-armswings | 4.45.54.2.2.1 | 562975 | Measure | Ramp function | Drive starts with ramp function
873 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.3 | 562976 | Cause | System shifts during scan | 4.45.54.2.3.1 | 562977 | Measure | User Manual - System shift during 3D | CAUTIONCause: System shifts during the 3D scanPossible consequences: Impaired 3D image acquisition!Remedy: All brakes need to be closed and the system shall not be touched by user to avoid that the system shifts.
874 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.4 | 562978 | Cause | C-arm angulation does not correspond to calibration area | 4.45.54.2.4.1 | 562979 | Measure | Retrieval of C-arm angulation | Start of 3D scan only possible if C-arm angulationÂ is inside of the defined allowed 3d scan angulation area (e.g. +/- 15Â°).
875 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.5 | 562980 | Cause | Image quality for image reconstuction impaired | 4.45.54.2.5.1 | 562981 | Measure | System calibration via calibration phantom | System calibration via calibration phantom. The system checks the date since the last calibration. After 24 month without recalibration the user receives a warning notice telling him to perform a recalibration.
876 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.5 | 562980 | Cause | Image quality for image reconstuction impaired | 4.45.54.2.5.2 | 562982 | Measure | Inspection of image quality | Service manual:Inspection of image quality during maintenace intervall
877 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.6 | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 4.45.54.2.6.1 | 562984 | Measure | Statement about fidelity in user manual | CAUTIONCause: Limited geometric fidelity of 3D reconstructionPossible consequences: Reduced 3D image quality or insufficient or incorrect basis for measurements, diagnostic or therapeutic decisionsRemedy: Note that the geometric fidelity of 3D reconstructions can only be maintained within a defined volume. Perform length and angle measurements in 3D images only within this volume and verify the results before diagnostic or therapeutic decisions are taken.
878 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.6 | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 4.45.54.2.6.2 | 562985 | Measure | Inspection of fidelity | Service DocumentsCAUTIONCause: Reduced accuracy in 3D for functions Measure and AnnotatePossible Consequences: Incorrect interventional decisions or treatment. Remedy:Â Inspect the geometric accuracy (in IQAP test) during regular maintenance interval.
879 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.6 | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 4.45.54.2.6.3 | 562986 | Measure | Encapsulation 3D image processing scale | Encapsulation and separate software tests for the handling of the model matrix.
880 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.7 | 562987 | Cause | Test procedure for inspection of IQ insufficient | 4.45.54.2.7.1 | 562988 | Measure | Application of an appropriate test procedure | Service DocumentsCAUTIONCause: 2D image quality tests are not performed prior to 3D image quality tests. Possible Consequences: Impaired image quality could lead to wrong decisions or treatments and / or additional doseRemedy:Â Perform 2DÂ image qualityÂ tests before performing theÂ 3DÂ image qualityÂ tests (3D IQ test procedures can only be performed, when 2D tests are all succeeded).
881 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.54 | 562967 | Function | 3D reconstruction | 4.45.54.2 | 562971 | Hazard | Impaired image acquisition | 4.45.54.2.7 | 562987 | Cause | Test procedure for inspection of IQ insufficient | 4.45.54.2.7.2 | 562989 | Measure | Encapsulation 3D handling of brightness / contrast related data | Encapsulation and separate software tests of handling the Brightness/Contrast related data (rescale slope, rescale intercept).
882 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.1 | 562991 | Hazard | Additional strain on patient | 4.45.55.1.1 | 562992 | Cause | Disturbance of data transfer | 4.45.55.1.1.1 | 562993 | Measure | Use of a defined protocol | Use of a defined protocol (NaviLink) based on TCP/IP
883 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.2 | 562994 | Hazard | Risk of injury by C-arm | 4.45.55.2.1 | 562995 | Cause | Unintended shifting of the C-arm | 4.45.55.2.1.1 | 562996 | Measure | User Manual - Non-Counterbalanced C-arm during Navigation | CAUTIONCause: Additional weightÂ after installation ofÂ parts for navigationPossible consequences: Risk of injuryRemedy: Note that the C-arm is no longer balanced after additional components are attached. It can move on its own after brakes are released.
884 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.2 | 562994 | Hazard | Risk of injury by C-arm | 4.45.55.2.2 | 562997 | Cause | Incorrectly mounted navigation components impair accuracy | 4.45.55.2.2.1 | 562998 | Measure | User Manual - Ensure Correct Mounting of Navigation System | CAUTIONCause: Incorrectly attached navigation system. Possible consequences: Risk of injury!Remedy: Before navigation-assisted operation, the user must ensure sufficient accuracy of the navigation system.
885 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.2 | 562994 | Hazard | Risk of injury by C-arm | 4.45.55.2.3 | 562999 | Cause | Mechanical connection between detector housing and adapter fails | 4.45.55.2.3.1 | 563000 | Measure | Reference - Design fit for fourfold weight | <XREF>563023</XREF>
886 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.3 | 563001 | Hazard | Risk of injury during operation/transportation | 4.45.55.3.1 | 563002 | Cause | Tipping over of the device | 4.45.55.3.1.1 | 563003 | Measure | Reference - Design of system | <XREF>562067</XREF>
887 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.3 | 563001 | Hazard | Risk of injury during operation/transportation | 4.45.55.3.1 | 563002 | Cause | Tipping over of the device | 4.45.55.3.1.2 | 563004 | Measure | User Manual - Maximum Weight | WARNINGCause: Heavy weight of external navigation components. Possible consequences: Risk of injury!Remedy: Note that stability in accordance with DIN EN/UL standards is guaranteed only if the external system (navigation) on the detector does not exceed the calculated maximum weight.
888 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.4 | 563005 | Hazard | Injury due to surface condition | 4.45.55.4.1 | 563006 | Cause | Sharp edges | 4.45.55.4.1.1 | 563007 | Measure | Reference - Compliance with IEC 60601-1 external | <XREF>563042</XREF>
889 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.5 | 563008 | Hazard | Risk of injury due to breakage | 4.45.55.5.1 | 563009 | Cause | Overloading of C-arm bearing and supporting elements for the detector | 4.45.55.5.1.1 | 563010 | Measure | Reference - Design of the movement unit | <XREF>562415</XREF>
890 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.5 | 563008 | Hazard | Risk of injury due to breakage | 4.45.55.5.1 | 563009 | Cause | Overloading of C-arm bearing and supporting elements for the detector | 4.45.55.5.1.2 | 563011 | Measure | User Manual - Break Resistance | CAUTIONCause: Overloading of C-arm. Possible consequences: Risk of injury!Remedy: Note that the braking action is guaranteed only if the external system (navigation) on the detector does not exceed the calculated maximum weight.
891 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.6 | 563012 | Hazard | Scattered radiation | 4.45.55.6.1 | 563013 | Cause | Unnecessary radiation exposure when additional objects in beam path. | 4.45.55.6.1.1 | 563014 | Measure | User Manual - Scattered Radiation | CAUTIONCause: Additional objects in beam path. Possible consequences: Risk of unnecessary radiation exposure!Remedy: Note that additional objects in the beam path increase scatter radiation.
892 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.7 | 563015 | Hazard | C-arm collides with people or goods | 4.45.55.7.1 | 563016 | Cause | Additional weight | 4.45.55.7.1.1 | 563017 | Measure | User Manual - Mounting of Additional Parts | CAUTIONCause: Changing the weight distributionPossible consequences: Risk of crushing!Remedy: Adding weight to the detector or single tank side means a loss in counterweight and can result in unintended movement of the C-arm.
893 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.8 | 563019 | Hazard | Markers of different nav. Systems can fall off | 4.45.55.8.1 | 563020 | Cause | Insufficient dimensioning of the mechanical interface | 4.45.55.8.1.1 | 563021 | Measure | Navigation Interface - Marker Frame Fixation | The dimensioning of the mechanical interface at the marker frame must be sufficient to enable a tight and secure fixation of the marker frame at the FD.
894 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.8 | 563019 | Hazard | Markers of different nav. Systems can fall off | 4.45.55.8.2 | 563022 | Cause | Reference ring loosenes | 4.45.55.8.2.1 | 563023 | Measure | Design fit for fourfold weight | Design fit for fourfold stability; according to IEC 60601-1-1:2005, clause 9. 8. 2 (tensile safety factor).
895 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.8 | 563019 | Hazard | Markers of different nav. Systems can fall off | 4.45.55.8.3 | 563024 | Cause | Markers insufficiently fixed | 4.45.55.8.3.1 | 563025 | Measure | Navigation Interface - Marker Fixation | The markers attached to the marker frame (e.g. spheres) must be fixed sufficiently to prevent markers from falling down.
896 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.9 | 563026 | Hazard | Risk of injury during operation/transportation | 4.45.55.9.1 | 563027 | Cause | Tipping over of the device in general | 4.45.55.9.1.1 | 563028 | Measure | Reference - Design of system | <XREF>562067</XREF>
897 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.9 | 563026 | Hazard | Risk of injury during operation/transportation | 4.45.55.9.1 | 563027 | Cause | Tipping over of the device in general | 4.45.55.9.1.2 | 563029 | Measure | Reference - Navigation Interface - Maximum Weight | <XREF>563734</XREF>
898 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.10 | 563033 | Hazard | Wrong treatment due to insufficient precision | 4.45.55.10.1 | 563034 | Cause | Dirty/Damaged Mechanical Detector Interface | 4.45.55.10.1.1 | 563035 | Measure | User Manual - Visual Inspection/Cleaning | CAUTIONCause: Insufficient navigation accuracy due to contaminated or defective detector interface. Possible consequences: Insufficient navigation accuracy, risk of injury. Remedy: Prior to using the navigation system check the required cleanliness and functionality.
899 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.11 | 563036 | Hazard | Risk of injury due to breakage | 4.45.55.11.1 | 563037 | Cause | Overloading of C-arm bearing and supporting elements for II | 4.45.55.11.1.1 | 563038 | Measure | Reference - Design of the movement unit | <XREF>562415</XREF>
900 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.12 | 563039 | Hazard | Electric schock | 4.45.55.12.1 | 563040 | Cause | Isolation error of nav. parts | 4.45.55.12.1.1 | 563041 | Measure | User Manual - Components Comply With IEC60601-1 | WARNINGCause: Use of insufficiently electrically-protected external navigation components. Possible consequences: Risk of injury!Remedy:Â Use only accessories corresponding to the CE Declaration of Conformity.
901 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.55 | 562990 | Function | 2D/3D Navigation | 4.45.55.12 | 563039 | Hazard | Electric schock | 4.45.55.12.1 | 563040 | Cause | Isolation error of nav. parts | 4.45.55.12.1.2 | 563042 | Measure | Navigation Interface - Comply with IEC60601-1 | The IEC 60601-1:2005 for navigation systems must be adhered to.
902 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.1 | 563044 | Hazard | Collision with catheters | 4.45.56.1.1 | 563045 | Cause | Operation error | 4.45.56.1.1.1 | 563046 | Measure | User Manual - Dry Run Manually | WARNING:Cause: Collision with e.g. catheters or other vital components in ORPossible consequences: Risk of injury!Remedy: Before starting the 3D scan, you must do a manual test drive to prevent collisions.
903 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.2 | 563047 | Hazard | 3D navigation accuracy reduced | 4.45.56.2.1 | 563048 | Cause | Start position 3D scan | 4.45.56.2.1.1 | 563049 | Measure | The accuracy of the mechanical marker interface | The geometry of the mechanical marker interface has to fulfill the tolerance +/-0,02 mm.
904 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.2 | 563047 | Hazard | 3D navigation accuracy reduced | 4.45.56.2.1 | 563048 | Cause | Start position 3D scan | 4.45.56.2.1.2 | 563050 | Measure | User Manual - Verify Accuracy | CAUTIONCause: Absolute position accuracy of the mechanical interface insufficient. Possible consequences: Risk of injuryRemedy: Accuracy has to be verified after 3D navigation system has received the 3D volume data.
905 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.3 | 563051 | Hazard | Wrong treatment due to insufficient precision | 4.45.56.3.1 | 563052 | Cause | Damage during transport | 4.45.56.3.1.1 | 563053 | Measure | Reference - User Manual - Visual Inspection/Cleaning | <XREF>563035</XREF>
906 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.3 | 563051 | Hazard | Wrong treatment due to insufficient precision | 4.45.56.3.2 | 563054 | Cause | Calibration-phantom interface damaged | 4.45.56.3.2.1 | 563055 | Measure | Reference - User Manual - Visual Inspection/Cleaning | <XREF>563035</XREF>
907 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.3 | 563051 | Hazard | Wrong treatment due to insufficient precision | 4.45.56.3.3 | 563056 | Cause | Recalibration for 3D-mode | 4.45.56.3.3.1 | 563057 | Measure | Admin Guide - Navigation system test after recalibration | CAUTIONCause: RecalibrationÂ for 3D modePossible consequences: Risk of injuryRemedy: After recalibrating in 3D mode, a test must be carried out using the navigation system.
908 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.3 | 563051 | Hazard | Wrong treatment due to insufficient precision | 4.45.56.3.3 | 563056 | Cause | Recalibration for 3D-mode | 4.45.56.3.3.2 | 563058 | Measure | Service Documentation - Navigation system test after recalibration | Service DocumentationCAUTIONCause: After a geometry calibration for 3D mode, or after modifications impairing the operation or accuracy of theÂ´navigation system, no recalibration with the relevant navigation system is performed. Possible consequences: Inaccurate navigation results, risk of injury for the patient. Remedy:- After a 3D geometry calibration or any modification impairing the operation or accuracy of the navigation system, a recalibration needs to be performed with the relevant navigation system. - The operator shall be informed verbally and / or in writing, that the accuracy of the navigation system must be checked. - The relevant parts that influence the navigation accuracy are listed in the Replacement of Parts document.
909 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.3 | 563051 | Hazard | Wrong treatment due to insufficient precision | 4.45.56.3.4 | 563059 | Cause | Faulty navigation system during calibration | 4.45.56.3.4.1 | 563060 | Measure | User Manual - Check Precision of Navigation | CAUTIONCause: Faulty navigation system during calibrationPossible consequences: Risk of injuryRemedy: Before performing an operation via navigation, the user must ensure that the navigation accuracy is adequate.
910 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.4 | 563061 | Hazard | Wrong interventional decision or treatment | 4.45.56.4.1 | 563062 | Cause | Wrong selection by user | 4.45.56.4.1.1 | 563063 | Measure | User Manual - Choose Correct Navigation | CAUTIONCause: Wrong selection of navigation system by user. Possible consequences: Risk of injury!Remedy: Before starting a scan, make sure that the correct navigation unit is selected.
911 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.5 | 563064 | Hazard | Additional strain on patient | 4.45.56.5.1 | 563065 | Cause | Incorrect registration of the marker | 4.45.56.5.1.1 | 563066 | Measure | No scan is possible without confirmation | Def. Handshake for registration in interface description: No scan is possible without confirmation from the nav. system
912 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.5 | 563064 | Hazard | Additional strain on patient | 4.45.56.5.1 | 563065 | Cause | Incorrect registration of the marker | 4.45.56.5.1.2 | 563067 | Measure | Navigation Interface - Scan Confirmation | A defined handshake for registration shall be used to ensure, that no scan is possible without confirmation from the navigation system.
913 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.5 | 563064 | Hazard | Additional strain on patient | 4.45.56.5.2 | 563068 | Cause | Rescan due to user error | 4.45.56.5.2.1 | 563069 | Measure | No additional measures required | No additional measures are required. The operators are trained.
914 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.5 | 563064 | Hazard | Additional strain on patient | 4.45.56.5.3 | 563070 | Cause | Rescan due to communication issue | 4.45.56.5.3.1 | 563071 | Measure | No additional measures required | No additional measures are required. The operators are trained. When communication is established again, the operator can rescan.
915 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.6 | 563072 | Hazard | Short-term delay | 4.45.56.6.1 | 563073 | Cause | Communication issues | 4.45.56.6.1.1 | 563074 | Measure | No additional measures required | No additional measures are required. In case the handshake fails, no scan is possible with navigation system. However 3D without navigation is possible.
916 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.6 | 563072 | Hazard | Short-term delay | 4.45.56.6.2 | 563075 | Cause | User error | 4.45.56.6.2.1 | 563076 | Measure | Reference - No scan is possible without confirmation | <XREF>563066</XREF>
917 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.56 | 563043 | Function | 3D Navigation | 4.45.56.6 | 563072 | Hazard | Short-term delay | 4.45.56.6.2 | 563075 | Cause | User error | 4.45.56.6.2.2 | 563077 | Measure | Reference - Navigation Interface - Scan Confirmation | <XREF>563067</XREF>
918 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.57 | 563078 | Function | Application Hosting | 4.45.57.1 | 563079 | Hazard | Influence on performance or functionality of hosted application | 4.45.57.1.1 | 563080 | Cause | Unauthorized update or execution of unauthorized software | 4.45.57.1.1.1 | 563081 | Measure | Intrusion detection solution | An intrusion detection solution based on whitelisting is provided by the system.
919 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.57 | 563078 | Function | Application Hosting | 4.45.57.2 | 563082 | Hazard | Application Hosting PC does not start | 4.45.57.2.1 | 563083 | Cause | Application PC does not start due to HW problem, boot problem | 4.45.57.2.1.1 | 563084 | Measure | None | The risk depends on the application. Application cannot be executed.
920 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.1 | 563086 | Hazard | Unwanted radiation | 4.45.58.1.1 | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 4.45.58.1.1.1 | 563088 | Measure | Label the wireless footswitch with location | In Service DocumentationCAUTIONCause: Radiation released from a footswitch in another roomPossible Consequences: Non-clinical exposure to ionizing radiationRemedy: Before performing any service activities with the wireless footswitch, checkÂ that the footswitchÂ belongs to the system and check that the indicating label is in place and properly visible.
921 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.1 | 563086 | Hazard | Unwanted radiation | 4.45.58.1.1 | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 4.45.58.1.1.2 | 563089 | Measure | User Manual - Label the wireless footswitch with location | CAUTIONCause: Radiation released from a footswitch in another roomRisk: Non-clinical exposure to ionizing radiation. Remedy: As part of the system installation, the wireless footswitchÂ is labeled according to whichÂ system it belongs to. When utilizingÂ theÂ wireless footswitch, check thatÂ it belongs to the systemÂ according to theÂ indicating labels.
922 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.1 | 563086 | Hazard | Unwanted radiation | 4.45.58.1.1 | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 4.45.58.1.1.3 | 563090 | Measure | Wireless footswitch and system are paired | An initial pairing has to be done. This pairing shall remain, even in the case of reboot, reconnection of the wireless footswitch or the system.
923 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.1 | 563086 | Hazard | Unwanted radiation | 4.45.58.1.1 | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 4.45.58.1.1.4 | 563091 | Measure | Label - wireless footswitch with location | A label shall be attached on wireless footswitch to which system he belongs to.
924 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.1 | 563086 | Hazard | Unwanted radiation | 4.45.58.1.2 | 563092 | Cause | Send/Receive modules of wireless footswitch are not reliable | 4.45.58.1.2.1 | 563093 | Measure | Safe hardware modules and signal transmission line | Usage of safe hardware modules and signal transmission line (e. g redundant hardware contacts and/or CRC for transmission line)
925 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.2 | 563094 | Hazard | Unnecessary radiation | 4.45.58.2.1 | 563095 | Cause | x-ray indicators and patient not visible for the user | 4.45.58.2.1.1 | 563096 | Measure | User Manual -Do not release x-ray out of vieweing distance | CAUTIONCause: X-ray release may be possible from remote locations. Risk: Unnecessary dose. Remedy: Do not release X-ray if you are not in viewing distance to the patient.
926 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.3 | 563097 | Hazard | Delayed or no treatment | 4.45.58.3.1 | 563098 | Cause | Interference with other wireless Footswitch | 4.45.58.3.1.1 | 563099 | Measure | User Manual - Interference wireless footswitch | CAUTIONCause: The wireless footswitch may be interfered by other equipment, including portable and fixed RF communication equipment, even if such equipment meets the applicable emissions requirements. Risk: Fluoro or acquisition may sporadically be interrupted. Remedy:- The operator must assure that other wireless device in the 2,4 GHz ISM band should not be operated in the vicinity of appr. 5m around theÂ system. - Please observe and verify normal operation of the wireless footswitch prior to using it.
927 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.3 | 563097 | Hazard | Delayed or no treatment | 4.45.58.3.2 | 563100 | Cause | Battery empty | 4.45.58.3.2.1 | 563101 | Measure | User Manual -Do not release x-ray out of vieweing distance | CAUTIONCause: Weak batteries of wireless footswitch. Risk: Delay in the clinical procedure. Remedy: Make sure the batteries areÂ charged prior to starting an examination.
928 | 4.45 | 562334 | Chapter | Perform Clinical Procedure | 4.45.58 | 563085 | Function | Wireless footswitch | 4.45.58.3 | 563097 | Hazard | Delayed or no treatment | 4.45.58.3.2 | 563100 | Cause | Battery empty | 4.45.58.3.2.2 | 563102 | Measure | High capacity battery | Usage of a battery with the possibility to work with wireless footswitchÂ forÂ 48 hours at least (see table 12 use case 4 for use profile). The user shall be informed via LEDsÂ about low capacity (usage<=24 hours under conditions of table 12 use case 4) to recharge or exchangeÂ the battery.
929 | 4.46 | 563103 | Chapter | Service System | 4.46.1 | 563104 | Function | Remote service | 4.46.1.1 | 563105 | Hazard | Malfunction of the examination routine | 4.46.1.1.1 | 563106 | Cause | Access via remote service | 4.46.1.1.1.1 | 563107 | Measure | Reference - Remote access | <XREF>563179</XREF>
930 | 4.46 | 563103 | Chapter | Service System | 4.46.1 | 563104 | Function | Remote service | 4.46.1.1 | 563105 | Hazard | Malfunction of the examination routine | 4.46.1.1.2 | 563108 | Cause | Install Software on wrong system | 4.46.1.1.2.1 | 563109 | Measure | Check system requirements before installation starts | It is required that the Software Package Source checks the needed system requirements before the installation starts. Therefore an installation on a wrong system is impossible. Â Note, seeÂ 11108300 EKO 000Â Configuration Management Plan
931 | 4.46 | 563103 | Chapter | Service System | 4.46.2 | 563110 | Function | Software modification incl. Bugfix | 4.46.2.1 | 563111 | Hazard | Delayed or no interventional decision | 4.46.2.1.1 | 563112 | Cause | Impermissible or faulty manipulation of SW from network. | 4.46.2.1.1.1 | 563113 | Measure | Reference - User Manual - Impermissible or faulty manipulation from network | <XREF>562154</XREF>
932 | 4.46 | 563103 | Chapter | Service System | 4.46.2 | 563110 | Function | Software modification incl. Bugfix | 4.46.2.1 | 563111 | Hazard | Delayed or no interventional decision | 4.46.2.1.5 | 563120 | Cause | Non approved SW/HW | 4.46.2.1.5.1 | 563121 | Measure | Reference - User Manual - Use of non-approved parts | <XREF>562152</XREF>
933 | 4.46 | 563103 | Chapter | Service System | 4.46.2 | 563110 | Function | Software modification incl. Bugfix | 4.46.2.1 | 563111 | Hazard | Delayed or no interventional decision | 4.46.2.1.5 | 563120 | Cause | Non approved SW/HW | 4.46.2.1.5.2 | 563122 | Measure | Use of unauthorized software | Use of unauthorized software or manipulations and alterations of theÂ software can cause the product to function incorrectly and cause injury to the patient and damage to the product. Use software with correct version approved by Siemens.
934 | 4.46 | 563103 | Chapter | Service System | 4.46.3 | 563123 | Function | Adjustment (e.g. dose rate adjustment) | 4.46.3.1 | 563124 | Hazard | Unwanted radiation released only by SW (closed HW loop in Service Mode, e.g. during calibration) | 4.46.3.1.1 | 563125 | Cause | SW or communication error | 4.46.3.1.1.1 | 563128 | Measure | Open HW-loop after Service | Service DocumentsCAUTIONCause: Hardware loop still active after service (radiation release switch S3) and rebootPossible Consequences: X-ray is prohibited, delay of procedureRemedy:Â Set the radiation release switch S3 to the correct positionÂ after service.
935 | 4.46 | 563103 | Chapter | Service System | 4.46.3 | 563123 | Function | Adjustment (e.g. dose rate adjustment) | 4.46.3.2 | 563129 | Hazard | Increased radiation exposure during service activities | 4.46.3.2.1 | 563130 | Cause | Radiation safety guidelines are not followed | 4.46.3.2.1.1 | 563131 | Measure | Service Manual - Radiation safety | Service DocumentsÂ WARNINGÂ Cause: Non-complianceÂ with radiation safety guidelinesÂ Possible consequences: Increased probability of deterministic and stochasticÂ radiation effects. Â Â Remedy:- Follow the radiation safety guidelines and the local regulations for radiation safety- Increase the distance to the radiation source- Release radiation onlyÂ when required- Use low radiation parameters- Use radiation protection,Â such asÂ protection clothing
936 | 4.46 | 563103 | Chapter | Service System | 4.46.4 | 563132 | Function | Replace electrical parts / Replace mechanical parts | 4.46.4.1 | 563133 | Hazard | Environmental hazard | 4.46.4.1.1 | 563134 | Cause | Improper disposal of hazardous materials | 4.46.4.1.1.1 | 563135 | Measure | Dismantling | <XREF>562246</XREF>
937 | 4.46 | 563103 | Chapter | Service System | 4.46.4 | 563132 | Function | Replace electrical parts / Replace mechanical parts | 4.46.4.3 | 563139 | Hazard | Thermal energy and high temperature | 4.46.4.3.1 | 563140 | Cause | Parts too warm | 4.46.4.3.1.1 | 563141 | Measure | Service Instruction - Hot Surface of TCP backside | Service DocumentsCAUTIONCause: Hot surface of TCP backsidePossible Consequences: Risk of burningsRemedy: While replacing trolley control panel (TCP), please note, surface on backside of TCP maybe hot. Handle with care!
938 | 4.46 | 563103 | Chapter | Service System | 4.46.4 | 563132 | Function | Replace electrical parts / Replace mechanical parts | 4.46.4.4 | 563142 | Hazard | ESD | 4.46.4.4.1 | 563143 | Cause | ESD related damages of electronic components | 4.46.4.4.1.1 | 563144 | Measure | Service Instruction - ESD | Service DocumentationÂ CAUTIONCause: ESDPossible Consequences: Damage of electronic components leading to deterioration in system functionality. Remedy: Please follow the ESD regulations during all work steps.
939 | 4.46 | 563103 | Chapter | Service System | 4.46.4 | 563132 | Function | Replace electrical parts / Replace mechanical parts | 4.46.4.5 | 563145 | Hazard | Spring jumps out when replacing the standard monitor holder on systems with a standard monitor column (non-motorized) | 4.46.4.5.1 | 563146 | Cause | Removing the snap ring without holding the spring | 4.46.4.5.1.1 | 563147 | Measure | Service Instruction - Hold spring of monitor holder | Service DocumentationCAUTIONCause: Removing the snap ring of the monitor holder without holding the spring. Possible Consequences: Spring can jump out, risk of injury. Remedy: While removing the snap ring, hold the spring with the other hand.
940 | 4.46 | 563103 | Chapter | Service System | 4.46.4 | 563132 | Function | Replace electrical parts / Replace mechanical parts | 4.46.4.6 | 563148 | Hazard | Inaccurate navigation reesult | 4.46.4.6.1 | 563149 | Cause | Removing and remounting the FD cover and optional mechanical interface of the navigation system | 4.46.4.6.1.1 | 563150 | Measure | Service Instruction - Inform customer to check navigation accuracy | Service DocumentationCAUTIONCause: Removing and remounting the FD cover and optional mechanical interface of the navigation system. Possible consequences: Inaccurate navigation results, risk of injury for the patient. Remedy:- After remounting the FD cover and the mechanical interface for the navigation system, check the correct mounting (clearance). - Inform the operator verbally and / or in writing, that the accuracy of the navigation system must be checked.
941 | 4.46 | 563103 | Chapter | Service System | 4.46.5 | 563151 | Function | UPS operation | 4.46.5.1 | 563152 | Hazard | Electrical shock | 4.46.5.1.1 | 563153 | Cause | Power is not switched off | 4.46.5.1.1.1 | 563154 | Measure | Service Instruction - Power Presence in System | Service DocumentsWARNINGCause: Power input connector of the imaging system PC is not unplugged or UPS is not switched off during service. Possible Consequences: Risk of death orÂ serious injury due to electric shock when electrified components are touched. Remedy: Prior to beginning service activities, unplug the power input connector in the imaging system PC or switch off the UPS.
942 | 4.46 | 563103 | Chapter | Service System | 4.46.5 | 563151 | Function | UPS operation | 4.46.5.1 | 563152 | Hazard | Electrical shock | 4.46.5.1.1 | 563153 | Cause | Power is not switched off | 4.46.5.1.1.2 | 563155 | Measure | Warning labeling | Warning labelingWarning labeling at the power inlet of the image system. Unplug the power input connector for service.
943 | 4.46 | 563103 | Chapter | Service System | 4.46.5 | 563151 | Function | UPS operation | 4.46.5.1 | 563152 | Hazard | Electrical shock | 4.46.5.1.1 | 563153 | Cause | Power is not switched off | 4.46.5.1.1.3 | 563156 | Measure | Label - Unplug the power input connector | A label shall be attached on Trolley housing to unplug the power input connector for service.
944 | 4.46 | 563103 | Chapter | Service System | 4.46.5 | 563151 | Function | UPS operation | 4.46.5.1 | 563152 | Hazard | Electrical shock | 4.46.5.1.2 | 563157 | Cause | Inappropiate devices connected to UPS | 4.46.5.1.2.1 | 563158 | Measure | User Manual - UPS | WARNINGCause: Unsuitable devices connected to the UPS. Possible consequences:Â Risk of electrical shock or damage to the system. Remedy: Do not connect any additional devices to the UPS sockets.
945 | 4.46 | 563103 | Chapter | Service System | 4.46.5 | 563151 | Function | UPS operation | 4.46.5.1 | 563152 | Hazard | Electrical shock | 4.46.5.1.2 | 563157 | Cause | Inappropiate devices connected to UPS | 4.46.5.1.2.2 | 563159 | Measure | Service Instruction UPS | Service DocumentsWARNINGCause: Inappropriate devices connected to UPSPossible consequences:Â Risk of electrical hazard or damage to the systemRemedy: Do not connect any additional device to the UPS power outlets
946 | 4.46 | 563103 | Chapter | Service System | 4.46.6 | 563160 | Function | Installation General | 4.46.6.1 | 563161 | Hazard | No hazard on system level | 4.46.6.1.1 | 563162 | Cause | None | 4.46.6.1.1.1 | 563163 | Measure | None
947 | 4.46 | 563103 | Chapter | Service System | 4.46.7 | 563164 | Function | Software installation | 4.46.7.1 | 563165 | Hazard | No hazard on system level | 4.46.7.1.1 | 563166 | Cause | None | 4.46.7.1.1.1 | 563167 | Measure | None
948 | 4.46 | 563103 | Chapter | Service System | 4.46.8 | 563168 | Function | Service General | 4.46.8.1 | 563169 | Hazard | No hazard on system level | 4.46.8.1.1 | 563170 | Cause | None | 4.46.8.1.1.1 | 563171 | Measure | None
949 | 4.46 | 563103 | Chapter | Service System | 4.46.9 | 563172 | Function | Report unexpected event | 4.46.9.1 | 563173 | Hazard | No hazard on system level | 4.46.9.1.1 | 563174 | Cause | None | 4.46.9.1.1.1 | 563175 | Measure | None
950 | 4.46 | 563103 | Chapter | Service System | 4.46.10 | 563176 | Function | Remote service | 4.46.10.1 | 563177 | Hazard | Malfunction of the examination routine | 4.46.10.1.1 | 563178 | Cause | Access via remote service | 4.46.10.1.1.1 | 563179 | Measure | Remote access | Only selected service functionalityÂ shall beÂ available for remote service in limited access mode. No customer functions shall be available during remote service in full access mode. Remote service shall be only possible when approved by the operator. An indicationÂ shall beÂ provided to the user that remote service with full access has been enabled. The user shall still be able to release radiation.
951 | 4.46 | 563103 | Chapter | Service System | 4.46.10 | 563176 | Function | Remote service | 4.46.10.2 | 563180 | Hazard | Remote software update | 4.46.10.2.1 | 563181 | Cause | Installing the hotfix or software update | 4.46.10.2.1.1 | 563182 | Measure | User Manual - Remote Software Update | CAUTIONCause: Software installation interferes with system functions. Possible consequences: Interruption of patient examination and loss of patient data. Remedy: Do not start the installation while working with patients.
952 | 4.46 | 563103 | Chapter | Service System | 4.46.10 | 563176 | Function | Remote service | 4.46.10.2 | 563180 | Hazard | Remote software update | 4.46.10.2.2 | 563183 | Cause | Failed installation of software | 4.46.10.2.2.1 | 563184 | Measure | User Manual - Remote Update Failure | CAUTIONCause: Failure installing the update package. Possible consequences: A failed software installation means that the system is in an undefined state. Remedy: Stop using the system and notify the UPTIME Service Center.
953 | 4.46 | 563103 | Chapter | Service System | 4.46.11 | 563185 | Function | Cables | 4.46.11.1 | 563186 | Hazard | Electric Shock | 4.46.11.1.1 | 563187 | Cause | Moving Components | 4.46.11.1.1.1 | 563188 | Measure | Reference - Cable conduit for movable cables | <XREF>562037</XREF>
954 | 4.46 | 563103 | Chapter | Service System | 4.46.12 | 563189 | Function | Service documentation | 4.46.12.1 | 563190 | Hazard | No hazard on system level | 4.46.12.1.1 | 563191 | Cause | None | 4.46.12.1.1.1 | 563192 | Measure | None
955 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.1 | 563194 | Hazard | Impaired vision through laser-beam | 4.46.13.1.1 | 563195 | Cause | Service staff looking into laser-beam | 4.46.13.1.1.1 | 563196 | Measure | Reference - Usage of laser class 1 | <XREF>562338</XREF>
956 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.2 | 563197 | Hazard | Excessive temperature, fire, explosion | 4.46.13.2.1 | 563198 | Cause | Improper exchange of batteries | 4.46.13.2.1.1 | 563199 | Measure | Warning Label - Exchange Battery UPS | A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries.
957 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.2 | 563197 | Hazard | Excessive temperature, fire, explosion | 4.46.13.2.1 | 563198 | Cause | Improper exchange of batteries | 4.46.13.2.1.2 | 563200 | Measure | Warning Label - Exchange Battery HICSwitch | A label shall be attached near the system batteries instructing the technician to consult the service manual before exchanging the batteries.
958 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.2 | 563197 | Hazard | Excessive temperature, fire, explosion | 4.46.13.2.2 | 563201 | Cause | Short-circuit of batteries during exchange | 4.46.13.2.2.1 | 563202 | Measure | Safe battery cases TrolleyU | Use battery cases or housings that prevent an accidental short-circuit
959 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.2 | 563197 | Hazard | Excessive temperature, fire, explosion | 4.46.13.2.2 | 563201 | Cause | Short-circuit of batteries during exchange | 4.46.13.2.2.2 | 563203 | Measure | Safe battery cases HICSwitch | Use battery cases or housings that prevent an accidental short-circuit
960 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.2 | 563197 | Hazard | Excessive temperature, fire, explosion | 4.46.13.2.2 | 563201 | Cause | Short-circuit of batteries during exchange | 4.46.13.2.2.3 | 563204 | Measure | Service Caution - Servicing battery components | Service DocumentationÂ CAUTIONÂ Cause: Incorrect handling of battery components during service,Â for example,Â wearing metal objects such as rings, watches or bracelets, or using tools with non-insulated handles when servicing the battery components, could lead to an accidentalÂ short-circuit of the battery contacts. Â Potential consequences: Risk of burns due to high short-circuit current and of destroying the battery component. Â Remedy: Before servicing the battery components, remove watches, rings, bracelets and all other metal objects from the hands and arms. Use tools with insulated handles.
961 | 4.46 | 563103 | Chapter | Service System | 4.46.13 | 563193 | Function | Maintenance, service / trouble shooting | 4.46.13.3 | 563205 | Hazard | Risk of Infection | 4.46.13.3.1 | 563206 | Cause | Opening Chassis / Trolley | 4.46.13.3.1.1 | 563207 | Measure | Service Instruction - Measures to prevent infection due to contaminated system parts - service | Warning in Service InstructionWARNINGCause: Service, disassembly or refurbishment activities on hardware parts ofÂ Cios systems which had been in clinical use and are potentially contaminated. Possible consequences: Increased risk of infection due to residues of infectious substances such as bodily fluids or deposits of organic matter colonized with microorganismsRemedy: Consequently apply precautionary measures for handling potentially infectious material, especially:- Wear protective gloves, closed and long-sleeved protective clothing, protective shoes, and protective glasses. - In case of severe contamination or strong infectiousness, in case the contamination may be dusting or emit droplets, apply additional measures such as face masks and disposable protective clothing, for example overalls and caps. - After finishing the activities on contaminated hardware parts, clean and disinfect potentially contaminated body parts, especially hands and face, as well as clothing and equipment.
962 | 4.46 | 563103 | Chapter | Service System | 4.46.14 | 563208 | Function | Imaging System, Field service | 4.46.14.1 | 563209 | Hazard | Wrong interventional decision | 4.46.14.1.1 | 563210 | Cause | Service parameters changed by unauthorized user | 4.46.14.1.1.1 | 563211 | Measure | Reference - User Manual - Trained Users | <XREF>561986</XREF>
963 | 4.46 | 563103 | Chapter | Service System | 4.46.14 | 563208 | Function | Imaging System, Field service | 4.46.14.1 | 563209 | Hazard | Wrong interventional decision | 4.46.14.1.1 | 563210 | Cause | Service parameters changed by unauthorized user | 4.46.14.1.1.2 | 563212 | Measure | Reference - Service training | <XREF>562186</XREF>
964 | 4.46 | 563103 | Chapter | Service System | 4.46.15 | 563213 | Function | 3D Service functions | 4.46.15.1 | 563214 | Hazard | Wrong interventional decision or treatment | 4.46.15.1.1 | 563215 | Cause | Unprecise 3D calibration due to a badly positioned calibration phantom | 4.46.15.1.1.1 | 563216 | Measure | Correct Position of 3D-Calibration Phantom | Service DocumentationCAUTIONCause: Imprecise 3DÂ calibrationPossible consequences: Wrong interventional decision or treatmentRemedy: Make sure that the calibration phantom is positioned correctly when calibrating the 3D system.
965 | 4.46 | 563103 | Chapter | Service System | 4.46.15 | 563213 | Function | 3D Service functions | 4.46.15.1 | 563214 | Hazard | Wrong interventional decision or treatment | 4.46.15.1.1 | 563215 | Cause | Unprecise 3D calibration due to a badly positioned calibration phantom | 4.46.15.1.1.2 | 580671 | Measure | Correct Position of 3D-Calibration Phantom | Admin GuideCAUTIONCause: Imprecise 3DÂ calibrationPossible consequences: Incorrect interventional decision or treatmentRemedy: When calibrating the 3D system, make sure that the calibration phantom is correctly positioned.
966 | 4.46 | 563103 | Chapter | Service System | 4.46.15 | 563213 | Function | 3D Service functions | 4.46.15.1 | 563214 | Hazard | Wrong interventional decision or treatment | 4.46.15.1.2 | 563217 | Cause | Wrong 3D-calibration due to wrong transformation matrices | 4.46.15.1.2.1 | 563218 | Measure | measuring instrument boards are positioned correctly | Service DocumentsCAUTIONCause:Â Incorrect 3D calibration (wrong transformation matrices) due to incorrect position of measuring instrumentsPossible Consequences: Wrong interventional decision or treatment. Remedy:Â Make sure that the measuring instrumentsÂ (navigation references like marker frame and marker plate) are positioned correctly in the measuring volume of the 3D camera when calibrating the 3D system.
967 | 4.46 | 563103 | Chapter | Service System | 4.46.15 | 563213 | Function | 3D Service functions | 4.46.15.2 | 563219 | Hazard | Risk of injury due to structure of surface | 4.46.15.2.1 | 563220 | Cause | Sharp edges | 4.46.15.2.1.1 | 563221 | Measure | Note in interface description | Note in interface description for nav. system:Â Attachable hardware shall have no sharp edges.
968 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.1 | 563224 | Hazard | Electrical safety, Electrical shock | 4.47.1.1.1 | 563225 | Cause | Electrical energy is present after shutdown. | 4.47.1.1.1.1 | 563226 | Measure | Warning label | Warning Labels show where voltage is present after shutdown.
969 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.1 | 563224 | Hazard | Electrical safety, Electrical shock | 4.47.1.1.1 | 563225 | Cause | Electrical energy is present after shutdown. | 4.47.1.1.1.2 | 563227 | Measure | Reference - Service Documentation - Voltage Presence | <XREF>563280</XREF>
970 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.2 | 563228 | Hazard | Loss of data | 4.47.1.2.1 | 563229 | Cause | System shutdown because of a software failure | 4.47.1.2.1.1 | 563230 | Measure | Reference - Protect shutdown | <XREF>563232</XREF>
971 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.2 | 563228 | Hazard | Loss of data | 4.47.1.2.2 | 563231 | Cause | User accidentally shuts off system during examination | 4.47.1.2.2.1 | 563232 | Measure | Protect shutdown | UPS must be available.
972 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.2 | 563228 | Hazard | Loss of data | 4.47.1.2.3 | 563233 | Cause | System restart after X-ray exposure | 4.47.1.2.3.1 | 563234 | Measure | No measure required
973 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.1 | 563223 | Function | Shut down system & restart | 4.47.1.3 | 563235 | Hazard | Delayed or no interventional decision | 4.47.1.3.1 | 563236 | Cause | System automatically switches on/off | 4.47.1.3.1.1 | 563237 | Measure | Reference - Display system status | <XREF>562150</XREF>
974 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.2 | 563238 | Function | Clean System | 4.47.2.1 | 563239 | Hazard | Risk of infection | 4.47.2.1.1 | 563240 | Cause | Improper cleaning | 4.47.2.1.1.1 | 563241 | Measure | User Manual - Cleaning of Wheels | CAUTIONCause: Improper cleaningPossible consequences: Risk of infection!Remedy: After dealing with patients with highly infectious diseases, clean the unit wheels, in addition to regular cleaning.
975 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.2 | 563238 | Function | Clean System | 4.47.2.1 | 563239 | Hazard | Risk of infection | 4.47.2.1.2 | 563242 | Cause | Contaminated device wheels | 4.47.2.1.2.1 | 563243 | Measure | Reference - User Manual - Cleaning of Wheels | <XREF>563241</XREF>
976 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.2 | 563238 | Function | Clean System | 4.47.2.1 | 563239 | Hazard | Risk of infection | 4.47.2.1.3 | 563244 | Cause | Insufficient cleaning | 4.47.2.1.3.1 | 563245 | Measure | Compliant with IEC60601-1 | The enclosure shall be designed in a way that it can be easily cleaned and wipe disinfected. Â The enclosure must be compliant with IEC60601-1:2005 chapter 11. 6. 6 & 7. 9. 2. 12 regarding cleaning and disinfection requirements. Â Fit criterion:All different enclosure surfaces must be compliant with MFR 890/009e - Resistance of surfaces to disinfection and cleaning agents.
977 | 4.47 | 563222 | Chapter | Shutdown System | 4.47.3 | 563246 | Function | Image data storage | 4.47.3.1 | 563247 | Hazard | Unnecessary radiation | 4.47.3.1.1 | 563248 | Cause | Image data loss | 4.47.3.1.1.1 | 563249 | Measure | The imaging system performs a controlled shutdown | The imaging system automatically performs a controlled shutdown, before the power of the Uninterrupted Power Supply runs out.
978 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.1 | 563251 | Function | Movements during patient positioning | 4.48.1.1 | 563252 | Hazard | Crushing of patient | 4.48.1.1.1 | 563253 | Cause | Contact of Patient with System | 4.48.1.1.1.1 | 563254 | Measure | User Manual - Crushing zones | CAUTIONCause: Movements of the C-arm during patient positioning. Possible Consequences: Crushing hazard for theÂ patient!Remedy: Be careful around the hazardous locations indicated.
979 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.3 | 563263 | Function | Virus Protection | 4.48.3.1 | 563264 | Hazard | Wrong or no interventional decision | 4.48.3.1.1 | 563265 | Cause | Degraded system performance | 4.48.3.1.1.1 | 563266 | Measure | Verify access to internet | There shall be only restricted access to internet possible for a clinical user (e.g. only some apphost related vendor URLs shall be allowed to access, but a normal internet usage shall not be possible). The service user shall be able to access internet via a special service environment
980 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.3 | 563263 | Function | Virus Protection | 4.48.3.1 | 563264 | Hazard | Wrong or no interventional decision | 4.48.3.1.1 | 563265 | Cause | Degraded system performance | 4.48.3.1.1.2 | 563267 | Measure | Verify access to BIOS | Access to BIOS needs to be password-protected.
981 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.1 | 563269 | Hazard | Hazards from fire or smoke | 4.48.4.1.1 | 563270 | Cause | Overheating | 4.48.4.1.1.1 | 563271 | Measure | Service Documentation usage of wrong fuses | Service Safety InstructionÂ CAUTIONCause: Usage of wrong fusesPossible Consequences:Â Risk ofÂ fireRemedy:Â Ensure the correct usage of fusesÂ taking intoÂ account the label next to the fuses
982 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.1 | 563269 | Hazard | Hazards from fire or smoke | 4.48.4.1.1 | 563270 | Cause | Overheating | 4.48.4.1.1.2 | 563272 | Measure | Reference to warning label next to fuses to use correct fuses | See HM_S1_hw_F2G1U2M1 in <XREF>578770</XREF>.
983 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.1 | 563269 | Hazard | Hazards from fire or smoke | 4.48.4.1.1 | 563270 | Cause | Overheating | 4.48.4.1.1.3 | 563273 | Measure | Reference to temperature monitoring of the ESU discharge resistors and inspection | See HM_S1_hw_F2G1U3M1 and HM_S1_hw_F2G1U3M2 in <XREF>578770</XREF>.
984 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.2 | 563274 | Hazard | Environmental hazard | 4.48.4.2.1 | 563275 | Cause | Improper disposal of hazardous materials | 4.48.4.2.1.1 | 563276 | Measure | Improper disposal of ESU (ultra capacitor cells with harmful and flammable liquid) | Which parts must be removed prior to scrapping and how these parts are disposed in an environmentalÂ friendly manner are documented in the Disposal Instructions (Service Documents). Â CAUTIONÂ Cause: ESU contains harmful and flammable liquid (Acetonitrile) and is not disposed correctlyPossible Consequences: Risk of injury and environmental damage. Remedy: For proper disposal, send the ESU back to the manufacturer, Siemens Healthineers.
985 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.2 | 563274 | Hazard | Environmental hazard | 4.48.4.2.1 | 563275 | Cause | Improper disposal of hazardous materials | 4.48.4.2.1.2 | 563277 | Measure | Reference to material request form and label on ESU | HM_S1_serv_F2G2U1M1 and HM_S1_esu_F2G2U1M2 in <XREF>578770</XREF>.
986 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.1 | 563279 | Cause | Voltage still present after switching-off | 4.48.4.3.1.1 | 563280 | Measure | Service Documentation - Voltage Presence | Service DocumentsWARNINGCause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Possible Consequences: Risk of death, serious physical injury or damage to property. Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels attached next to conducting parts.
987 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.1 | 563279 | Cause | Voltage still present after switching-off | 4.48.4.3.1.2 | 596998 | Measure | Factory Documentation - Voltage Presence | FactoryÂ DocumentsÂ WARNINGÂ Cause: Even after switching off the power supply, electrical voltage is present in the system. Live-line working can cause electric shock or mechanical damage through short circuiting. Possible Consequences: Risk of death, serious physical injury or damage to property. Remedy: De-energize the system. If checks or measurements require live-line working, cover energized parts to prevent unintentional contact. Perform live-line working with extreme caution. Use only measurement equipment and devices which are suited to the task. Read and observe the warning labels attached next to conducting parts.
988 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.1 | 563279 | Cause | Voltage still present after switching-off | 4.48.4.3.1.3 | 563282 | Measure | Reference to discharge circuits | See HM_S1_hw_F6G1U4M1 in <XREF>578770</XREF>
989 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.2 | 563283 | Cause | ESU requires discharge after switch-off | 4.48.4.3.2.1 | 563284 | Measure | Service Documentation ESU discharge | Service Safety InstructionÂ WARNINGÂ Cause: If the ESU is not discharged completely, electrical voltage is present and is accessible when e.g. main unit covers are removed. Possible consequences: Risk of death orÂ serious injury due to electric shock when electrified components are touched. Remedy: Discharge the ESU immediately after removing main unit covers before starting to work on relevant parts of the main unit, especially when removing or working on the ESU itself. If one of the fuses F21 or F20 is broken, the fast discharging procedure is not working. Verify that no voltage on the fuse holder is present before replacing broken fuses!
990 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.2 | 563283 | Cause | ESU requires discharge after switch-off | 4.48.4.3.2.2 | 563285 | Measure | Supply Chain Documentation ESU discharge | Supply Chain Safety InstructionÂ Warning:Â Â Cause: If the ESU is not discharged completely, electrical voltage is present. Live-line working can cause electric shock. Â Possible consequences: Risk of death, serious physical injury or damage to property. Â Remedy: Discharge the ESU prior to removing it.
991 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.2 | 563283 | Cause | ESU requires discharge after switch-off | 4.48.4.3.2.3 | 563286 | Measure | Reference to ESU discharge circuit, display of voltage, warning labels | SeeHM_S1_hw_F6G1U1M2HM_S1_esu_F6G1U1M3HM_S1_hw_F6G1U1M4HM_S1_hw_F6G1U1M5HM_S1_hw_F6G1U1M6HM_S1_mec_F6G1U1M1,HM_S1_mec_F6G1U2M1,HM_S1_mec_F6G1U2M2,HM_S1_hw_F6G1U3M1,HM_S1_manu_F6G1U3M2in <XREF>578770</XREF>.
992 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.3 | 580657 | Cause | Discharge unit defective, intermediate circuit not discharged completely | 4.48.4.3.3.1 | 580658 | Measure | Warning - Service Documentation - Check discharge of intermediate circuit | Service InstructionÂ WARNINGÂ Cause: If the intermediate circuit is not discharged completely, electrical voltage is present. Live-line working can cause electric shock. Possible consequences: Risk of death,Â serious injury or damage to property. Remedy: Check the intermediate circuit with suitable voltmeter on connector board M100.
993 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.3 | 563278 | Hazard | Electric shock | 4.48.4.3.3 | 580657 | Cause | Discharge unit defective, intermediate circuit not discharged completely | 4.48.4.3.3.2 | 580659 | Measure | Warning - Factory Documentation - Check discharge of intermediate circuit | Factory DocumentationÂ WARNINGÂ Cause: If the intermediate circuit is not discharged completely, electrical voltage is present. Live-line working can cause electric shock. Possible consequences: Risk of death,Â serious injury or damage to property. Remedy: Check the intermediate circuit with suitable voltmeter on connector board M100.
994 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.4 | 563287 | Hazard | Oil leaks out of the monoblock | 4.48.4.4.1 | 563288 | Cause | Component error | 4.48.4.4.1.1 | 563289 | Measure | Check leakage of monobloc | Service DocumentationCAUTIONCause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil)Possible consequences: Functional disturbances, leaking of liquids from the systemRemedy: Check the monoblock and the cooling unitÂ for leakage of liquids during service activities and replace defective parts.
995 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.4 | 563287 | Hazard | Oil leaks out of the monoblock | 4.48.4.4.1 | 563288 | Cause | Component error | 4.48.4.4.1.2 | 563290 | Measure | Reference to measures against leakage of oil | HM_S1_manu_F3G1U1M1, HM_S1_hw_F3G1U1M2, HM_S1_hw_F3G1U1M3, HM_S1_manu_F3G1U1M4, HM_S1_hw_F3G1U2M1, HM_S1_mec_F3G1U3M1 in <XREF>578770</XREF>.
996 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.4 | 563287 | Hazard | Oil leaks out of the monoblock | 4.48.4.4.1 | 563288 | Cause | Component error | 4.48.4.4.1.3 | 596995 | Measure | Check monobloc leakage | FactoryÂ DocumentationÂ CAUTIONÂ Cause: Undetected leakage of the monoblock or cooling unit (cooling liquid, oil)Possible consequences: Functional disturbances, leaking of liquids from the systemRemedy: Check the monoblock and the cooling unitÂ for leakage of liquids during service activities and replace defective parts.
997 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.5 | 563291 | Hazard | Leaking cooling liquid | 4.48.4.5.1 | 563292 | Cause | Component error or service error (Refill of Cooling Unit with connected Compensation Bag) | 4.48.4.5.1.1 | 563293 | Measure | Service Documentation - Disconnect Compensation Bag | CAUTIONCause: If the compensation bag is not disconnected from the cooling unit, the compensation bag can lacerate. Possible Consequences: Risk of burning or infection. Remedy: Prior to refilling the cooling unit, disconnect the compensation bag.
998 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.5 | 563291 | Hazard | Leaking cooling liquid | 4.48.4.5.1 | 563292 | Cause | Component error or service error (Refill of Cooling Unit with connected Compensation Bag) | 4.48.4.5.1.2 | 563294 | Measure | Reference to measures against leakage of cooling liquid | See HM_S1_mec_F3G1U4M1, HM_S1_manu_F3G1U4M2, HM_S1_manu_F3G1U4M3, HM_S1_serv_F3G1U4M4, HM_S1_hw_F3G1U4M5 in <XREF>578770</XREF>.
999 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.4 | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 4.48.4.6 | 563295 | Hazard | Unnecessary radiation | 4.48.4.6.1 | 563296 | Cause | Missing error signal about image loss | 4.48.4.6.1.1 | 563297 | Measure | Missing HV_ON signal | The system must ensure that missing HV_ON signal shall not lead to a patient harm. [See also HM_S1_hw_F5G3U1M1 in <XREF>578770</XREF>]
1000 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.1 | 563299 | Hazard | Risk of burnings | 4.48.5.1.1 | 563300 | Cause | Too high temperature of the detector | 4.48.5.1.1.1 | 563301 | Measure | User Manual - Detector Temperature | CAUTIONCause:Â Detector temperature is not within recommended range. Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of X-ray radiation exposure without any diagnostic value!Remedy: Regularly check the temperature icon on the display.
1001 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.1 | 563299 | Hazard | Risk of burnings | 4.48.5.1.1 | 563300 | Cause | Too high temperature of the detector | 4.48.5.1.1.2 | 563302 | Measure | Reference to System Warning - Temperature of the Detector | <XREF>562226</XREF>.
1002 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.2 | 563303 | Hazard | Malfunction of imaging system during operation due to defective HW | 4.48.5.2.1 | 563304 | Cause | Defective FLC-Hardware | 4.48.5.2.1.1 | 563305 | Measure | User Manual - Defective FLC HW | CAUTIONCause: Defective imaging system hardwarePossible consequences: Risk of X-ray radiation without a stored imageRemedy: In case of image failures stop exposures and switch off the system.
1003 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.3 | 563306 | Hazard | Risk of incorrect interventional decision caused by wrong image processing | 4.48.5.3.1 | 563307 | Cause | Wrong LUT | 4.48.5.3.1.1 | 563308 | Measure | User Manual - Wrong LUT | CAUTIONCause: Wrong or extreme image processing parameters. Possible consequences: Risk of limited image quality in the resulting X-ray images. Risk of incorrect decision during intervention or artifacts. Remedy:- Apply image processing parameters carefully, such as edge enhancement and windowing, and always check them carefully. - Always check the image quality before reaching an interventional decision.
1004 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.1 | 563310 | Cause | Plugging/Unplugging USB devices interfere with performance of imaging system | 4.48.5.4.1.1 | 563311 | Measure | User Manual - Plug/Unplug USB Devices | CAUTIONCause: Risk of connecting/disconnecting USB devices during the exposure. Possible consequences: System instability, impact onÂ or disturbance of running acquisition tasks, loss of data. Remedy: Do not insert or disconnect USB devices during exposure or other critical processes.
1005 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.2 | 563312 | Cause | Unauthorized installation of SW/HW | 4.48.5.4.2.1 | 563313 | Measure | Reference to User Manual - Unauthorized Software | <XREF>562152</XREF>
1006 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.3 | 563314 | Cause | Network failure | 4.48.5.4.3.1 | 563315 | Measure | User Manual - Network Failure | CAUTIONCause: Network connection may not be available continuously. Possible consequences: Reference images not available during intervention. Remedy: Before beginning the intervention, the user should load the required images from the network archive.
1007 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.4 | 563316 | Cause | Software lockup | 4.48.5.4.4.1 | 563317 | Measure | User Manual - Software Lockup | CAUTIONCause: Software crash. Possible consequences: Working with the imaging system is no longer possible. Remedy: Restart the system.
1008 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.5 | 563318 | Cause | Restart missing after Service session | 4.48.5.4.5.1 | 563319 | Measure | User Manual - Reboot after Service | CAUTIONCause: No restart after service session. Possible consequences: Changes to the service environment may not go into effect without a system restart. Remedy: Restart the system after a service session.
1009 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.4 | 563309 | Hazard | Delayed or no interventional decision | 4.48.5.4.5 | 563318 | Cause | Restart missing after Service session | 4.48.5.4.5.2 | 563320 | Measure | Message to request reboot after service | A message is displayed to restart the system when the Service environment is left.
1010 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.5 | 563321 | Hazard | Danger of incorrect interventional decision | 4.48.5.5.1 | 563322 | Cause | Assignment of series/ scenes/ images to the wrong patient | 4.48.5.5.1.1 | 563323 | Measure | User Manual - Mix-up of Patients | CAUTIONCause: Risk of mistaking patients. Possible consequences: The wrong patient is being treated. Remedy: The user must ensure the correct patient is being treated before he or she is registered on the system.
1011 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.5 | 563321 | Hazard | Danger of incorrect interventional decision | 4.48.5.5.2 | 563324 | Cause | Scale displayed for non-calibrated images | 4.48.5.5.2.1 | 563325 | Measure | Reference to Caution for 2D measurements | <XREF>562590</XREF>.
1012 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.6 | 563326 | Hazard | loss of images or patient data | 4.48.5.6.1 | 563327 | Cause | Postponed loss of recorded patient data | 4.48.5.6.1.1 | 563329 | Measure | User Manual - Loss of Data USB | CAUTIONCause: Disconnecting USB devices during ongoing data transfer. Possible consequences: Loss of data and damage to the data medium possible. Remedy:Â Â - Do not disconnect the USB device during active data transfer (such as export). -Â Use removable devices as transfer media only. Removable devices are not permitted for archiving purposes.
1013 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.6 | 563326 | Hazard | loss of images or patient data | 4.48.5.6.2 | 563331 | Cause | Harddisk or file system defect | 4.48.5.6.2.1 | 563332 | Measure | User Manual - Hard Disk Failure | CAUTIONCause: Use the imaging system hard disk as a long-term archive. Possible consequences: Loss of image dataRemedy: Do not use the imaging system hard disk as a long-term archive. Follow the local regulations regarding the archiving of X-ray images, while taking into account the use of DICOM archive nodes.
1014 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.5 | 563298 | Function | Documentation user manual for Imaging System | 4.48.5.7 | 563333 | Hazard | loss of image quality | 4.48.5.7.1 | 563334 | Cause | Incorrect calibration data used | 4.48.5.7.1.1 | 563335 | Measure | Note in service documentation | Note in service documentation:Correct FD calibration in regular intervals
1015 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.1 | 563337 | Hazard | Electric Shock | 4.48.6.1.1 | 563338 | Cause | Intrusion of Liquids and Instruments | 4.48.6.1.1.1 | 563339 | Measure | Reference - User Manual - Insufficient cleaning | <XREF>562240</XREF>
1016 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.1 | 563337 | Hazard | Electric Shock | 4.48.6.1.1 | 563338 | Cause | Intrusion of Liquids and Instruments | 4.48.6.1.1.2 | 563340 | Measure | Reference - Spillage of Liquids on System | <XREF>562549</XREF>
1017 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.1 | 563337 | Hazard | Electric Shock | 4.48.6.1.1 | 563338 | Cause | Intrusion of Liquids and Instruments | 4.48.6.1.1.3 | 563341 | Measure | User Manual - Intrusion of Instruments xCP | WARNING:Cause: Instruments falling into the gap between the control unit and its housing. Possible Consequences: Short circuit, electric shock, and/or fire!Remedy: Make sure no instruments, conductive material, or fluids penetrate into the openings/joints between the control unit and its housing.
1018 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.2 | 563342 | Hazard | Overheating | 4.48.6.2.1 | 563343 | Cause | Inappropriate Draping | 4.48.6.2.1.1 | 563344 | Measure | User Manual - Use Specified Drapes Only | CAUTIONCause:Â Remote control unit covered by unsuitable covers. Possible Consequences: Overheating of the remote control unit!Remedy: Only use the sterile covers specified.
1019 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.2 | 563342 | Hazard | Overheating | 4.48.6.2.2 | 563345 | Cause | Hot Surface of RCP | 4.48.6.2.2.1 | 563346 | Measure | User Manual - Avoid Patient Contact with RCP | CAUTIONCause: The surfaces of the remote control unit become hot. Possible Consequences: Risk of burns!Remedy:- Make sure that the patient does not touch the remote control unit for more than 10 minutes or that after 10 minutes can no longer come in contact with it. - As a rule, contact between the patient and the remote control unit should be avoided.
1020 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.2 | 563342 | Hazard | Overheating | 4.48.6.2.2 | 563345 | Cause | Hot Surface of RCP | 4.48.6.2.2.2 | 563347 | Measure | User Manual - Positioning of RCP | <XREF>562487</XREF>
1021 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.2 | 563342 | Hazard | Overheating | 4.48.6.2.2 | 563345 | Cause | Hot Surface of RCP | 4.48.6.2.2.3 | 563348 | Measure | Reference - Compliance with IEC 60601-1 | Compliance with IEC 60601-1, 04S451630 in <XREF>561856</XREF>
1022 | 4.48 | 563250 | Chapter | Additional hazards from subsystems | 4.48.6 | 563336 | Function | User Manual for Touch Interfaces (xCP) | 4.48.6.3 | 563349 | Hazard | Risk of injury | 4.48.6.3.1 | 563350 | Cause | Falling Down of RCP | 4.48.6.3.1.1 | 563351 | Measure | User Manual - Falling Down | CAUTIONCause: Remote control unit falls down. Possible Consequences: Risk of injury/crushing!Remedy: It is recommended to keep the remote control unit in the holder provided. If a holder is not available, place the remote control unit on the instrument table.
1023 | 4.50 | 563353 | Chapter | Processes employed | 4.50.1 | 563354 | Function | Refurbishment | 4.50.1.1 | 563355 | Hazard | Biological contamination or infection during refurbishment activities | 4.50.1.1.1 | 563356 | Cause | Contact with contaminated system parts during deinstallation or cleaning for refurbishment | 4.50.1.1.1.1 | 563357 | Measure | Measures to prevent infection due to contaminated system parts - refurbishment | Warning in instruction for refurbishment:Â Dienstleisterspezifikation <XREF>561903</XREF>. Â For text, see <XREF>563207</XREF>.
1024 | 4.50 | 563353 | Chapter | Processes employed | 4.50.2 | 563358 | Function | Service / maintenance | 4.50.2.1 | 563359 | Hazard | Biological contamination or infection during service activities | 4.50.2.1.1 | 563360 | Cause | Contact with contaminated system parts during service activities | 4.50.2.1.1.1 | 563361 | Measure | Reference to measures to prevent infection due to contaminated system parts - service | <XREF>563207</XREF>.
1025 | 4.50 | 563353 | Chapter | Processes employed | 4.50.2 | 563358 | Function | Service / maintenance | 4.50.2.2 | 563362 | Hazard | Undetected damage of the cooling unit | 4.50.2.2.1 | 563363 | Cause | Overpressure test of cooling unit assemly not performed correctly | 4.50.2.2.1.1 | 563364 | Measure | Service Documentation - Overpressure test | Service DocumentationCAUTIONCause: Overpressure test is performed with cooling unit components and hose line including the flow indicator while connected to the system. Possible consequences: Undetected damage of the cooling unit components. Remedy: Make sure that the cooling unit assembly and the hose line including the flow indicatorÂ are disconnected from the system during overpressure test.
1026 | 4.50 | 563353 | Chapter | Processes employed | 4.50.3 | 563365 | Function | Manufacturing | 4.50.3.1 | 578701 | Hazard | None | 4.50.3.1.1 | 578702 | Cause | None | 4.50.3.1.1.1 | 578703 | Measure | None
1027 | 4.50 | 563353 | Chapter | Processes employed | 4.50.4 | 563366 | Function | Supply chain management process | 4.50.4.1 | 578704 | Hazard | None | 4.50.4.1.1 | 578705 | Cause | None | 4.50.4.1.1.1 | 578706 | Measure | None
1028 | 4.50 | 563353 | Chapter | Processes employed | 4.50.5 | 563367 | Function | Packaging and transport | 4.50.5.1 | 578707 | Hazard | None | 4.50.5.1.1 | 578708 | Cause | None | 4.50.5.1.1.1 | 578709 | Measure | None
1029 | 4.50 | 563353 | Chapter | Processes employed | 4.50.6 | 563368 | Function | Disposal | 4.50.6.1 | 578710 | Hazard | None | 4.50.6.1.1 | 578711 | Cause | None | 4.50.6.1.1.1 | 578712 | Measure | None
1030 | 4 | 561938 | Chapter | Evaluation of Functions, Hazards, Risks and Measures
1031 | 4.2 | 563439 | Chapter | Startup main unit in standalone mode
1032 | 4.5 | 563432 | Chapter | Drape the system
1033 | 4.6 | 563433 | Chapter | Connect trolley to line power
1034 | 4.7 | 563434 | Chapter | Connect trolley and main unit
1035 | 4.9 | 563435 | Chapter | Register patient
1036 | 4.11 | 563437 | Chapter | Position patient
1037 | 4.13 | 563440 | Chapter | Move main unit axles - manual
1038 | 4.14 | 563441 | Chapter | Store position
1039 | 4.16 | 563442 | Chapter | Adjust the field of view
1040 | 4.17 | 563443 | Chapter | Select and adjust examination settings and imaging parameters
1041 | 4.20 | 563460 | Chapter | Display dose related information during radiation
1042 | 4.21 | 563454 | Chapter | Postprocessing - 2D images
1043 | 4.22 | 563455 | Chapter | Postprocessing - Subtraction
1044 | 4.23 | 563456 | Chapter | Postprocessing - 3D
1045 | 4.24 | 563457 | Chapter | Postprocessing - Panorama
1046 | 4.26 | 563462 | Chapter | Perform 2D navigation
1047 | 4.28 | 563464 | Chapter | Operate application on application host
1048 | 4.29 | 563458 | Chapter | Protocol examination data
1049 | 4.30 | 563453 | Chapter | Administer patient data
1050 | 4.31 | 563466 | Chapter | Close patient
1051 | 4.32 | 563465 | Chapter | Delete images
1052 | 4.33 | 563459 | Chapter | Shutdown system
1053 | 4.34 | 563469 | Chapter | Disconnect trolley and main unit
1054 | 4.35 | 563470 | Chapter | Disconnect trolley from line power
1055 | 4.38 | 563471 | Chapter | Install the system
1056 | 4.39 | 563472 | Chapter | Service the system
1057 | 4.40 | 563473 | Chapter | Refurbish the system
1058 | 4.41 | 563474 | Chapter | Dispose the system
1059 | 4.49 | 563352 | Chapter | Accessories and combinations
1060 | 4.51 | 563370 | Chapter | Usability aspects
1061 | 4.52 | 563369 | Chapter | Warnings and cautions from supplied components
1062 | 4.53 | 563371 | Chapter | Hazards to environment
1063 | 4.54 | 563372 | Chapter | Biological and chemical hazards
1064 | 4.55 | 563373 | Chapter | Hazards through outsourcing of processes
1065 | 4.56 | 563374 | Chapter | Hazards by building services
1066 | 4.57 | 563376 | Chapter | Hazards by environmental conditions during usage
1067 | 4.58 | 595404 | Chapter | Hazard by to visual electromagnetic radiation
1068 | ENDE | 0 | Chapter