## result_table fuer map
Chapter ID | Chapter Type | Chapter Requirement | Function ID | Function Type | Function Requirement | Hazard ID | Hazard Type | Hazard Requirement | Cause ID | Cause Type | Cause Requirement | Measure ID | Measure Type | Measure Requirement | Measure Description
563430 | Chapter | Move trolley | 569143 | Function | Move trolley | 561941 | Hazard | Crushing due to trolley tipping over | 561942 | Cause | Movement over inclined surface or monitor movement changes the center of gravity of the trolley | 561943 | Measure | Trolley design prevents overbalance
563430 | Chapter | Move trolley | 569143 | Function | Move trolley | 561944 | Hazard | Crushing through tilting of the device | 561947 | Cause | Increase of tilt moment due to additional weight | 561949 | Measure | Design of the trolley
563430 | Chapter | Move trolley | 569143 | Function | Move trolley | 561944 | Hazard | Crushing through tilting of the device | 561950 | Cause | Moving the trolley over a doorstep, against a doorframe, wall or other equipment | 561951 | Measure | Rough handling test trolley
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564256 | Cause | Short circuit of the battery (external) | 564257 | Measure | Coded battery connector
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564256 | Cause | Short circuit of the battery (external) | 564258 | Measure | Battery housing preventing accidental short circuit
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564256 | Cause | Short circuit of the battery (external) | 564259 | Measure | Wiring check battery in manufacturing
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564267 | Cause | Overcharging of the battery (voltage, current) | 564260 | Measure | Prevent overcharging of battery
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 571858 | Cause | Overcharging of the battery, sleeping error in P path | 571859 | Measure | Self-test of battery overcharging protection
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564268 | Cause | Internal short circuits and temperature increase of the battery packs | 564261 | Measure | Prevent charging of battery with high temperature
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564268 | Cause | Internal short circuits and temperature increase of the battery packs | 577242 | Measure | Ventilation battery housing
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 571860 | Cause | Internal short circuits and temperature increase of the battery packs, sleeping error in P path | 571861 | Measure | Self-test of protection against charging of battery with high temperature
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564269 | Cause | Mechanical destruction | 564262 | Measure | Battery housing mechanical strength
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564269 | Cause | Mechanical destruction | 564263 | Measure | Battery discharge current limitation
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564270 | Cause | Thermal runaway of battery pack (internal short circuit, overheating) | 564265 | Measure | Battery housing fireproof and preventing flame break-out
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 564255 | Hazard | Explosion, fire (HS_056, HS_030) | 564270 | Cause | Thermal runaway of battery pack (internal short circuit, overheating) | 564264 | Measure | Thermal isolation of single battery cells
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 577236 | Cause | Increased temperature of the battery packs | 577237 | Measure | Reference - Prevent charging of battery with high temperature
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 577236 | Cause | Increased temperature of the battery packs | 577241 | Measure | Reference - Prevent overcharging of battery
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 577236 | Cause | Increased temperature of the battery packs | 577238 | Measure | Reference - Battery housing mechanical strength
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 577236 | Cause | Increased temperature of the battery packs | 577239 | Measure | Reference - Battery discharge current limitation
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 577220 | Hazard | Inhalation of toxic fumes when battery material is burnt or exposed to fire (HS_084) | 577236 | Cause | Increased temperature of the battery packs | 577240 | Measure | Reference - Thermal isolation of single battery cells
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 569265 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569264 | Cause | Activation of control elements by non-authorized persons (e. g. passer-by in the corridor) | 564347 | Measure | Lock movement initiation from chassis handle
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 569265 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569264 | Cause | Activation of control elements by non-authorized persons (e. g. passer-by in the corridor) | 569266 | Measure | Caution - Lock chassis movement or switch off
563431 | Chapter | Move main unit chassis in standalone mode | 564254 | Function | Move main unit chassis in standalone mode | 569291 | Hazard | Severe limitation of movement functionality (HS_111) | 569292 | Cause | Chassis battery runs empty, no more motorized movements possible | 568840 | Measure | Indicate battery charge for standalone mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564116 | Measure | Collision avoidance in fast drive mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 562463 | Measure | Reference - Set speed values
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 569215 | Measure | Reference - Set acceleration values
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564118 | Measure | Additional restrictions for fast drive mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564119 | Measure | Timeout for fast drive mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564120 | Measure | Reference - Maximum overtravel values
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564121 | Measure | Reference - System speed limits
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564122 | Measure | Reference - Collision detection
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564123 | Measure | Reference - Collision detection at rear side of the chassis
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564124 | Measure | Reference - Warning - General risk of collisions
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564115 | Cause | User does not pay attention to obstacles in the movement path during transport movement (intentional) | 564125 | Measure | Reference - Emergency stop at every point of motorized movement activation
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 571908 | Measure | Reference - Self-test of protection module hardware
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 578162 | Measure | Reference - Safe power supply of protection module hardware
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 571909 | Measure | Reference - System speed limits
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571907 | Cause | Wrong supervision limits are applied due to error or sleeping or unexpected error in P path | 571910 | Measure | Reference - Additional restrictions for transport mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 571914 | Measure | Fast drive button protected against unintended activation
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 571912 | Measure | Reference - Timeout for transport mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571911 | Cause | Unintentional activation of fast drive mode by the operator | 571913 | Measure | Indicate active fast drive mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 578308 | Cause | Unintentional activation of fast drive mode by technical error | 578309 | Measure | Safe activation fast drive mode
564113 | Chapter | Move main unit chassis in transport mode | 564112 | Function | Move main unit chassis in transport mode | 564114 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 578308 | Cause | Unintentional activation of fast drive mode by technical error | 571915 | Measure | Check fast drive mode button state
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 562344 | Measure | Check movement release signal status at startup
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 562378 | Measure | Control elements protected against unintended movement or x-ray activation
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 562377 | Measure | Continuous activation of movements with redundant activation paths
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563480 | Cause | Unintentional activation of movement during startup (by operator, e. g. control element clamped, or by technical error) | 562345 | Measure | Emergency stop at every point of motorized movement activation
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 571864 | Measure | Self-test of protection module hardware
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 571866 | Measure | Reference - Check movement release signal status at startup
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 578161 | Measure | Safe power supply of protection module hardware
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 563479 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571863 | Cause | Unintentional activation of movement during startup, sleeping or unexpected error in P path | 578307 | Measure | General principles protective path
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 569274 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 569275 | Cause | X-ray chain is not functional (e. g. communication error, hardware error) | 564352 | Measure | X-ray chain selftest at startup
563429 | Chapter | Startup the system via trolley | 563478 | Function | Startup the system via trolley | 569274 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 569297 | Cause | Delayed boot-up or restart time to recover system functions (e. g. after error, SW freeze) | 568860 | Measure | Recovery management
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 563627 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or high contact pressure (HS_051) | 563629 | Cause | Mechanical support structures of the wheels break, the system suddenly tips to one side or tips over completely | 561959 | Measure | Mechanical design drive and wheels
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 578748 | Hazard | Limited imaging functionality | 578749 | Cause | Performance of the imaging system is degraded due to system mangement application | 578750 | Measure | Boundary conditions for imaging system performance test
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 597004 | Measure | Label electromagnetic fields
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 597005 | Measure | User Manual: Electromagnetic fields Chassis
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 597007 | Measure | Service Doc: Electromagnetic fields Chassis
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 597002 | Hazard | Effect on people with implants, e.g. pacemaker | 597003 | Cause | Interference of electromagnetic fields into implant/s electronics | 597011 | Measure | Factory Doc: Electromagnetic fields Chassis
563628 | Chapter | System operation - general | 563631 | Function | System operation - general | 597013 | Hazard | User has skin irritation | 597014 | Cause | System contains non-biocompatible material | 597019 | Measure | Biocompatible material
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 563485 | Measure | Reference - Control elements protected against unintended movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 578331 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 578316 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 578332 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 563488 | Measure | Prerequisites for fast drive mode
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 578291 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 578358 | Measure | Reference - Collision detection at rear side of the chassis and FD
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 571916 | Measure | Collision detection via position change of other axles
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 563493 | Measure | Warning - General risk of collisions
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 563494 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562462 | Cause | Unintentional activation of control elements by the operator (lean over, hang something over control elements, falling parts ) | 563495 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 571869 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 571870 | Measure | Positioning validity of motorized axes
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 578163 | Measure | Independent power supply of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 578164 | Measure | Reference - Safe power supply of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571868 | Cause | Speed limits are exceeded due to sleeping or unexpected error in P path | 595374 | Measure | Reference - Prevent incorrect movement execution - mecanum
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571872 | Cause | Collision detection fails due to sleeping error in P path | 571874 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571872 | Cause | Collision detection fails due to sleeping error in P path | 571873 | Measure | Self-test of the collision sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571872 | Cause | Collision detection fails due to sleeping error in P path | 571876 | Measure | Reliable design of collision sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571872 | Cause | Collision detection fails due to sleeping error in P path | 563492 | Measure | Warning - Check collision sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563645 | Cause | Unintentional activation of movement due to technical failure (without the operator's request / approval) | 563646 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563645 | Cause | Unintentional activation of movement due to technical failure (without the operator's request / approval) | 578359 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563622 | Cause | User does not pay attention to feet in the movement path during movement (intended): System drives over toes | 563625 | Measure | Mechanical safety - allowed gaps mecanum
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563622 | Cause | User does not pay attention to feet in the movement path during movement (intended): System drives over toes | 599924 | Measure | Reference - warning label risk of crushing foot
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563638 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563639 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563640 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563641 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563642 | Measure | Reference - Collision detection at rear side of the chassis
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563643 | Measure | Reference - Warning - General risk of collisions
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563636 | Cause | User does not pay attention to obstacles in the movement path during movement (intended): User reaction, uncontrolled speed, etc | 563644 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563674 | Cause | System does not move according to control parameters: Resulting system speed too high due to technical error | 563676 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 563680 | Measure | System acceleration and deceleration limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 563681 | Measure | Limit system centrifugal acceleration
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563679 | Cause | System does not move according to control parameters: System accelerates hazardously (e. g. max. possible by design) | 563682 | Measure | Reference - Plausibility check for sensors in movement control elements
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 571878 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 578166 | Measure | Reference - Safe power supply of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 571879 | Measure | Reference - Self-test of the protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 571880 | Measure | Reference - Mounting of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571877 | Cause | System does not move according to control parameters: System accelerates hazardously, sleeping or unexpected error in P path | 578165 | Measure | Reference - Safe power supply of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 580816 | Cause | System does not move according to control parameters: During quickstop, the system violates trajectory trueness | 596883 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 563719 | Measure | Behaviour in collision
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 563722 | Measure | Torque limits mecanum
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 563723 | Measure | Reference - Warning - General risk of collisions
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563720 | Cause | Collision sensor is active, operator moves further into collision by mistake | 576754 | Measure | Reference - Override angular / orbital brakes manually
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 571894 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 578167 | Measure | Reference - Safe power supply of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 571895 | Measure | Reference - Self-test of the protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 571896 | Measure | Reference - Mounting of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571893 | Cause | Collision sensor is active, behaviour in collision is incorrect due to sleeping or unexpected error of P path | 578168 | Measure | Reference - Safe power supply of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571881 | Cause | Collision sensor is active, operator moves further into collision by technical error (e. g. joystick deflection sensor) | 571882 | Measure | Reference - Behaviour in collision
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571881 | Cause | Collision sensor is active, operator moves further into collision by technical error (e. g. joystick deflection sensor) | 571883 | Measure | Reference - Torque limits mecanum
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563747 | Cause | System does not stop when user releases control element or stops too late | 563748 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563747 | Cause | System does not stop when user releases control element or stops too late | 563749 | Measure | Safe stop motorized movements
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563747 | Cause | System does not stop when user releases control element or stops too late | 563750 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 571885 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 578169 | Measure | Reference - Safe power supply of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 571886 | Measure | Reference - Self-test of the protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 578170 | Measure | Reference - Safe power supply of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 571887 | Measure | Reference - Mounting of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 576755 | Measure | Self-test STO functionality
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571884 | Cause | System does not stop when user releases control element or stops too late due to sleeping error in P path | 576756 | Measure | Reference - Self-test of the brakes
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 563759 | Measure | Reference - Control elements protected against unintended movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 563760 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 563761 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563758 | Cause | System does not stop when user releases control element: Erroneous prolongation of movement due stuck movement key | 563762 | Measure | Reference - Warning - General risk of collisions
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 562459 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 563763 | Cause | System does not stop when user releases control element: Overtravel is increased due to slip , e. g. fluid on the floor | 563765 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 563727 | Measure | Reference - System acceleration and deceleration limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 563728 | Measure | Reference - Limit system centrifugal acceleration
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 563729 | Measure | Prevent incorrect movement execution - mecanum
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563726 | Cause | System tips over during movement: One drive blocks, one brake gets activated, random HW error | 563730 | Measure | Prevent tip over by mechanical design
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 571889 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 578171 | Measure | Reference - Safe power supply of protection module hardware
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 571890 | Measure | Reference - Self-test of the protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 571891 | Measure | Reference - Mounting of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 571888 | Cause | System tips over during movement, incorrect mecanum movement, sleeping or unexpected error in P path | 578172 | Measure | Reference - Safe power supply of protective position sensors
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563731 | Cause | System tips over during movement or during braking: Safety mechanism is activated (e. g. quickstop is not initiated at all) | 563732 | Measure | Reference - Prevent tip over by mechanical design
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 563734 | Measure | Weight restriction mounting external components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 563735 | Measure | Warning - Weight limit for additionally mounted components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 563770 | Measure | Reference - Prevent tip over by mechanical design
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 578343 | Measure | Warning - Inclined surface restriction
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563733 | Cause | System tips over during movement or during braking: Additional weight is mounted at the C-arm | 596712 | Measure | Warning - Service instruction for inclined surfaces
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563736 | Cause | System tips over during movement or during braking process: Curve radius too small in relation to speed | 563737 | Measure | Reference - Limit system centrifugal acceleration
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563766 | Cause | System tips over during movement or during braking process: Normal braking ramp is too harsh | 563767 | Measure | Reference - Prevent tip over by mechanical design
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 563766 | Cause | System tips over during movement or during braking process: Normal braking ramp is too harsh | 563768 | Measure | Reference - System acceleration and deceleration limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563725 | Hazard | Crushing by or collision with falling or tipping objects with high potential energy and / or contact force / pressure (HS_051) | 580818 | Cause | System lifts at one side during braking process | 580819 | Measure | Reference - Prevent tip over by mechanical design
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 563634 | Measure | Only one active movement control element at a time
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 563635 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 569307 | Measure | No movement take over across controls
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563633 | Cause | Two operators attempt to initiate system movements at the same time (intentionally) | 569308 | Measure | Redundant signals from same control button
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563648 | Measure | Movement preview for position movements
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563650 | Measure | Warning - Collisions during position movements without test run
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563651 | Measure | Orientation indication for remote movement controls
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 564191 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563652 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563655 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563654 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563656 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563657 | Measure | Reference - Collision detection at rear side of the chassis
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563658 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563647 | Cause | System moves in wrong direction due to operating error (intentional movement): mix-up of directions or buttons etc. | 563659 | Measure | Reference - Warning - General risk of collisions
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563671 | Cause | System moves in wrong direction due to technical error (intentional movement): force handles | 563672 | Measure | Deactivate whole chassis handle when error in movement activation is detected
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563671 | Cause | System moves in wrong direction due to technical error (intentional movement): force handles | 563673 | Measure | Plausibility check for sensors in movement control elements
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563684 | Measure | Main Unit movement shock detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563685 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 564164 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563687 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563688 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563689 | Measure | Cable deflectors and wheel protection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 563018 | Measure | Reference - Weight restriction mounting external components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563683 | Cause | System does not move according to control parameters, does not follow the planned movement path due to external factors | 561948 | Measure | Reference - Warning - Weight limit for additionally mounted components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563691 | Measure | Reference - Main Unit shock detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 599809 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563692 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563693 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563695 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563696 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 563697 | Measure | Reference - Mechanical design drive and wheels
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 568836 | Measure | Classification for mechanical shock and vibration - operation
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 568848 | Measure | Movement check in manufacturing
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563690 | Cause | System does not move according to control parameters, does not follow the planned movement path due to technical errors | 568847 | Measure | Movement check after service / maintenance
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 563699 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 599810 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 563700 | Measure | Reference - Mechanical design drive and wheels
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563698 | Cause | System does not move according to control parameters, does not follow the planned movement path due blocking of one drive | 563702 | Measure | Reference - Cable deflectors and wheel protection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 563755 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 563756 | Measure | Reference - Weight restriction mounting external components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 563757 | Measure | Reference - Warning - Weight limit for additionally mounted components
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 563751 | Cause | System does not move according to control parameters, does not follow the planned movement path: One wheel incorrect | 563752 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 571853 | Cause | Brakes have been opened manually, subsequent motorized movement is performed with brakes open, arbitrary movement of axle | 571854 | Measure | Activate other brakes for motorized movement
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563632 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 580821 | Cause | During braking process, orbital and angular axle are moved unintentionally due to inertia (overridden brakes) | 580822 | Measure | Reference - Braking force
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 563706 | Cause | Chassis movement is blocked: E. g. Brakes cannot be unlocked due to power supply failure, motor not functional | 562205 | Measure | Warning - Establish emergency procedures
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 563718 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 563721 | Measure | Reference - Behaviour in collision
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563704 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 563718 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 571892 | Measure | No additional measures are required
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 580797 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 580798 | Cause | Unintended movement activation with activated e-stop: Operator releases manual brakes during active e-stop | 580799 | Measure | Reference - Behaviour in collision
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563705 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 563707 | Cause | System gets damaged due to a previous collision | 563708 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563705 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 563741 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 563742 | Measure | Reference - Prevent tension at trolley-chassis cable
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563709 | Hazard | Critical influence on external electric devices (HS_011) | 563710 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 563712 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563709 | Hazard | Critical influence on external electric devices (HS_011) | 563710 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 562061 | Measure | Warning - EMC disturbances of other equipment
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563714 | Hazard | Non-critical influence on external electric devices (HS_010) | 563715 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 563716 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563714 | Hazard | Non-critical influence on external electric devices (HS_010) | 563715 | Cause | Electro-magnetic emissions of the drives in combination with the other emissions of the system | 563717 | Measure | Reference - Warning - EMC disturbances of other equipment
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 563739 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 563740 | Measure | Prevent tension at trolley-chassis cable
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 563739 | Cause | System is moved away from the trolley, exceeding the trolley-chassis cable length, so that cable and / or connector gets damaged | 576777 | Measure | Behaviour with activated trolley-chassis cable tension detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563738 | Hazard | Contact with high level electrical energy (HS_002) | 576778 | Cause | System is moved away from the trolley, trolley-chassis cable tension detection is not working | 576779 | Measure | Self-test trolley-chassis cable tension detection
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563623 | Hazard | Short-term delay of the clinical procedure (HS_092) | 563624 | Cause | Collision is detected, the system is stopped, but the operator is not aware of the situation and is confused about the stop | 563626 | Measure | Collision state indication
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 563743 | Hazard | Severe limitation of movement functionality (HS_111) | 563744 | Cause | Collision sensor fails during any other operation (direction movement, standstill,Ã¢Â€Â¦): permanently activated collision sensor | 563745 | Measure | No additional measure required
563438 | Chapter | Move main unit | 563475 | Function | Move main unit - motorized, general | 576792 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 576793 | Cause | Regulatory source to skin distance is decreased when collision detection mechanism of X-ray source cover is activated | 576794 | Measure | No measure required
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 564224 | Measure | Pairing of hand control and system
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 564225 | Measure | Optical indication of system and hand control pairs
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 596993 | Measure | Label wireless hand control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 600715 | Measure | Optical indication of system and hand controls - collision
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564223 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates other system (e. g. in another room) | 600716 | Measure | Labeling wireless hand control - collision
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564234 | Cause | Control is dropped and function gets activated | 564235 | Measure | Reference - Control elements protected against unintended movement activation
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 564238 | Measure | Deactivate movement and radiation at wireless hand control receiver
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 595749 | Measure | Deactivate movement and radiation at wired hand control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 564239 | Measure | Indication of e-stop functionality
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 571957 | Measure | User message about non-functional wireless hand control and instruction to put it in the drawer
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 595744 | Measure | User message if wired hand control is non-functional
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 564240 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 564241 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564237 | Cause | Emergency stop on the wireless or wired hand control is not functional during a hazardous situation | 564242 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 578630 | Cause | Emergency stop on the wireless hand control is not functional when control goes to sleep mode | 578632 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564246 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 564247 | Measure | Reference - Deactivate movement and radiation at wireless hand control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564246 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 564248 | Measure | Reference - Alternative deactivation of emergency stop - hand control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568773 | Cause | Error in wireless data transmission: Delay, loss of data packet(s) | 568776 | Measure | Reference - Deactivate movement and radiation at wireless hand control receiver
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568774 | Cause | Error in wireless data transmission: Wrong sequence of data packets, repetition of invalid packets | 568779 | Measure | Reference - Safe wireless transfer of safety relevant information
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568775 | Cause | Error in wireless data transmission: Wrong information transmitted | 564349 | Measure | Safe wireless transfer of safety relevant information
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 576770 | Cause | Error in wireless data transmission: Sleeping error in protection mechanisms of wireless transfer | 576771 | Measure | Self-test safe wireless transfer components
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 576770 | Cause | Error in wireless data transmission: Sleeping error in protection mechanisms of wireless transfer | 578173 | Measure | Safe power supply for safe wireless transfer components
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568777 | Cause | Error in wireless data transmission: Internal error in receiver | 564351 | Measure | Safe state wireless hand control receiver
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568778 | Cause | Error in movement release circuit of the wireless control | 578324 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568778 | Cause | Error in movement release circuit of the wireless control | 578325 | Measure | Reference - Plausibility check for sensors in movement control elements
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564221 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568780 | Cause | Wireless control comes into reach or connect. reestablished and mov. button is activ. or tech. error in mov. release circuit | 568781 | Measure | Block movement when controls connecting
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 591782 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 591783 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates non-correctly calibrated control | 591784 | Measure | Reference - Pairing of hand control and system
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 591782 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 591783 | Cause | Remote controls are interchanged e. g. due to cleaning or moving the systems: Operates non-correctly calibrated control | 591785 | Measure | Reference - Optical indication of system and hand control pairs
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564227 | Cause | Wireless remote control is not functional: Battery runs empty | 564228 | Measure | Check battery charge of wireless remote control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564227 | Cause | Wireless remote control is not functional: Battery runs empty | 564229 | Measure | Wired remote control mandatory
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564230 | Cause | Wireless remote control is not functional: Liquid intrudes into the electronics | 564231 | Measure | IP protection class remote controls
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564230 | Cause | Wireless remote control is not functional: Liquid intrudes into the electronics | 568818 | Measure | Caution - Correct change of batteries wireless remote control
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564232 | Cause | Wireless remote control is not functional: Control is dropped and gets damaged | 564233 | Measure | Reference - Control elements protected against unintended movement activation
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564243 | Cause | Emergency stop on the wired / wireless remote control is activated unintentionally | 564244 | Measure | Prevent unintended activation of e-stop
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 564243 | Cause | Emergency stop on the wired / wireless remote control is activated unintentionally | 564245 | Measure | Indication of e-stop activation and location
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564226 | Hazard | Short-term delay of the clinical procedure (HS_092) | 591780 | Cause | Wireless remote control comes into wireless range or becomes functional again with a pressed emergency stop | 591781 | Measure | No additional measure required
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564249 | Hazard | Severe limitation of movement functionality (HS_111) | 564250 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 564252 | Measure | Reference - Deactivate movement and radiation at wireless hand control receiver
563438 | Chapter | Move main unit | 564220 | Function | Perform movements via wireless hand control | 564249 | Hazard | Severe limitation of movement functionality (HS_111) | 564250 | Cause | Emergency stop on the wired / wireless remote control is activated and control looses connection / gets unplugged | 564251 | Measure | Alternative deactivation of emergency stop - hand control
563438 | Chapter | Move main unit | 569231 | Function | Perform movements via wired hand control | 569232 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562466 | Cause | Disconnect wired hand control with activated emergency stop, system starts moving | 571978 | Measure | Reference - Alternative deactivation of emergency stop - hand control
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564079 | Cause | Control system does not brake / brakes too slowy (Quickstop not initiated or wrong braking ramp) | 564080 | Measure | Reference - Safe stop motorized movements
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564081 | Cause | After activating the emergency stop, movements stop, operator deactivates emergency stop, movement is initiated again | 564082 | Measure | No automatic movement release after e-stop deactivation
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562460 | Cause | Emergency stop not functional | 563483 | Measure | Emergency stop self-test and failure exclusions
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562460 | Cause | Emergency stop not functional | 563484 | Measure | Warning - Check emergency stop
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 562460 | Cause | Emergency stop not functional | 571900 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 564077 | Function | Activate emergency stop | 564078 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571898 | Cause | After activating the emergency stop, movement is initiated again, sleeping error in P path | 571899 | Measure | Reference - Self-test of protection module hardware
563438 | Chapter | Move main unit | 564083 | Function | Maintain stable system position | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564084 | Cause | System moves unintentionally during standstill | 564086 | Measure | Braking force
563438 | Chapter | Move main unit | 564083 | Function | Maintain stable system position | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564084 | Cause | System moves unintentionally during standstill | 564087 | Measure | Self-test of the brakes
563438 | Chapter | Move main unit | 564083 | Function | Maintain stable system position | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571856 | Cause | System moves unintentionally during brake test | 594854 | Measure | Reference protect unintentional activation rescue user
563438 | Chapter | Move main unit | 564083 | Function | Maintain stable system position | 564085 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 571856 | Cause | System moves unintentionally during brake test | 594855 | Measure | Reference - Inhibit motorized movements of chassis - rescue function
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564154 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 564155 | Measure | Positioning accuracy
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564156 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564158 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564156 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 599811 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564159 | Cause | Deviation of target position too large: External influences (slip, obstacles,Ã¢Â€Â¦) | 564160 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564162 | Measure | Reference - Movement preview for position movements
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564193 | Measure | Timeout for position movement selection
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 569216 | Measure | Movement range limit park and return
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564170 | Measure | Reference - Warning - Collisions during position movements without test run
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564192 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564163 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564165 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564166 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564167 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564161 | Cause | Operator has selected an incorrect target position from position storage or is not aware of selection | 564168 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564152 | Function | Perform position movement: Park and Return | 564153 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564171 | Cause | Operator initiates park and return on outdated / invalid position data | 564172 | Measure | Prerequisites park and return and position memory
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564200 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 564201 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564202 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564203 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564202 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 599812 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564205 | Cause | Deviation of target position too large: External influences (slip, obstacles,Ã¢Â€Â¦) | 564206 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564208 | Measure | Reference - Movement preview for automatic movements
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564209 | Measure | Reference - Timeout for automatic movement selection
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564219 | Measure | Movement range limit position memory
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564210 | Measure | Reference - Warning - Collisions during automatic movements without test run
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564211 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564212 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564213 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564214 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564215 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564207 | Cause | Operator has selected an incorrect target position from position storage | 564216 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564198 | Function | Perform position movement: Move to position from Position Memory (all motorized axles, simultaneously) | 564199 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564217 | Cause | Operator initiates movement to stored position on outdated / invalid position data | 564218 | Measure | Reference - Prerequisites park and return and position memory
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564175 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 564176 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564178 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 599813 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 594826 | Measure | Reference - no movement takeover across controls
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564177 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 594827 | Measure | Reference - only one active control element
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564179 | Cause | Deviation of target position too large: External influences (slip, obstacles,Ã¢Â€Â¦) | 564180 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564182 | Measure | Reference - Orientation indication for remote movement controls
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564183 | Measure | Reference - Warning - Collisions during automatic movements without test run
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564195 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564184 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564185 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564186 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564187 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 564173 | Function | Move along center beam (chassis and lift simultaniously) | 564174 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564181 | Cause | Operator has selected an incorrect movement function | 564188 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564312 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 564313 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564315 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 599814 | Measure | Reference - Positioning accuracy
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564316 | Measure | Reference - no movement takeover across controls
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564314 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 594828 | Measure | Reference - only one active control element
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564317 | Cause | Deviation of target position too large: External influences (slip, obstacles,Ã¢Â€Â¦) | 564318 | Measure | Reference - Main Unit movement shock detection
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564320 | Measure | Reference - Movement preview for automatic movements
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564321 | Measure | Reference - Timeout for position movement selection
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 569237 | Measure | Movement range limit click and go
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564322 | Measure | Reference - Warning - Collisions during automatic movements without test run
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564323 | Measure | Reference - Continuous activation of movements with redundant paths
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564324 | Measure | Reference - Set speed values
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564325 | Measure | Reference - Maximum overtravel values
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564326 | Measure | Reference - System speed limits
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564327 | Measure | Reference - Collision detection
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564319 | Cause | Operator has selected an incorrect target position on the image | 564328 | Measure | Reference - Emergency stop at every point of motorized movement activation
563438 | Chapter | Move main unit | 564310 | Function | Perform position movement: Click and Go (chassis and lift simultaneously) | 564311 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564329 | Cause | Operator initiates click and go on outdated / invalid position / image data | 564190 | Measure | Prerequisites click and go
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 564129 | Measure | Provide rescue button
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 562449 | Measure | Override angular / orbital brakes manually
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564128 | Cause | Chassis is not movable: Power supply, random hardware error (brakes are activated) | 564130 | Measure | Caution - Use rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 564132 | Measure | Reference - Mechanical design drive and wheels
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 564133 | Measure | Rotate chassis with blocked mecanum wheel
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 564140 | Measure | Reference - Provide rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 564134 | Measure | Reference - Override angular / orbital brakes manually
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564131 | Cause | Chassis is not movable: Mechanical failure of gears or wheels | 564135 | Measure | Reference - Caution - Use rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 564137 | Measure | Reference - Behaviour in collision
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 564139 | Measure | Reference - Collision detection
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 564141 | Measure | Reference - Provide rescue button
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 564142 | Measure | Reference - Override angular / orbital brakes manually
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564127 | Hazard | Patient rescue hindered (HS_140) | 564136 | Cause | Chassis movement is blocked: Collision sensor remains activated in collision, more than one collision sensor activated | 564143 | Measure | Reference - Caution - Use rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564150 | Cause | Technical failure in the rescue function activates the function and releases the mecanum brakes, system rolls | 564151 | Measure | Prevent unintentional activation of rescue function by design
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564150 | Cause | Technical failure in the rescue function activates the function and releases the mecanum brakes, system rolls | 576703 | Measure | Check rescue functionality during planned maintenance
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 564146 | Measure | Prevent unintentional activation of rescue function by user error
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 564147 | Measure | Inhibit motorized movements of chassis - rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564145 | Cause | Operator misuses the rescue function intentionally or unintentionally | 564148 | Measure | Warning - Correct use of rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 576759 | Cause | Rescue function is activated during activated emergency stop | 576760 | Measure | Reference - Inhibit motorized movements of chassis - rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 564149 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 580834 | Cause | Operator misuses the rescue function intentionally or unintentionally | 580835 | Measure | Reference - Inhibit motorized movements of chassis - rescue function
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 576758 | Cause | Unexpected arbitrary movement | 576762 | Measure | Reference - Prevent unintentional activation of rescue function by user error
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 576761 | Cause | Rescue function is activated during collision - collision detection ""non-functional"" | 576763 | Measure | Reference - Prevent unintentional activation of rescue function by user error
563467 | Chapter | Rescue or transfer patient (e. g. CPR, relocate) | 564126 | Function | Rescue or transfer patient (e. g. CPR, relocate) | 576757 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 576761 | Cause | Rescue function is activated during collision - collision detection ""non-functional"" | 576764 | Measure | Reference - Warning - Correct use of rescue function
569262 | Chapter | Acquire images | 563444 | Function | Acquire images - general | 580656 | Hazard | Unnecessary radiation, dose not as desired | 580654 | Cause | Radiation is switched off later than required in regulatory limits (IEC 60601-2-54:2018), normal conditions | 580655 | Measure | No additional measures required
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 568783 | Cause | Wireless control comes into reach and radiation button is activated or technical error in radiation release circuit | 562292 | Measure | Block radiation when controls connecting
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 596992 | Measure | Labeling wireless hand control radiation
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 596634 | Measure | Optical indication of system and hand controls radiation
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 600718 | Measure | Reference pairing wireless hand control
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 600719 | Measure | Reference Optical indication of system and hand control pairs
569262 | Chapter | Acquire images | 569260 | Function | Acquire images via wireless hand control | 568782 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 600717 | Cause | User mistakenly uses a wireless hand control from another system | 600720 | Measure | Reference Label wireless hand control
569262 | Chapter | Acquire images | 569279 | Function | Acquire images via wired footswitch | 569277 | Hazard | Medium-term delay of the clinical procedure (HS_091) | 569278 | Cause | Footswitch gets damaged due to ingress of fluid | 568819 | Measure | IP protection class footswitch
569262 | Chapter | Acquire images | 563447 | Function | Acquire Roadmap images | 591776 | Hazard | Unnecessary exposure to limited levels of ionizing radiation (HS_061) | 591777 | Cause | Position mismatch between subtraction and roadmap phases | 591778 | Measure | No system movement during or between subtraction and roadmap phases
569262 | Chapter | Acquire images | 563447 | Function | Acquire Roadmap images | 591776 | Hazard | Unnecessary exposure to limited levels of ionizing radiation (HS_061) | 591777 | Cause | Position mismatch between subtraction and roadmap phases | 591779 | Measure | Warning - Prevent movements during roadmap workflow
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564090 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 564091 | Measure | Reference - Braking force
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564090 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 564092 | Measure | Reference - Automatic brake test (selftest)
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564102 | Cause | Operator moves chassis (motorized) or other motorized axle during collision check or shortly after collision check | 564103 | Measure | Repeat collision check
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564089 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 564105 | Cause | Other motorized movement is triggered by technical fault during 3D scan | 564106 | Measure | Block other axles during 3D scan
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564097 | Hazard | Severe limitation of the imaging functionality (HS_100) | 564098 | Cause | Collision sensor fails or is activated during 3D scan, scan is aborted, restart possible | 564101 | Measure | Collision detection during 3D
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564093 | Hazard | Severely reduced image quality (HS_120) | 564094 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 564095 | Measure | Reference - Braking force
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 564093 | Hazard | Severely reduced image quality (HS_120) | 564094 | Cause | System moves unintentionally during 3D (brakes too weak for inertial force, brake not functional) | 564096 | Measure | Reference - Automatic brake test (selftest)
569262 | Chapter | Acquire images | 563448 | Function | Acquire 3D images | 578312 | Hazard | Unnecessary exposure to low levels of ionizing radiation (HS_060) | 578313 | Cause | Movement stops during acquisition | 578315 | Measure | Stop radiation when movement is stopped
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564272 | Hazard | Severely reduced image quality (HS_120) | 564273 | Cause | Wrong stiching of images: Increased mechanical tolerances, compromised position resolution, vibrations (systematic) | 564274 | Measure | Trajectory accuracy
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564272 | Hazard | Severely reduced image quality (HS_120) | 564273 | Cause | Wrong stiching of images: Increased mechanical tolerances, compromised position resolution, vibrations (systematic) | 564291 | Measure | Reduce vibrations during panorama
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564272 | Hazard | Severely reduced image quality (HS_120) | 564275 | Cause | Wrong stiching of images: Source to object information is not known | 564276 | Measure | No additional measure required
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564272 | Hazard | Severely reduced image quality (HS_120) | 564285 | Cause | Wrong synchronization at panorama imaging: Technical error | 564286 | Measure | Reference - Main Unit movement shock detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564272 | Hazard | Severely reduced image quality (HS_120) | 564288 | Cause | Wrong stiching of images: Software error | 564289 | Measure | Save all images panorama
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 564278 | Cause | Operator initates panorama imaging on any (outdated) image data | 564280 | Measure | Prerequisites panorama imaging
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569290 | Cause | Vibrations during scan loosens connectors | 596956 | Measure | Reference - class_shock_vibration_operation
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 564281 | Cause | Wrong synchronization at panorama imaging: Technical error | 564283 | Measure | Reference - Main Unit movement shock detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 564282 | Cause | Wrong synchronization at panorama imaging: External influences (slip, obstacles,...) | 564290 | Measure | Reference - Main Unit movement shock detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 596954 | Measure | Reference - movement check after service/maintenance
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 596955 | Measure | Reference - check_movement_manufacturing
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 569288 | Measure | Reference - Main Unit movement shock detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 569289 | Measure | Reference - Mechanical design drive and wheels
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564277 | Hazard | Severe limitation of the imaging functionality (HS_100) | 569287 | Cause | Undetected error in mecanum mechanics ( e. g. one castor is damaged) | 599815 | Measure | Reference - Positioning accuracy
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564294 | Measure | Reference - Movement preview for automatic movements
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 568854 | Measure | Restricted movement range for panorama imaging
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564301 | Measure | Reference - Timeout for automatic movement selection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564295 | Measure | Reference - Set speed values
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564296 | Measure | Reference - Maximum overtravel values
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564297 | Measure | Reference - System speed limits
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564298 | Measure | Reference - Collision detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564299 | Measure | Reference - Emergency stop at every point of motorized movement activation
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564293 | Cause | User error, planning of movement path is not done correctly (mix-up of directionsÃ¢Â€Â¦) | 564300 | Measure | Reference - Warning - Collisions during automatic movements without test run
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564302 | Cause | Deviation of target position too large: Tolerances too large, compromised resolution (systematic) | 564303 | Measure | Reference - Positioning accuracy
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564305 | Measure | Reference - Main Unit movement shock detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 599816 | Measure | Reference - Positioning accuracy
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 564306 | Measure | Reference - no movement takeover across controls
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 564304 | Cause | Deviation of target position too large: Technical error (wrong axle positions, wrong direction, wrong trajectory) | 594829 | Measure | Reference - only one active control element
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 577223 | Measure | Reference - Set speed values
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 577225 | Measure | Reference - Collision detection
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 577226 | Measure | Reference - System speed limits
569262 | Chapter | Acquire images | 564271 | Function | Acquire Panorama images | 564292 | Hazard | Crushing by or collision with objects moving with limited kinetic energy and / or limited contact force or pressure (HS_041) | 569228 | Cause | Emergency stop too far away when panorama image is acquired via footswitch | 577224 | Measure | Reference - Emergency stop at every point of motorized movement activation
563461 | Chapter | Display acquired image(s) | 569299 | Function | Display acquired image(s) | 569300 | Hazard | Slightly reduced image quality (HS_121) | 569301 | Cause | Impaired displayed image quality due to dirt on the touch monitor | 569148 | Measure | Reference - Cleaning of the touch monitor
563436 | Chapter | Save images | 569280 | Function | Save images | 569282 | Hazard | Image data is unavailable or lost (HS_130) | 569283 | Cause | Storage rate of images during acquisition is reduced | 594856 | Measure | Reference - full storage rate radiography
563463 | Chapter | Perform 3D navigation | 569284 | Function | Perform 3D navigation | 569285 | Hazard | Contact with high level electrical energy (HS_002) | 569286 | Cause | Damaged 3rd party devices are connected to the system | 568831 | Measure | Electrical safety external interfaces
563463 | Chapter | Perform 3D navigation | 569284 | Function | Perform 3D navigation | 569285 | Hazard | Contact with high level electrical energy (HS_002) | 569286 | Cause | Damaged 3rd party devices are connected to the system | 568833 | Measure | Warning - Electrical safety USB
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 568824 | Measure | Caution - Close mecanum cover after cleaning
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 571852 | Measure | Reference - Lock all motorized movements
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568822 | Cause | Cleaning cover of mecanum wheels is not closed after cleaning | 580832 | Measure | Label - Hand injury mecanum
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568823 | Cause | Unintended movement activation during cleaning activities of the mecanum wheels | 578106 | Measure | Reference - Warning - Close mecanum cover after cleaning
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568821 | Hazard | Crushing by or collision with objects moving with high kinetic energy and / or high contact force or pressure (HS_042) | 568823 | Cause | Unintended movement activation during cleaning activities of the mecanum wheels | 578107 | Measure | Reference - Lock all motorized movements
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568826 | Hazard | Contact with high level electrical energy (HS_002) | 568827 | Cause | Contact with electrical energy during mecanum wheel cleaning (e. g. ingress of cleaning agent) | 568828 | Measure | Protection against ingress of fluids - mecanum
563468 | Chapter | Clean the system | 568820 | Function | Clean the mecanum wheels | 568826 | Hazard | Contact with high level electrical energy (HS_002) | 568827 | Cause | Contact with electrical energy during mecanum wheel cleaning (e. g. ingress of cleaning agent) | 578318 | Measure | Reference - IEC 60601-1 Clause 8.7 and 8.8
569267 | Chapter | Transport the system | 569269 | Function | Transport the system | 569268 | Hazard | Complete loss of movement functionality (e. g. chassis) (HS_110) | 569270 | Cause | System gets damaged due to transport (e. g. vibrations, rough handling) | 568837 | Measure | No additional measure required
569267 | Chapter | Transport the system | 569269 | Function | Transport the system | 569271 | Hazard | Nonrecoverable or long-term loss of imaging functionality (HS_090) | 569272 | Cause | System gets damaged due to transport (e. g. vibrations, rough handling) | 569273 | Measure | Reference - Classification for mechanical shock and vibration - transport
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561954 | Hazard | Pinching through breakage of supporting parts | 561955 | Cause | Breakage of welded joints or wheels | 561956 | Measure | Design of the welded joints
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 561961 | Cause | Brake failure | 561962 | Measure | User Manual - Brake failure
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 561963 | Cause | Braking power not sufficient | 561964 | Measure | Reference - Design of the brakes
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 561963 | Cause | Braking power not sufficient | 561965 | Measure | User Manual - Parking Trolley on inclined Plane (i.e. Ramp)
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561960 | Hazard | Crushing through autonomous movements e.g. on inclined plane (i.e. ramp) | 561963 | Cause | Braking power not sufficient | 561966 | Measure | Warning Label - Parking Position Trolley
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561967 | Hazard | Risk of injury through damaged components | 561968 | Cause | Crossing a step | 561969 | Measure | Reference - Rough handling test trolley
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561970 | Hazard | Crushing when adjusting heights of monitors | 561971 | Cause | Motor-driven movement | 561972 | Measure | Distance in case of fault according to IEC 60601-1
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561970 | Hazard | Crushing when adjusting heights of monitors | 561971 | Cause | Motor-driven movement | 561973 | Measure | Attention note in user manual
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561974 | Hazard | Electric shock | 561975 | Cause | Protective earth broken | 597040 | Measure | Reference Cable protection
561939 | Chapter | Setup Trolley | 561940 | Function | Move Trolley into OR | 561980 | Hazard | Injury through dropping parts | 561981 | Cause | During transportation not secured parts may drop | 561982 | Measure | Built-in parts have to be secured because of the risk for dropping
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561984 | Hazard | Delayed or no interventional decision | 561985 | Cause | Untrained users / Missing knowledge of the system | 561986 | Measure | User Manual - Trained Users
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561984 | Hazard | Delayed or no interventional decision | 561987 | Cause | Not intended use of system | 561988 | Measure | Reference - Operator Manual - Not intended use
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561989 | Hazard | Loss of functionality | 561990 | Cause | Complexity of the system (component failures), mains supply failure or voltage out of specification | 561991 | Measure | Reference to User Manual - Establish emergency procedures
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561989 | Hazard | Loss of functionality | 561992 | Cause | (Cooling) Liquid is stored in the trolley and leaks, damaging the electronics | 561993 | Measure | No fluid storage in trolley
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561989 | Hazard | Loss of functionality | 561994 | Cause | Foreseeable misuse | 561995 | Measure | Reference - OperatorManual-Trained Users
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561989 | Hazard | Loss of functionality | 561996 | Cause | Some parts in system could lead to a hazard by wear | 561997 | Measure | Reference - OperatorManual-Abrasion of wear
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 561999 | Cause | Overheating and Flammability Materials | 562000 | Measure | Reference - IEC 60601-1 Standards
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 561999 | Cause | Overheating and Flammability Materials | 562001 | Measure | Reference - IT equipment safety
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 561999 | Cause | Overheating and Flammability Materials | 562002 | Measure | Reference - User Manual - Switch off the system in the event of fire User Manual
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 562003 | Cause | Auto-ignition or Melting of Material | 562004 | Measure | No measure required
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 562005 | Cause | Explosion | 562006 | Measure | User Manual - Danger of explosion
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 561998 | Hazard | Hazards from fire or smoke | 562007 | Cause | Too high temperature of the detector | 562008 | Measure | Reference - System Warning - Temperature of the Detector
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562009 | Hazard | Chemical hazards | 562010 | Cause | Toxic materials and fluids | 562011 | Measure | Reference - Technical instruction for the disposal procedure
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562009 | Hazard | Chemical hazards | 562010 | Cause | Toxic materials and fluids | 562012 | Measure | Service Instruction - Toxic Fluid
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562013 | Hazard | electromagnetic disturbances | 562014 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 562016 | Measure | User Manual - Degradation of essential performance by electromagnetic interferences
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562013 | Hazard | electromagnetic disturbances | 562014 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 562017 | Measure | Reference to generator measure for IEC 60601-1-2
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562018 | Hazard | Hearing impairment of patients and/or staff | 562019 | Cause | Excessive audible noise | 562020 | Measure | No measure required
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562021 | Hazard | Injuries due to monitor dropping | 562022 | Cause | Breakage of the monitor support | 562023 | Measure | Compliance with DIN EN 60601-1
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562024 | Hazard | Risk of injury through impaired parts | 562025 | Cause | Transport damages e.g. through collision | 562026 | Measure | Protection of monitor during transport
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562024 | Hazard | Risk of injury through impaired parts | 562025 | Cause | Transport damages e.g. through collision | 596969 | Measure | Robust touch glass of monitor
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562028 | Cause | High bleeder resistor | 562029 | Measure | Leakage current measurement
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562028 | Cause | High bleeder resistor | 562030 | Measure | Work step in maintenance documents
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562028 | Cause | High bleeder resistor | 562031 | Measure | Attention note in service manual
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562032 | Cause | Interruption of the protective earth | 562033 | Measure | Protective earth measurement
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562032 | Cause | Interruption of the protective earth | 562034 | Measure | Attention note in service manual
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562032 | Cause | Interruption of the protective earth | 562035 | Measure | Work step in service instructions
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562036 | Cause | Mechanically impaired cables/wiring | 562037 | Measure | Cable conduit for movable cables
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562038 | Cause | Static charge | 562039 | Measure | Create conductance on the trolley and main unit
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562038 | Cause | Static charge | 562040 | Measure | Work step in maintenance documents
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 562042 | Measure | Reference to compliance to IEC 60601-1
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 562043 | Measure | Reference to Service Documentation - Trained service personnel
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562027 | Hazard | Electric shock | 562041 | Cause | Residual charge in components after power off or impaired electrical integrity of the system | 562044 | Measure | Caution in Service Documentation - Discharge capacitor
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562045 | Hazard | Risk of injury due to bad surface finish | 562046 | Cause | Sharp edges | 562047 | Measure | Develop design without sharp edges
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562045 | Hazard | Risk of injury due to bad surface finish | 562046 | Cause | Sharp edges | 562048 | Measure | Reference to measures
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562045 | Hazard | Risk of injury due to bad surface finish | 562049 | Cause | Rough surface finish | 562050 | Measure | Specify smooth surface finishes
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562052 | Hazard | Effect on people, e.g. tissue injury, reddened skin | 562053 | Cause | Electromagnetic radiation | 562054 | Measure | No measure required
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562055 | Hazard | Effect on people with implants e.g. pacemaker | 562056 | Cause | Interference into the implant\s electronics | 562058 | Measure | User Manual - Interference into Implants
561939 | Chapter | Setup Trolley | 561983 | Function | Startup Trolley | 562055 | Hazard | Effect on people with implants e.g. pacemaker | 562056 | Cause | Interference into the implant\s electronics | 576774 | Measure | Reference - EMC compliance - emissions and immunity and pass / fail criteria
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562065 | Hazard | Crushing due to system tipping over | 562066 | Cause | C-arm movement shifts over-all center of gravity | 562067 | Measure | Design of system
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562065 | Hazard | Crushing due to system tipping over | 562068 | Cause | Weight put on in addition changes over-all center of gravity | 562069 | Measure | User Manual - Weight Balancing
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562065 | Hazard | Crushing due to system tipping over | 562070 | Cause | Transport of the device over an inclined plane (i.e. ramp) | 562071 | Measure | Design of system
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562065 | Hazard | Crushing due to system tipping over | 562072 | Cause | Crossing a step | 562073 | Measure | Reference - Rough handling test (main unit)
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562074 | Hazard | Crushing through breakage | 562075 | Cause | Breakage of welded joints | 562076 | Measure | Reference - Design of the welded joints
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562077 | Hazard | Crushing through breakage of the wheels | 562078 | Cause | Carrying capacity of the wheels too small | 562079 | Measure | Reference - Design of the wheels
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562080 | Hazard | Crushing due to autonomous movements on inclined plane (i.e. ramp) | 562081 | Cause | Brake failure | 562082 | Measure | Reference - self test of the brakes
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562080 | Hazard | Crushing due to autonomous movements on inclined plane (i.e. ramp) | 562083 | Cause | Braking power not sufficient | 562084 | Measure | Reference - Specify sufficient braking power
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562085 | Hazard | Risk of injury through damaged components | 562086 | Cause | Moving the system over a doorstep, against a doorframe, wall or other equipment | 562087 | Measure | Rough handling test (main unit)
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562085 | Hazard | Risk of injury through damaged components | 562086 | Cause | Moving the system over a doorstep, against a doorframe, wall or other equipment | 594861 | Measure | Reference - steep ramp
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562090 | Hazard | Injury through damaged control element | 562091 | Cause | Control element falls on the floor during transportation | 562092 | Measure | Drop test
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562093 | Hazard | Biological and Chemical Hazards | 562094 | Cause | Detector is damaged, service personnel comes into contact with CsI | 562096 | Measure | Supply Chain Documentation - Toxic Materials
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562093 | Hazard | Biological and Chemical Hazards | 562094 | Cause | Detector is damaged, service personnel comes into contact with CsI | 562097 | Measure | Service Instruction - Toxic Materials
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562093 | Hazard | Biological and Chemical Hazards | 562098 | Cause | Detector not disposed correctly | 562099 | Measure | Service Instruction - Toxic Materials - Disposal
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562093 | Hazard | Biological and Chemical Hazards | 569205 | Cause | Detector is damaged, operator comes into contact with CsI | 562095 | Measure | User Manual - Detector Cover Broken
562063 | Chapter | Setup Main Unit | 562064 | Function | Move Main Unit in OR | 562093 | Hazard | Biological and Chemical Hazards | 569205 | Cause | Detector is damaged, operator comes into contact with CsI | 569206 | Measure | No additional measure required
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562102 | Cause | Damage of the cable | 562103 | Measure | Certified mains power cable
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562102 | Cause | Damage of the cable | 569245 | Measure | Robust trolley-chassis cable
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562102 | Cause | Damage of the cable | 562104 | Measure | Cable anchorage
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562102 | Cause | Damage of the cable | 562105 | Measure | User Manual - Damage of Cable
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562106 | Cause | Interruption of the protective earth | 562107 | Measure | Reference - Protective earth measurement
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562106 | Cause | Interruption of the protective earth | 562108 | Measure | Reference - Attention note in service instructions
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562109 | Cause | Excess current or short-circuit | 562110 | Measure | Fuse for the supply line
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562109 | Cause | Excess current or short-circuit | 562111 | Measure | Reference - Warning label
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562112 | Cause | Static charge | 562113 | Measure | Reference - Create conductance on the trolley and main unit
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562101 | Hazard | Electric shock | 562112 | Cause | Static charge | 562114 | Measure | Reference - Work step in maintenance documents
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562115 | Hazard | Risk to fall | 562116 | Cause | Tripping over cables | 562117 | Measure | User Manual - Installing Cables
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562119 | Cause | Overcharge of rechargeable batteries | 562120 | Measure | Provide protection from overcharge HICTouchCon
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562119 | Cause | Overcharge of rechargeable batteries | 562121 | Measure | Provide protection from overcharge HICSwitch
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562119 | Cause | Overcharge of rechargeable batteries | 562122 | Measure | Use approved batteries
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562123 | Cause | Battery Gas Emission during Charge | 562124 | Measure | Batteries with Temperature Sensors
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562125 | Cause | Overcharge of the capacitors | 562126 | Measure | Mounted under covers
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562125 | Cause | Overcharge of the capacitors | 562127 | Measure | Reference - Protection of the energy storage
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562118 | Hazard | Excessive temperature, fire, explosion | 562125 | Cause | Overcharge of the capacitors | 562128 | Measure | Reference - Rupture joint by Overcharge and pressure relief valve
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562129 | Hazard | Foreseeable missuse | 562130 | Cause | Misplugging of cables | 562131 | Measure | Design considering misplugging
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 562134 | Measure | Check radiation release signal status at startup
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 562135 | Measure | Emergency stop radiation
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 562136 | Measure | User Manual - Emergency Stop Button
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 562137 | Measure | User Manual - Shutdown switch
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562133 | Cause | Exposure release switch is pressed by mistake during start up | 562138 | Measure | Reference to radiation indication signal
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562139 | Cause | Radiation indicator fails | 568841 | Measure | Monitoring of radiation indicators
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562139 | Cause | Radiation indicator fails | 562140 | Measure | Plug with stablized connection
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562139 | Cause | Radiation indicator fails | 562141 | Measure | User Manual - Radiation Indicator
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562132 | Hazard | Unintentional Radiation | 562142 | Cause | Radiation indicator LED fails | 562143 | Measure | No measure required
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562145 | Cause | Prevent startup | 562146 | Measure | Reference - User Manual - Wear and tear
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562147 | Cause | Data loss because the System is not safely turned off | 562148 | Measure | User Manual - ShutDown during X-Ray
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562149 | Cause | System not in a proper condition e.g. error in a component | 562150 | Measure | Display system status
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562151 | Cause | User loads non-approved SW or add non approved HW to the system | 562152 | Measure | User Manual - Use of non-approved parts
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 562154 | Measure | User Manual - Impermissible or faulty manipulation from network
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 562155 | Measure | Reference to measures against unauthorized installation of non-approved software
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 569133 | Measure | Imaging system access restrictions
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 569135 | Measure | Disable autorun for removable media
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562153 | Cause | Impermissible or faulty manipulation of SW from network. | 569136 | Measure | Imaging system access restrictions for internet and BIOS
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562156 | Cause | Incorrect treatment | 562157 | Measure | Reference - User Manual - Trained Users
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562158 | Cause | Harddisk or file system defect | 562159 | Measure | User manual - No Archive Medium
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562160 | Cause | Electric power failure | 562161 | Measure | Reference - The imaging system performs a controlled shutdown
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562162 | Cause | Insufficient charge (power) of UPS | 562163 | Measure | No measure required
562063 | Chapter | Setup Main Unit | 562100 | Function | Connect Trolley and Main Unit | 562144 | Hazard | Delayed or no interventional decision | 562164 | Cause | Charging capacity of the Uninterruptible Power Supply worn out | 562165 | Measure | No measure required
562063 | Chapter | Setup Main Unit | 562166 | Function | Performing system relevant check | 562167 | Hazard | No hazard on system level | 562168 | Cause | None | 562169 | Measure | None
562063 | Chapter | Setup Main Unit | 562170 | Function | Performing quality control | 562171 | Hazard | No hazard on system level | 562172 | Cause | None | 562173 | Measure | None
562063 | Chapter | Setup Main Unit | 562174 | Function | Calibration | 562175 | Hazard | Unnecessary radiation | 562176 | Cause | Incorrect parameters | 562177 | Measure | Automatic procedure to perform FD calibration correctly
562063 | Chapter | Setup Main Unit | 562174 | Function | Calibration | 562175 | Hazard | Unnecessary radiation | 562178 | Cause | Operating error | 562179 | Measure | Admin Guide - Calibration Procedure
562063 | Chapter | Setup Main Unit | 562174 | Function | Calibration | 562180 | Hazard | Image artifacts - Wrong interventional decision possible | 562181 | Cause | Bad calibration | 562182 | Measure | Reference - Automatic procedure to perform calibration correctly
562063 | Chapter | Setup Main Unit | 562174 | Function | Calibration | 562180 | Hazard | Image artifacts - Wrong interventional decision possible | 562181 | Cause | Bad calibration | 562183 | Measure | Reference - User Manual - Calibration Procedure
562063 | Chapter | Setup Main Unit | 562174 | Function | Calibration | 562184 | Hazard | Unwanted radiation | 562185 | Cause | Person in room during calibration | 562186 | Measure | Service training
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562188 | Hazard | Danger for patient or user | 562189 | Cause | Wrong Intended Use | 562190 | Measure | User Manual - Intended use only
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562188 | Hazard | Danger for patient or user | 562191 | Cause | Crash of subsystems | 562192 | Measure | Architectural Design of Subsystems and Components
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562188 | Hazard | Danger for patient or user | 562191 | Cause | Crash of subsystems | 562193 | Measure | User Manual - Reset procedure for a freezed imaging SW
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562188 | Hazard | Danger for patient or user | 562191 | Cause | Crash of subsystems | 562194 | Measure | Block keyboard / mouse inputs during radiation
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562195 | Hazard | Loss of functionality | 562196 | Cause | Complexity of the system or line voltage below the specified threshold | 562197 | Measure | Reference to User Manual - Establish emergency procedures
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562195 | Hazard | Loss of functionality | 562198 | Cause | Foreseeable misuse | 562199 | Measure | Reference - User Manual - Trained Users
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562195 | Hazard | Loss of functionality | 562200 | Cause | Some parts in system could lead to a hazard by wear | 562201 | Measure | User Manual - Wear and tear
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562195 | Hazard | Loss of functionality | 562202 | Cause | Operational hazard due to missing knowledge of the system | 562203 | Measure | Provide training for technician
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562206 | Hazard | Danger to electric high voltage or line voltage | 562207 | Cause | Problems with protective ground | 562208 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562206 | Hazard | Danger to electric high voltage or line voltage | 562207 | Cause | Problems with protective ground | 562209 | Measure | Start up instruction
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562206 | Hazard | Danger to electric high voltage or line voltage | 562210 | Cause | Contact with Parts Carrying Voltage | 562211 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562206 | Hazard | Danger to electric high voltage or line voltage | 562210 | Cause | Contact with Parts Carrying Voltage | 562212 | Measure | Warning Label - Voltage is present
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562206 | Hazard | Danger to electric high voltage or line voltage | 562213 | Cause | Isolation defect | 562214 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562215 | Hazard | Hazards from fire or smoke | 562216 | Cause | Fire inside the system | 562217 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562215 | Hazard | Hazards from fire or smoke | 562216 | Cause | Fire inside the system | 562218 | Measure | User Manual - Switch off the system in the event of fire
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562221 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562222 | Measure | Surface temperature applied parts
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562225 | Measure | Temperature sensor self test
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562226 | Measure | System Warning - Temperature of the Detector
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562228 | Measure | User Manual - Hot tube
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562220 | Cause | Parts too warm | 562229 | Measure | ServiceDoc - Hot tube housing
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562230 | Cause | Leakage of hot oil from tube assembly | 562231 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562219 | Hazard | Thermal energy and high temperature | 562230 | Cause | Leakage of hot oil from tube assembly | 562232 | Measure | Pressure switch in tube
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562237 | Hazard | Patient is infected by a virus due to a bad disinfection, system is damaged | 562238 | Cause | Incorrect cleaning of system after each patient, liquids or blood enter the system | 562239 | Measure | Reference - Spillage of Liquids on System
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562237 | Hazard | Patient is infected by a virus due to a bad disinfection, system is damaged | 562238 | Cause | Incorrect cleaning of system after each patient, liquids or blood enter the system | 562240 | Measure | User Manual - Cleaning and draping
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562241 | Hazard | Environmental hazard | 562242 | Cause | Explosion | 562243 | Measure | Reference - User Manual - Danger of explosion
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562244 | Hazard | Chemical hazards | 562245 | Cause | Toxic materials and fluids | 562246 | Measure | Technical instruction for the disposal procedure
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562247 | Hazard | Electro magnetic disturbances | 562248 | Cause | Delayed or disturbed transmission of images, or other functional disturbance | 562249 | Measure | Reference - EMC compliance - emissions, immunity and pass / fail criteria
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562250 | Hazard | Danger by releasing of stored mechanical energy | 562251 | Cause | Rotation of c-arm when power fails | 562252 | Measure | User Manual - Rotation of c-arm
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562250 | Hazard | Danger by releasing of stored mechanical energy | 562253 | Cause | Cable c-arm falling down | 562254 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562250 | Hazard | Danger by releasing of stored mechanical energy | 562255 | Cause | Collimator comes loose | 562256 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562260 | Cause | Unintentional activation of control elements | 578758 | Measure | Reference - Emergency stop at every point of motorized movement activation
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562262 | Cause | Manual or motorized movements of any parts | 562263 | Measure | User Manual - Risk of Crushing
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562262 | Cause | Manual or motorized movements of any parts | 562264 | Measure | User Manual - Danger zones
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562262 | Cause | Manual or motorized movements of any parts | 562265 | Measure | Limit motorized orbital operating radius
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562262 | Cause | Manual or motorized movements of any parts | 562266 | Measure | Warning Label - Risk of crushing hand
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562267 | Cause | Motorized movements of lifting column | 562268 | Measure | Reference - User Manual - Risk of crushing
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562267 | Cause | Motorized movements of lifting column | 562269 | Measure | Reference - User Manual - Danger Zones
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562267 | Cause | Motorized movements of lifting column | 562270 | Measure | Reference - User Manual - Trained Users
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562267 | Cause | Motorized movements of lifting column | 562271 | Measure | Warning Label - Risk of crushing foot
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562259 | Hazard | Collision with operator, patient or equipment, crushing | 562267 | Cause | Motorized movements of lifting column | 562272 | Measure | Warning Label - Risk of crushing hand
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562273 | Hazard | Injuries because of falling parts | 562274 | Cause | Wrong calculated max load, unforeseeable load. | 562275 | Measure | Reference - Compliance with IEC 60601-1
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562276 | Hazard | Injury or crushing due to emergency stop out of order | 562277 | Cause | Defect in button or cable harness | 562279 | Measure | Check emergency stop
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562280 | Hazard | Operator fails to actuate emergency stop | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 578759 | Measure | Reference - Emergency stop at every point of motorized movement activation
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562280 | Hazard | Operator fails to actuate emergency stop | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 578760 | Measure | Reference - Warning - General risk of collisions
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562280 | Hazard | Operator fails to actuate emergency stop | 562281 | Cause | Emergency stop is not readily accessible and easily identifiable | 562284 | Measure | User Manual - Emergency stop location
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562285 | Hazard | Malfunction of the system | 562286 | Cause | Inappropriate accessory | 562287 | Measure | User Manual - Inappropriate accessory
562063 | Chapter | Setup Main Unit | 562187 | Function | System Operation | 562285 | Hazard | Malfunction of the system | 562288 | Cause | Any defect parts in the system | 562289 | Measure | User Manual - Defect of system components
562063 | Chapter | Setup Main Unit | 562295 | Function | Stop button | 562296 | Hazard | Emergency stop out of order | 562297 | Cause | Defect in button or cable harness | 562298 | Measure | Reference - Check emergency stop - maintenance
562063 | Chapter | Setup Main Unit | 562299 | Function | Clean system | 562300 | Hazard | Patient collision | 562301 | Cause | Swapped wireless footswitch | 562302 | Measure | Reference to measures regarding footswitch
562063 | Chapter | Setup Main Unit | 562299 | Function | Clean system | 562303 | Hazard | Unwanted radiation | 562304 | Cause | Swapped wireless footswitch | 562305 | Measure | Reference to measures regarding footswitch
562063 | Chapter | Setup Main Unit | 562306 | Function | Store patient data | 562307 | Hazard | Unnecessary minor exposition to radiation | 562308 | Cause | Image store does not keep pace with acquisition under error conditions | 562309 | Measure | Warning Message
562063 | Chapter | Setup Main Unit | 562306 | Function | Store patient data | 562307 | Hazard | Unnecessary minor exposition to radiation | 562310 | Cause | Storage Limit Reached (License) | 591790 | Measure | Reference - Storage capacity exhausted
562063 | Chapter | Setup Main Unit | 562306 | Function | Store patient data | 562307 | Hazard | Unnecessary minor exposition to radiation | 562310 | Cause | Storage Limit Reached (License) | 591791 | Measure | Reference - Display storage space
562063 | Chapter | Setup Main Unit | 562306 | Function | Store patient data | 562312 | Hazard | Wrong interventional decision | 562313 | Cause | Image quality cannot be improved after acquisition | 562314 | Measure | Store raw data
562315 | Chapter | Connect to Network | 562316 | Function | HIS/RIS | 562317 | Hazard | Non-harmonized configuration | 562318 | Cause | Incorrect Configuration of HIS/RIS | 562319 | Measure | User Manual - HIS/RIS Configuration
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562323 | Measure | Use of reliable data transfer standards
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562324 | Measure | Prompt by local deleting of data
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562325 | Measure | Data integrity of xray image transfer via NXS
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562326 | Measure | NXS protocol with heartbeat
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562327 | Measure | NXS protocol with heartbeat for extapp
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562328 | Measure | Additional information of xray image transfer via NXS
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 562329 | Measure | Additional information of xray image transfer via NXS for extapp
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 568849 | Measure | Image compression threshold
562315 | Chapter | Connect to Network | 562320 | Function | Transfer exam | 562321 | Hazard | Delayed or no interventional decision | 562322 | Cause | Network disturbance | 569294 | Measure | Reference - Image display delay time - external
562315 | Chapter | Connect to Network | 562330 | Function | Connectivity | 562331 | Hazard | No hazard on system level | 562332 | Cause | None | 562333 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562336 | Hazard | Impaired vision through laser-beam | 562337 | Cause | People looking into laser-beam | 562338 | Measure | Usage of laser class 1
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562347 | Cause | Unintended movement activation when wired control element is connected to the system | 578761 | Measure | Reference - Warning - General risk of collisions
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562347 | Cause | Unintended movement activation when wired control element is connected to the system | 578766 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562351 | Cause | Movement without intentional activation on remote console | 578762 | Measure | Reference - Warning - General risk of collisions
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562351 | Cause | Movement without intentional activation on remote console | 578767 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562354 | Cause | Movement does not stop | 578768 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562354 | Cause | Movement does not stop | 578320 | Measure | Reference - Continuous activation of movements with redundant paths
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562342 | Hazard | Collision with patient, operator or equipment | 562354 | Cause | Movement does not stop | 578763 | Measure | Reference - Warning - General risk of collisions
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562358 | Hazard | Electrical shock | 562359 | Cause | Protective earth issues | 562360 | Measure | Reference - Compliance with IEC 60601-1
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562358 | Hazard | Electrical shock | 562359 | Cause | Protective earth issues | 562361 | Measure | Reference - Start up instruction
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562358 | Hazard | Electrical shock | 562362 | Cause | Contact with parts carrying voltage. | 562363 | Measure | Reference - Compliance with IEC 60601-1
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562358 | Hazard | Electrical shock | 562364 | Cause | Defective insulation | 562365 | Measure | IEC 60601-1 Clause 8.7 and 8.8
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562358 | Hazard | Electrical shock | 562364 | Cause | Defective insulation | 562366 | Measure | IEC 60601-1 Clause 9.8
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562367 | Hazard | Burns | 562368 | Cause | Hot surfaces | 562369 | Measure | Reference - Compliance with IEC 60601-1
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562370 | Hazard | Injuries because of falling parts | 562371 | Cause | Inadequate layout or construction of mechanical system parts. | 562372 | Measure | Reference - Compliance with IEC 60601-1 Standards
562334 | Chapter | Perform Clinical Procedure | 562335 | Function | Position Patient | 562370 | Hazard | Injuries because of falling parts | 562371 | Cause | Inadequate layout or construction of mechanical system parts. | 562373 | Measure | Reference to mounting of the monoblock
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562376 | Cause | Unintentional activation of movement | 562380 | Measure | No additional measure required - Reference to set values motorized speed
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562381 | Cause | Malfunction of the control element | 578094 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562384 | Measure | Stop vertical lift
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562385 | Measure | Acoustic signal
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562386 | Measure | Reference - User Manual - Guide Openings
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562387 | Measure | Reference - User Manual - Risk of Crushing
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562388 | Measure | User Manual - Check Audio Signal
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562389 | Measure | Supply Chain Documentation - Assembly Crushing Zone Switch
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562383 | Cause | Falling below the safety distance during vertical movement | 562390 | Measure | User Manual - Meaning of Audio Signals
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562391 | Cause | Movement speed too high | 578771 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562394 | Cause | Overtravel (stopping distance) too long | 578326 | Measure | Reference - Maximum overtravel values
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562375 | Hazard | Crushing through motor-driven movement | 562394 | Cause | Overtravel (stopping distance) too long | 562396 | Measure | Force limit orbital / angular motorized movement
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562399 | Cause | Weight causes drive to lower autonomously | 562400 | Measure | Provide brake or self-locking of the drive
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562401 | Cause | C-arm not weight balanced | 562402 | Measure | Design C-arm weight balanced
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562403 | Cause | Additional weight changes weight balance | 562404 | Measure | User Manual - Unbalanced C-arm
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562405 | Cause | Weakening of magnetic force | 562406 | Measure | Service Instruction - periodic maintenance
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562407 | Cause | Angular / orbital brake not functional | 562408 | Measure | Reference to Caution Operator Manual - Check brakes
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562398 | Hazard | Crushing through arbitrary movement of individual axles | 562407 | Cause | Angular / orbital brake not functional | 562409 | Measure | Reference to weight balanced design
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562414 | Cause | Carrying capacity of the unit too small | 562415 | Measure | Design of the movement unit
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562416 | Cause | Motor-driven movement and collision with fixed parts | 562417 | Measure | Design of the movement unit collision
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562418 | Cause | Breakage or wear of the spindle nut | 562419 | Measure | Additional safety mechanism - lift column
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562418 | Cause | Breakage or wear of the spindle nut | 562420 | Measure | Activation of safety mechanism blocks lift drive
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562418 | Cause | Breakage or wear of the spindle nut | 562421 | Measure | Labeling of the safety equipment
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562413 | Hazard | Crushing through breakage | 562418 | Cause | Breakage or wear of the spindle nut | 562422 | Measure | User Manual - Spindle Nut
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 562424 | Cause | Overtravel / movement range exceeds range | 562425 | Measure | End stop
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 562426 | Cause | Automatic 3D-scan | 578319 | Measure | Reference - Continuous activation of movements with redundant signals
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 562428 | Cause | Weight of the device | 562429 | Measure | User Manual - Fast Braking
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562423 | Hazard | Risk of crushing or falling parts, impaired mechanical integrity | 562430 | Cause | Mounting RCP on OR table holder | 562431 | Measure | Design without crushing zones
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 562433 | Cause | Brake failure | 562434 | Measure | User Manual - Daily Inspection
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 562435 | Cause | Braking power not sufficient | 562436 | Measure | Specify sufficient braking power
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562432 | Hazard | Crushing due to arbitrary movements of the chassis on inclined plane (i.e. ramp) | 562435 | Cause | Braking power not sufficient | 562437 | Measure | User Manual - Parking Chassis on inclined Plane (i.e. Ramp)
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562439 | Hazard | Risk of crushing during movement execution | 562440 | Cause | Distance between operator rail and other components too small | 562441 | Measure | Compliance with distances according to IEC 60601-1
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562439 | Hazard | Risk of crushing during movement execution | 562440 | Cause | Distance between operator rail and other components too small | 562442 | Measure | User Manual - Distance Operator Rail
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562439 | Hazard | Risk of crushing during movement execution | 562440 | Cause | Distance between operator rail and other components too small | 562443 | Measure | User Manual - Appropriate Distance
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562439 | Hazard | Risk of crushing during movement execution | 562440 | Cause | Distance between operator rail and other components too small | 562444 | Measure | User Manual - Guide Openings
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562439 | Hazard | Risk of crushing during movement execution | 562445 | Cause | Activation of movement without intervisibility | 562446 | Measure | User Manual - Observe Patient/Persons during Movement
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562450 | Hazard | Injuries through OR-instruments | 562451 | Cause | C-arm movements during operations | 562452 | Measure | User Manual - Collision with OR Instruments
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562453 | Hazard | Misjudgment due to poor image quality | 562454 | Cause | calibration incorrect | 562455 | Measure | Calibration cycles
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562456 | Hazard | Injury through dropping parts | 562457 | Cause | Collision with other systems | 562458 | Measure | User Manual - Collisions due to misuse
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562468 | Hazard | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) | 562469 | Cause | Restricted movement functionality (e. g. complete loss of vertical, angular orbital movement) | 562470 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562374 | Function | Position System | 562471 | Hazard | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) | 562472 | Cause | Slightly restricted movement functionality (e. g. complete loss of horizontal or swivel movement) | 562473 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562474 | Function | Wrap c-arm with Sterile Cover | 562475 | Hazard | Risk of infection | 562476 | Cause | Improperly installed sterile drape | 562477 | Measure | Reference to User Manual - Cleaning and draping
562334 | Chapter | Perform Clinical Procedure | 562478 | Function | Wrap Foot Switch with Covers | 562479 | Hazard | Risk of infection | 562480 | Cause | Improperly installed sterile drape | 562481 | Measure | Reference to User Manual - Cleaning and draping
562334 | Chapter | Perform Clinical Procedure | 562482 | Function | Wrap RCP with Covers | 562483 | Hazard | Risk of infection | 562484 | Cause | Damage of sterile drape | 562485 | Measure | Avoid sharp edges
562334 | Chapter | Perform Clinical Procedure | 562482 | Function | Wrap RCP with Covers | 562483 | Hazard | Risk of infection | 562486 | Cause | Inappropriate positioning of RCP | 562487 | Measure | User Manual - Positioning of RCP
562334 | Chapter | Perform Clinical Procedure | 562482 | Function | Wrap RCP with Covers | 562483 | Hazard | Risk of infection | 562488 | Cause | Self-loosening of OR table holder | 562489 | Measure | Fixation on OR table rail system
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562492 | Cause | Increasing temperature due to extended use | 562493 | Measure | Provide Monobloc Housing
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562492 | Cause | Increasing temperature due to extended use | 562494 | Measure | Reference - Pressure switch in tube
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562492 | Cause | Increasing temperature due to extended use | 562495 | Measure | Temperature monitoring monoblock
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562497 | Cause | Oil leakage out of leaky monoblock | 562499 | Measure | Pressure Compensation Membrane
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562497 | Cause | Oil leakage out of leaky monoblock | 578321 | Measure | Reference - Temperature monitoring monoblock
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562497 | Cause | Oil leakage out of leaky monoblock | 562500 | Measure | Reference - Compliance with IEC 60601-1
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562497 | Cause | Oil leakage out of leaky monoblock | 562501 | Measure | Reference - Pressure switch in tube
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562502 | Cause | Error of cooling system | 562503 | Measure | Reference - Temperature Control
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562491 | Hazard | Risk of burnings | 562504 | Cause | Longer operation heats up detector | 562505 | Measure | Provide contact protection/access cover for detector
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562507 | Hazard | Risk of infection | 562508 | Cause | Oil leakage out of leaky monoblock | 562509 | Measure | Stability test
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562507 | Hazard | Risk of infection | 562508 | Cause | Oil leakage out of leaky monoblock | 562510 | Measure | Reference - Pressure switch in tube
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562507 | Hazard | Risk of infection | 562508 | Cause | Oil leakage out of leaky monoblock | 562511 | Measure | Reference - Temperature control
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562507 | Hazard | Risk of infection | 562508 | Cause | Oil leakage out of leaky monoblock | 562512 | Measure | Reference - Check leakage of monobloc
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562507 | Hazard | Risk of infection | 562513 | Cause | Cooling circuit leaky, cooling medium drips into OR-field | 562514 | Measure | Reference to Cooling Unit Leakage Check
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562515 | Hazard | Injury through expelled parts | 562516 | Cause | Breakage rotating anode | 562517 | Measure | Reference - Stability test
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562515 | Hazard | Injury through expelled parts | 562518 | Cause | Implosion | 562519 | Measure | Reference to stability test and constructive measures
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 562522 | Measure | Reference - Mounted under covers
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 562523 | Measure | Protection of the energy storage
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 562524 | Measure | Rupture joint by Overcharge
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 562525 | Measure | Pressure relief valve
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562521 | Cause | Overcharge of the capacitors for large energy storage / Elko | 562526 | Measure | Reference to monitoring of the intermediate circuit voltage
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562527 | Cause | Short-circuit of electrical components | 562528 | Measure | Protection against excessive current and voltage
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562527 | Cause | Short-circuit of electrical components | 562529 | Measure | Protection against excessive current and voltage - PowerUnit
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562520 | Hazard | Excessive temperature, fire, explosion | 562527 | Cause | Short-circuit of electrical components | 562530 | Measure | Reference to generator measures to UL listed materials and fuses in intermediate circuit
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562531 | Hazard | Implosion | 562532 | Cause | Collision/impact against component | 562533 | Measure | Mechanical strength of the X-ray generation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562531 | Hazard | Implosion | 562532 | Cause | Collision/impact against component | 562534 | Measure | Reference - Stability test
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562531 | Hazard | Implosion | 562535 | Cause | Breakage rotating anode | 562536 | Measure | Housing has to stay integral
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562531 | Hazard | Implosion | 562535 | Cause | Breakage rotating anode | 562537 | Measure | Reference - Stability test
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562531 | Hazard | Implosion | 562535 | Cause | Breakage rotating anode | 562538 | Measure | Reference to constructive measures
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562539 | Hazard | Risk of injury due to bad surface finish | 562540 | Cause | Sharp edges | 562541 | Measure | Develop design without sharp edges
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562539 | Hazard | Risk of injury due to bad surface finish | 562542 | Cause | Rough surface finish | 562543 | Measure | Reference - Specify smooth surface finishes
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562544 | Hazard | Injury due to damaged housing | 562545 | Cause | Transport damages e.g. collision | 562546 | Measure | User Manual - Professional Repair
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562547 | Hazard | Electric shock | 562548 | Cause | Intrusion of liquids | 562549 | Measure | Spillage of Liquids on System
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562547 | Hazard | Electric shock | 562548 | Cause | Intrusion of liquids | 562552 | Measure | User Manual - Draping of FD
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562554 | Cause | No radiation protection behind the entering side of radiation | 562555 | Measure | Compliance with DIN EN 60601-1-3
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562556 | Cause | Increased leakage radiation monoblock | 562557 | Measure | Reference to constructive measures and tests to comply with leakage radiation requirements
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562558 | Cause | Stay in endangered area | 562559 | Measure | User Manual - Unwanted Radiation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562560 | Cause | Accidental release of radiation | 562561 | Measure | Design of X-ray release controls
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562560 | Cause | Accidental release of radiation | 562562 | Measure | Display Radiation On
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562564 | Cause | Malfunction of exposure release circuit. | 562565 | Measure | Hardware signal
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562564 | Cause | Malfunction of exposure release circuit. | 580167 | Measure | Quasi-continuous activation of radiation with redundant activation paths
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562564 | Cause | Malfunction of exposure release circuit. | 562567 | Measure | Reference to generator safety circuit and test at start-up
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562564 | Cause | Malfunction of exposure release circuit. | 578773 | Measure | Reference - Emergency stop radiation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562568 | Cause | Loss of connection between FD and imaging system | 562569 | Measure | Reference - Imaging system hardware test at startup
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562570 | Cause | High radiation dose due to excessive procedure duration and / or heavily inadequate radiation settings without clinical benefi | 562571 | Measure | Reference to Warning - Trained Users
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562572 | Cause | Bad image quality of single image - make single image without having done flluoro before | 562573 | Measure | No additional measures required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562574 | Cause | Failed subtraction / roadmap due to operator error or configuration error (licensing,...) | 562575 | Measure | No additional measures required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562553 | Hazard | Unwanted radiation | 562576 | Cause | Radiation is not switched off or switched off too late (e. g. pulses) due to HW / SW errors | 562577 | Measure | Reference to measures for radiation switch off
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562579 | Cause | Image orientation mixed up due to image rotation | 562580 | Measure | User Manual - Image Orientation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562581 | Cause | Image orientation mixed up due to flipped image | 562582 | Measure | User Manual - Label each side
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562583 | Cause | Image orientation mixed up | 562584 | Measure | User Manual - Image Orientation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562586 | Cause | Wrong image orientation applied by imaging system | 562587 | Measure | Encapsulation for image flip and rotate
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562588 | Cause | Wrong calibration display | 562589 | Measure | Update calibration
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562588 | Cause | Wrong calibration display | 562590 | Measure | User Manual - 2D measurements
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562591 | Cause | Wrong interpretation of results after calibration | 562592 | Measure | Reference to Caution for 2D measurements
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562593 | Cause | Wrong mask image for subtraction | 562594 | Measure | Correct mask calculation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562595 | Cause | Movements (either extrinsic or intrinsic) or image manipulations (e.g. AAA line) | 562596 | Measure | User Manual - Mismatch of AAA Line
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562599 | Cause | Image orientation reset due to application change | 562600 | Measure | User Manual - Image orientation reset
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562601 | Cause | LIH is confused with live image | 562602 | Measure | Clearance of LIH during radiation
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562603 | Cause | Live image covered by messages or UI elements | 562604 | Measure | Keep live imaging in foreground
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562605 | Cause | Incorrect K-wire display | 562606 | Measure | User Manual - Incorrect K-wire display
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562607 | Cause | overlay hide K-wire on image | 562608 | Measure | User Manual - overlay hide K-wire on image
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562578 | Hazard | Wrong interventional decision | 562607 | Cause | overlay hide K-wire on image | 562609 | Measure | Visibility of top of K-wire
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562610 | Hazard | Loss of images | 562611 | Cause | Insufficient image storage space | 562612 | Measure | Storage capacity exhausted
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562610 | Hazard | Loss of images | 562611 | Cause | Insufficient image storage space | 562613 | Measure | Display storage space
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562610 | Hazard | Loss of images | 562614 | Cause | Images / series lost in imaging system | 569132 | Measure | Immediate image storage on hard disk
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562610 | Hazard | Loss of images | 562614 | Cause | Images / series lost in imaging system | 578317 | Measure | Full storage rate radiography
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562610 | Hazard | Loss of images | 562616 | Cause | Reducing the storage rate or being confused about replay | 562617 | Measure | Reference - Immediate image storage on hard disk.
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562618 | Hazard | Unwanted Injection | 562619 | Cause | Unwanted radiation release | 562620 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562618 | Hazard | Unwanted Injection | 562621 | Cause | Triggersignal failure | 562622 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 580664 | Hazard | Unintended injection | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 580666 | Measure | Reference - User Manual - Injector Readiness
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 580664 | Hazard | Unintended injection | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 581162 | Measure | Injector trigger in dedicated modes and conditions only
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 580664 | Hazard | Unintended injection | 580665 | Cause | Trigger e. g. fluoro with injector in ready state, not knowing that every X-ray mode triggers the injector | 581195 | Measure | Display when injector is armed
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562623 | Hazard | Radiation without injection | 562624 | Cause | Acquire image, but injector not prepared correctly | 562625 | Measure | User Manual - Injector Readiness
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562623 | Hazard | Radiation without injection | 581193 | Cause | Injector interface is damaged due to motorized platform movement with plugged injector connection | 581194 | Measure | Restricted platform movement range with plugged injector
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562623 | Hazard | Radiation without injection | 581193 | Cause | Injector interface is damaged due to motorized platform movement with plugged injector connection | 595189 | Measure | Unplug injector cable during movement
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562626 | Hazard | Radiation interrupted | 562627 | Cause | Exceeding maximum delay time | 568834 | Measure | Image display delay time
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562626 | Hazard | Radiation interrupted | 562627 | Cause | Exceeding maximum delay time | 568835 | Measure | Image display delay time - external
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562626 | Hazard | Radiation interrupted | 562627 | Cause | Exceeding maximum delay time | 596894 | Measure | Caution in user manual to image display delay
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562629 | Hazard | Restricted imaging functionality due to low-level arcing | 562630 | Cause | Arcing in the X-ray tube | 562631 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562632 | Hazard | Restricted imaging functionality due to medium-level arcing | 562633 | Cause | Arcing in the X-ray tube | 562634 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562635 | Hazard | Image quality is reduced | 562636 | Cause | FLC hardware is defective | 562637 | Measure | Imaging system hardware test at startup
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562635 | Hazard | Image quality is reduced | 562636 | Cause | FLC hardware is defective | 599844 | Measure | User Manual - High-Energy-Isotopes
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562635 | Hazard | Image quality is reduced | 562638 | Cause | Numerical calculation error or parameterization issue in imaging pipeline | 562639 | Measure | Consistency check in imaging pipeline
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562635 | Hazard | Image quality is reduced | 562640 | Cause | Detector heats up out of calibration range | 562641 | Measure | Reference - System Warning - Temperature of the Detector
562334 | Chapter | Perform Clinical Procedure | 562490 | Function | Acquire Image | 562635 | Hazard | Image quality is reduced | 562640 | Cause | Detector heats up out of calibration range | 562642 | Measure | Reference to Caution in Operator Manual - Detector temperature
562334 | Chapter | Perform Clinical Procedure | 562643 | Function | Patient browser | 562644 | Hazard | No hazard on system level | 562645 | Cause | None | 562646 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562647 | Function | Patient registration | 562648 | Hazard | Misinterpretation of patient data | 562649 | Cause | Patient-ID already in use | 562650 | Measure | User Manual - Patient ID
562334 | Chapter | Perform Clinical Procedure | 562647 | Function | Patient registration | 562651 | Hazard | Emergency login not possible | 562652 | Cause | Authenticated login, access not possible | 562653 | Measure | Emergency login without password protection
562334 | Chapter | Perform Clinical Procedure | 562654 | Function | APS management | 562655 | Hazard | No hazard on system level | 562656 | Cause | None | 562657 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562661 | Measure | Reliable data transfer standards and protocols
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562662 | Measure | NGS protocol with speed detection
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562663 | Measure | NGS protocol with link loss detection
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562664 | Measure | NGS protocol with latency detection
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562665 | Measure | NGS protocol with update rate detection
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 562666 | Measure | NGS protocol with data integrity
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 569295 | Measure | Reference - Image display delay time
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 569296 | Measure | Reference - Image display delay time - external
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562660 | Cause | Communication problems | 569238 | Measure | Caution - Reduced image quality streaming
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562659 | Hazard | Delayed or no interventional decision | 562667 | Cause | Lossy organprogram | 562668 | Measure | Reference - User Manual - Wrong exposure parameters
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562669 | Hazard | Unnecessary radiation | 562670 | Cause | Communication network disturbance -incorrect parameters | 562671 | Measure | Reference - Reliable data transfer standards and protocols
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562669 | Hazard | Unnecessary radiation | 562672 | Cause | Operating error | 562673 | Measure | Reference - User Manual - Trained Users
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562669 | Hazard | Unnecessary radiation | 562672 | Cause | Operating error | 568855 | Measure | Indication of loading factors
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562674 | Hazard | Dose not as desired | 562675 | Cause | Incorrect parameters due to communication/network disturbance | 562676 | Measure | Reference - Ready for radiation
562334 | Chapter | Perform Clinical Procedure | 562658 | Function | Administration, selection and changing of applications | 562674 | Hazard | Dose not as desired | 562675 | Cause | Incorrect parameters due to communication/network disturbance | 562677 | Measure | Reference - Reliable data transfer standards and protocols
562334 | Chapter | Perform Clinical Procedure | 562678 | Function | Display of patient relevant data on user interface | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 562681 | Measure | Patient Data Display
562334 | Chapter | Perform Clinical Procedure | 562678 | Function | Display of patient relevant data on user interface | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 562682 | Measure | Reference to measures for patient or image data refresh and data consistency
562334 | Chapter | Perform Clinical Procedure | 562678 | Function | Display of patient relevant data on user interface | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 569126 | Measure | Refresh display after patient or image data correction
562334 | Chapter | Perform Clinical Procedure | 562678 | Function | Display of patient relevant data on user interface | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 569127 | Measure | Imaging system data consistency
562334 | Chapter | Perform Clinical Procedure | 562678 | Function | Display of patient relevant data on user interface | 562679 | Hazard | Mismatch of patient data, incorrect treatment | 562680 | Cause | Disturbance in processing, incorrect patient data displayed | 569128 | Measure | Transfer images of current patient only
562334 | Chapter | Perform Clinical Procedure | 562683 | Function | Process patient exam list | 562684 | Hazard | No hazard on system level | 562685 | Cause | None | 562686 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562687 | Function | Presetting of collimator position | 562688 | Hazard | Unnecessary radiation | 562689 | Cause | Communication network disturbance -incorrect parameters | 562690 | Measure | Reference - Reliable data transfer standards and protocols
562334 | Chapter | Perform Clinical Procedure | 562687 | Function | Presetting of collimator position | 562688 | Hazard | Unnecessary radiation | 562691 | Cause | Use Presettings of Collimator Position for Automatic Virtual Shutter | 562692 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562693 | Function | Exposure field | 562694 | Hazard | Unnecessary radiation | 562695 | Cause | Position of collimator leaves out of control due to stepping errors | 562696 | Measure | Perform Calibration of Collimator Leaves
562334 | Chapter | Perform Clinical Procedure | 562697 | Function | Wireless Remote control for movement functions | 562698 | Hazard | Injury by crushing | 562707 | Cause | Other wireless devices act as remote control | 562708 | Measure | Remote control and system are paired
562334 | Chapter | Perform Clinical Procedure | 562697 | Function | Wireless Remote control for movement functions | 562709 | Hazard | Interfere with other devices | 562710 | Cause | Electromagnetic field disturbs life-supporting devices | 562711 | Measure | Reference to Operator Manual - Interference with life-supporting devices or critical equipment
562334 | Chapter | Perform Clinical Procedure | 562697 | Function | Wireless Remote control for movement functions | 562709 | Hazard | Interfere with other devices | 562710 | Cause | Electromagnetic field disturbs life-supporting devices | 562712 | Measure | Reference - EMC compliance - emissions, immunity and pass / fail criteria
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562715 | Cause | Dose controlling | 562716 | Measure | IQAP measurement
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562715 | Cause | Dose controlling | 562717 | Measure | Switch off radiation
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 562719 | Measure | Admin Guide - System calibration
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 569225 | Measure | Reference - Ready for radiation
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 562721 | Measure | Display clinical procedure
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 562722 | Measure | User Manual - Wrong exposure parameters
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 562723 | Measure | User Manual - Reduced Image Quality due to Filtering
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562718 | Cause | Incorrect parameters | 562724 | Measure | Encapsulation of brightness value calculation
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562725 | Cause | Reduced image quality | 562726 | Measure | User Manual - System calibration
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562725 | Cause | Reduced image quality | 562727 | Measure | User Manual - Reduced Image Quality due to Filtering
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 562729 | Measure | Reference to measures to generator self-tests during start-up / initialisation and filtration test
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 578746 | Measure | Ready for radiation
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 562731 | Measure | Reference to measure to ensure changing application is possible by operator
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 562732 | Measure | Reference - User Manual - Wrong exposure parameters
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 562733 | Measure | Reference - User Manual - Reduced Image Quality due to Filtering
562334 | Chapter | Perform Clinical Procedure | 562713 | Function | Dose control | 562714 | Hazard | Unnecessary radiation, dose not as desired | 562728 | Cause | Incorrect parameters due to error in generator or tube | 562734 | Measure | Reference - Encapsulation of brightness value calculation
562334 | Chapter | Perform Clinical Procedure | 562735 | Function | Grid selection | 562736 | Hazard | Unnecessary radiation | 562737 | Cause | Operating error | 562738 | Measure | User Manual - Incorrect Grid Status
562334 | Chapter | Perform Clinical Procedure | 562735 | Function | Grid selection | 562739 | Hazard | Crushing between the housings of the detector case. | 562740 | Cause | Changing the grid | 562741 | Measure | User Manual - Changing the grid
562334 | Chapter | Perform Clinical Procedure | 562742 | Function | Store fluoro | 562743 | Hazard | Loss of images | 562744 | Cause | Insufficient image storage space | 562745 | Measure | Reference - Display storage space
562334 | Chapter | Perform Clinical Procedure | 562746 | Function | Position FOV | 562747 | Hazard | Dose not as desired | 562748 | Cause | Incorrect collimator setting | 562749 | Measure | Adjustment of the collimator plate
562334 | Chapter | Perform Clinical Procedure | 562746 | Function | Position FOV | 562747 | Hazard | Dose not as desired | 562748 | Cause | Incorrect collimator setting | 562750 | Measure | User Manual - Collimator Check
562334 | Chapter | Perform Clinical Procedure | 562746 | Function | Position FOV | 562747 | Hazard | Dose not as desired | 562748 | Cause | Incorrect collimator setting | 562751 | Measure | Adjustment of the collimator plate
562334 | Chapter | Perform Clinical Procedure | 562746 | Function | Position FOV | 562747 | Hazard | Dose not as desired | 562748 | Cause | Incorrect collimator setting | 562752 | Measure | Reference - Perform calibration of collimator leaves
562334 | Chapter | Perform Clinical Procedure | 562746 | Function | Position FOV | 562753 | Hazard | Eye injury (visual disturbance) | 562754 | Cause | Unintended look into the laser beam for a lengthy time | 562755 | Measure | Usage of laser class 1
562334 | Chapter | Perform Clinical Procedure | 562756 | Function | Zoom format change | 562757 | Hazard | Dose not as desired | 562758 | Cause | Error in FD zoom change | 562759 | Measure | Adjustment of the collimator plate
562334 | Chapter | Perform Clinical Procedure | 562760 | Function | Exchange grid | 562761 | Hazard | Dose not as desired | 562762 | Cause | Incorrect grid status, inserted, removed | 562763 | Measure | Detect grid status
562334 | Chapter | Perform Clinical Procedure | 562760 | Function | Exchange grid | 562761 | Hazard | Dose not as desired | 562762 | Cause | Incorrect grid status, inserted, removed | 562764 | Measure | User Manual - Incorrect Grid Status
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562766 | Hazard | Infection: A patient is infected by virus/bacteria due to bad disinfection of the system | 562767 | Cause | Insufficient cleaning of system between patients | 562768 | Measure | Reducing openings and gaps
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562766 | Hazard | Infection: A patient is infected by virus/bacteria due to bad disinfection of the system | 562767 | Cause | Insufficient cleaning of system between patients | 562769 | Measure | User Manual - Cleaning of Wheels
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562770 | Hazard | Delayed or no treatment | 562771 | Cause | Blood or other liquids damage system | 562772 | Measure | Reference - Spillage of Liquids on System
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562770 | Hazard | Delayed or no treatment | 562773 | Cause | Cleaning and disinfection substances damage system | 562774 | Measure | User Manual - Insufficient cleaning
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562775 | Hazard | Undesired radiation | 562776 | Cause | Cleaning person accidentally steps on footswitch | 562777 | Measure | User Manual - System switch-off
562334 | Chapter | Perform Clinical Procedure | 562765 | Function | Release patient | 562775 | Hazard | Undesired radiation | 562776 | Cause | Cleaning person accidentally steps on footswitch | 562778 | Measure | EmergencyStop
562334 | Chapter | Perform Clinical Procedure | 562779 | Function | Select patient from patient list | 562780 | Hazard | No hazard on system level | 562781 | Cause | None | 562782 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562783 | Function | Image display | 562784 | Hazard | Incorrect treatment basis | 562785 | Cause | Confusing of, mix-up of images/series from different patients. | 562786 | Measure | Always Display Patient Data
562334 | Chapter | Perform Clinical Procedure | 562783 | Function | Image display | 562784 | Hazard | Incorrect treatment basis | 562785 | Cause | Confusing of, mix-up of images/series from different patients. | 562787 | Measure | Correct image assignment
562334 | Chapter | Perform Clinical Procedure | 562783 | Function | Image display | 562784 | Hazard | Incorrect treatment basis | 562788 | Cause | Images are sorted or numbered incorrectly | 562789 | Measure | Image number storage in header
562334 | Chapter | Perform Clinical Procedure | 562790 | Function | Close Patient | 562791 | Hazard | No hazard on system level | 562792 | Cause | None | 562793 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 562796 | Cause | Erroneous release of stored energy (ESU) | 562797 | Measure | Electrical separation of accessible parts not protected by MOOP
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 562796 | Cause | Erroneous release of stored energy (ESU) | 562798 | Measure | Check of electrical separation functionality during manufacturing - generator
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 562795 | Hazard | Hazardous voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 562796 | Cause | Erroneous release of stored energy (ESU) | 562799 | Measure | Check of electrical separation functionality during regular maintenance
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 573395 | Measure | Reference - Electrical separation of accessible parts not protected by MOOP
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 573396 | Measure | Reference - Check of electrical separation functionality during manufacturing - generator
562334 | Chapter | Perform Clinical Procedure | 562794 | Function | Disconnect Trolley and Main Unit | 573393 | Hazard | Limited voltage present at the trolley-chassis connection after disconnecting trolley and main unit | 573394 | Cause | Erroneous release of stored energy (ELKO buffer) | 573397 | Measure | Reference - Check of electrical separation functionality during regular maintenance
562334 | Chapter | Perform Clinical Procedure | 562800 | Function | Quick Clean | 562801 | Hazard | Contamination through improper cleaning | 562802 | Cause | Usage of improper cleaning agents | 562803 | Measure | User Manual - Improper Cleaning Agents
562334 | Chapter | Perform Clinical Procedure | 562804 | Function | Retrieve and modify patient | 562805 | Hazard | No hazard on system level | 562806 | Cause | None | 562807 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562809 | Hazard | Loss of image data | 562810 | Cause | Unintentional deletion of images/series | 562811 | Measure | Reference - Prompt by local deleting of data
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562809 | Hazard | Loss of image data | 562810 | Cause | Unintentional deletion of images/series | 562812 | Measure | Security dialog for any deletion by the user
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562809 | Hazard | Loss of image data | 562813 | Cause | Destruction of the patient directory and loss of patient data and images | 562814 | Measure | File system of imaging system
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562809 | Hazard | Loss of image data | 562813 | Cause | Destruction of the patient directory and loss of patient data and images | 569130 | Measure | Imaging system recovery
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562815 | Hazard | Delayed or no interventional decision | 562816 | Cause | Unintended local deleting of patient data | 562817 | Measure | Reference - Prompt by local deleting of data
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562815 | Hazard | Delayed or no interventional decision | 562816 | Cause | Unintended local deleting of patient data | 562818 | Measure | Security dialog for any deletion by the user
562334 | Chapter | Perform Clinical Procedure | 562808 | Function | Image delete | 562815 | Hazard | Delayed or no interventional decision | 562816 | Cause | Unintended local deleting of patient data | 562819 | Measure | Storage commitment archive flag
562334 | Chapter | Perform Clinical Procedure | 562820 | Function | Calibrate detector | 562821 | Hazard | Image artifacts | 562822 | Cause | Bad calibration | 562823 | Measure | Automatic procedure to perform calibration correctly
562334 | Chapter | Perform Clinical Procedure | 562820 | Function | Calibrate detector | 562821 | Hazard | Image artifacts | 562822 | Cause | Bad calibration | 562824 | Measure | Reference - User Manual - Calibration Procedure
562334 | Chapter | Perform Clinical Procedure | 562820 | Function | Calibrate detector | 562821 | Hazard | Image artifacts | 562822 | Cause | Bad calibration | 562825 | Measure | Reference - User Manual - System calibration
562334 | Chapter | Perform Clinical Procedure | 562820 | Function | Calibrate detector | 562826 | Hazard | Unwanted radiation exposure | 562827 | Cause | Patient or other person in exam room during calibration | 562828 | Measure | Reference - Service training
562334 | Chapter | Perform Clinical Procedure | 562829 | Function | Stop button | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562832 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562833 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562834 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562835 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562836 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562837 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562838 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562839 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562840 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562841 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562842 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562843 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562844 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562845 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562846 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562847 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562848 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562849 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562850 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562851 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562852 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562853 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562854 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562855 | Function | Native language support | 562830 | Hazard | Unit stops too late in a critical situation | 562831 | Cause | In a panic situation, the user does not find the emergency stop switch quickly enough | 562856 | Measure | User Manual - Emergency Stop Location
562334 | Chapter | Perform Clinical Procedure | 562855 | Function | Native language support | 562852 | Hazard | Delayed or no interventional decision | 562853 | Cause | Network disturbance | 562854 | Measure | Reference - Use of reliable data transfer standards
562334 | Chapter | Perform Clinical Procedure | 562855 | Function | Native language support | 562852 | Hazard | Delayed or no interventional decision | 562853 | Cause | Network disturbance | 562855 | Measure | Reference - Prompt by local deleting of data
562334 | Chapter | Perform Clinical Procedure | 562856 | Function | Virus Protection | 562857 | Hazard | Delayed or no interventional decision | 562858 | Cause | System gehts infected by malware | 562865 | Measure | Intrusion detection solution
562334 | Chapter | Perform Clinical Procedure | 562856 | Function | Virus Protection | 562857 | Hazard | Delayed or no interventional decision | 562858 | Cause | System gehts infected by malware | 562866 | Measure | Reference - Provide SVPs regularly
562334 | Chapter | Perform Clinical Procedure | 562867 | Function | Help functionality | 562868 | Hazard | No hazard on system level | 562869 | Cause | None | 562870 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562871 | Function | Measure Area Dose Product | 562872 | Hazard | Unnecessary dose | 562873 | Cause | Wrong interpretation of applied dose, dose reporting affected | 562874 | Measure | Maintenance of the dose area product meter
562334 | Chapter | Perform Clinical Procedure | 562875 | Function | IHE Profiles | 562876 | Hazard | No hazard on system level | 562877 | Cause | None | 562878 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562879 | Function | Data Model | 562880 | Hazard | No hazard on system level | 562881 | Cause | None | 562882 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562883 | Function | Worklist | 562884 | Hazard | Patient data mixed up | 562885 | Cause | Mixing of data caused by RIS data transfer | 562886 | Measure | Reference - Imaging system data consistency
562334 | Chapter | Perform Clinical Procedure | 562883 | Function | Worklist | 562884 | Hazard | Patient data mixed up | 562885 | Cause | Mixing of data caused by RIS data transfer | 569129 | Measure | Check worklist data
562334 | Chapter | Perform Clinical Procedure | 562887 | Function | Storage | 562888 | Hazard | No hazard on system level | 562889 | Cause | None | 562890 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562891 | Function | Archive to offline media | 562892 | Hazard | No hazard on system level | 562893 | Cause | None | 562894 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562895 | Function | Query/retrieve | 562896 | Hazard | No hazard on system level | 562897 | Cause | None | 562898 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562899 | Function | MPPS | 562900 | Hazard | No hazard on system level | 562901 | Cause | None | 562902 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562903 | Function | DICOM dose reporting | 562904 | Hazard | Wrong dose reported by MPPS | 562905 | Cause | Misleading Information | 562906 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562907 | Function | Standard display function | 562908 | Hazard | No hazard on system level | 562909 | Cause | None | 562910 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562911 | Function | Verification | 562912 | Hazard | No hazard on system level | 562913 | Cause | None | 562914 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562915 | Function | Help functionality | 562916 | Hazard | No hazard on system level | 562917 | Cause | None | 562918 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562919 | Function | General requirements | 562920 | Hazard | No hazard on system level | 562921 | Cause | None | 562922 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 562925 | Cause | Loss of 2D images throughout the 3D scan | 562926 | Measure | Save all 2D images data set
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 562925 | Cause | Loss of 2D images throughout the 3D scan | 562927 | Measure | User Manual - Reduced accuracy of 3D reconstruction
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562924 | Hazard | Functional Hazards 3D Image Quality is reduced | 562928 | Cause | C-arm cable influences c-arm geometry | 562929 | Measure | No external C-arm cable
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562930 | Hazard | Patient risk for biological hazard | 562931 | Cause | C-arm cable moves into sterile area | 562932 | Measure | Reference - C-arm cable avoids sterile area for all positions
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562933 | Hazard | Risk of injury during motorized orbital movement | 562934 | Cause | Collision of C-arm with people and goods | 562935 | Measure | Reference to workflow
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562933 | Hazard | Risk of injury during motorized orbital movement | 562934 | Cause | Collision of C-arm with people and goods | 562936 | Measure | Motorized operation only after manual test drive
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562933 | Hazard | Risk of injury during motorized orbital movement | 562934 | Cause | Collision of C-arm with people and goods | 562937 | Measure | Reference - Continuous activation
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562933 | Hazard | Risk of injury during motorized orbital movement | 562934 | Cause | Collision of C-arm with people and goods | 562938 | Measure | Reference - Emergency stop at every point of motorized movement activation
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562933 | Hazard | Risk of injury during motorized orbital movement | 562934 | Cause | Collision of C-arm with people and goods | 562939 | Measure | Reference to set max. speed 3D scan
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562940 | Hazard | Patient cannot be salvaged | 562941 | Cause | Orbital drive is failing | 562942 | Measure | C-arm can be moved manually
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562943 | Hazard | Treatment on the wrong side | 562944 | Cause | Image orientation confused | 562945 | Measure | Reference to User Manual - Correct Position of Patient
562334 | Chapter | Perform Clinical Procedure | 562923 | Function | 3D Scan and acquisistion | 562946 | Hazard | 3D scan is unavailable | 562947 | Cause | Orbital drive fails or other error | 562948 | Measure | No measure required
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 562951 | Cause | Artifacts that might be mistaken for anatomical structures | 562952 | Measure | Artifacts Image Quality test for MPR
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 562951 | Cause | Artifacts that might be mistaken for anatomical structures | 562953 | Measure | Artifacts Maintenance recommendation in the user manual
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 562954 | Cause | Alignment error | 562955 | Measure | Alignment error Image Quality test for MPR
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562950 | Hazard | Multi Planar Reconstruction (MPR) | 562954 | Cause | Alignment error | 562956 | Measure | Reference - Artifacts Maintenance recommendation in the user manual
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562957 | Hazard | Wrong orientation of the reconstructed 3D volume | 562958 | Cause | Wrong input of patient orientation | 562959 | Measure | User Manual - Verify Patient Position and Orientation
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562957 | Hazard | Wrong orientation of the reconstructed 3D volume | 562958 | Cause | Wrong input of patient orientation | 562961 | Measure | Orientation of Image and P-matrices
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562962 | Hazard | 3D Imaging incorrect | 562963 | Cause | Acquired set of 2D image assigned to wrong P-matrix | 562964 | Measure | Assure correct assignment
562334 | Chapter | Perform Clinical Procedure | 562949 | Function | 3D Imaging | 562962 | Hazard | 3D Imaging incorrect | 562965 | Cause | Wrong assignment of P-matrices to scan protocol | 562966 | Measure | Correct assignment of scan config
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562968 | Hazard | Functional Hazards 3D Imaging insufficient | 562969 | Cause | Artifact level too high | 562970 | Measure | Reference - Artifacts Maintenance recommendation in the user manual
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562972 | Cause | Patient moves | 562973 | Measure | User Manual - patient movement during 3D
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562974 | Cause | System/C-armswings | 562975 | Measure | Ramp function
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562976 | Cause | System shifts during scan | 562977 | Measure | User Manual - System shift during 3D
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562978 | Cause | C-arm angulation does not correspond to calibration area | 562979 | Measure | Retrieval of C-arm angulation
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562980 | Cause | Image quality for image reconstuction impaired | 562981 | Measure | System calibration via calibration phantom
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562980 | Cause | Image quality for image reconstuction impaired | 562982 | Measure | Inspection of image quality
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 562984 | Measure | Statement about fidelity in user manual
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 562985 | Measure | Inspection of fidelity
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562983 | Cause | Insufficient geometric fidelity when using the viewer | 562986 | Measure | Encapsulation 3D image processing scale
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562987 | Cause | Test procedure for inspection of IQ insufficient | 562988 | Measure | Application of an appropriate test procedure
562334 | Chapter | Perform Clinical Procedure | 562967 | Function | 3D reconstruction | 562971 | Hazard | Impaired image acquisition | 562987 | Cause | Test procedure for inspection of IQ insufficient | 562989 | Measure | Encapsulation 3D handling of brightness / contrast related data
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 562991 | Hazard | Additional strain on patient | 562992 | Cause | Disturbance of data transfer | 562993 | Measure | Use of a defined protocol
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 562994 | Hazard | Risk of injury by C-arm | 562995 | Cause | Unintended shifting of the C-arm | 562996 | Measure | User Manual - Non-Counterbalanced C-arm during Navigation
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 562994 | Hazard | Risk of injury by C-arm | 562997 | Cause | Incorrectly mounted navigation components impair accuracy | 562998 | Measure | User Manual - Ensure Correct Mounting of Navigation System
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 562994 | Hazard | Risk of injury by C-arm | 562999 | Cause | Mechanical connection between detector housing and adapter fails | 563000 | Measure | Reference - Design fit for fourfold weight
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563001 | Hazard | Risk of injury during operation/transportation | 563002 | Cause | Tipping over of the device | 563003 | Measure | Reference - Design of system
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563001 | Hazard | Risk of injury during operation/transportation | 563002 | Cause | Tipping over of the device | 563004 | Measure | User Manual - Maximum Weight
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563005 | Hazard | Injury due to surface condition | 563006 | Cause | Sharp edges | 563007 | Measure | Reference - Compliance with IEC 60601-1 external
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563008 | Hazard | Risk of injury due to breakage | 563009 | Cause | Overloading of C-arm bearing and supporting elements for the detector | 563010 | Measure | Reference - Design of the movement unit
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563008 | Hazard | Risk of injury due to breakage | 563009 | Cause | Overloading of C-arm bearing and supporting elements for the detector | 563011 | Measure | User Manual - Break Resistance
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563012 | Hazard | Scattered radiation | 563013 | Cause | Unnecessary radiation exposure when additional objects in beam path. | 563014 | Measure | User Manual - Scattered Radiation
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563015 | Hazard | C-arm collides with people or goods | 563016 | Cause | Additional weight | 563017 | Measure | User Manual - Mounting of Additional Parts
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563019 | Hazard | Markers of different nav. Systems can fall off | 563020 | Cause | Insufficient dimensioning of the mechanical interface | 563021 | Measure | Navigation Interface - Marker Frame Fixation
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563019 | Hazard | Markers of different nav. Systems can fall off | 563022 | Cause | Reference ring loosenes | 563023 | Measure | Design fit for fourfold weight
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563019 | Hazard | Markers of different nav. Systems can fall off | 563024 | Cause | Markers insufficiently fixed | 563025 | Measure | Navigation Interface - Marker Fixation
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563026 | Hazard | Risk of injury during operation/transportation | 563027 | Cause | Tipping over of the device in general | 563028 | Measure | Reference - Design of system
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563026 | Hazard | Risk of injury during operation/transportation | 563027 | Cause | Tipping over of the device in general | 563029 | Measure | Reference - Navigation Interface - Maximum Weight
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563033 | Hazard | Wrong treatment due to insufficient precision | 563034 | Cause | Dirty/Damaged Mechanical Detector Interface | 563035 | Measure | User Manual - Visual Inspection/Cleaning
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563036 | Hazard | Risk of injury due to breakage | 563037 | Cause | Overloading of C-arm bearing and supporting elements for II | 563038 | Measure | Reference - Design of the movement unit
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563039 | Hazard | Electric schock | 563040 | Cause | Isolation error of nav. parts | 563041 | Measure | User Manual - Components Comply With IEC60601-1
562334 | Chapter | Perform Clinical Procedure | 562990 | Function | 2D/3D Navigation | 563039 | Hazard | Electric schock | 563040 | Cause | Isolation error of nav. parts | 563042 | Measure | Navigation Interface - Comply with IEC60601-1
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563044 | Hazard | Collision with catheters | 563045 | Cause | Operation error | 563046 | Measure | User Manual - Dry Run Manually
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563047 | Hazard | 3D navigation accuracy reduced | 563048 | Cause | Start position 3D scan | 563049 | Measure | The accuracy of the mechanical marker interface
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563047 | Hazard | 3D navigation accuracy reduced | 563048 | Cause | Start position 3D scan | 563050 | Measure | User Manual - Verify Accuracy
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563051 | Hazard | Wrong treatment due to insufficient precision | 563052 | Cause | Damage during transport | 563053 | Measure | Reference - User Manual - Visual Inspection/Cleaning
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563051 | Hazard | Wrong treatment due to insufficient precision | 563054 | Cause | Calibration-phantom interface damaged | 563055 | Measure | Reference - User Manual - Visual Inspection/Cleaning
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563051 | Hazard | Wrong treatment due to insufficient precision | 563056 | Cause | Recalibration for 3D-mode | 563057 | Measure | Admin Guide - Navigation system test after recalibration
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563051 | Hazard | Wrong treatment due to insufficient precision | 563056 | Cause | Recalibration for 3D-mode | 563058 | Measure | Service Documentation - Navigation system test after recalibration
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563051 | Hazard | Wrong treatment due to insufficient precision | 563059 | Cause | Faulty navigation system during calibration | 563060 | Measure | User Manual - Check Precision of Navigation
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563061 | Hazard | Wrong interventional decision or treatment | 563062 | Cause | Wrong selection by user | 563063 | Measure | User Manual - Choose Correct Navigation
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563064 | Hazard | Additional strain on patient | 563065 | Cause | Incorrect registration of the marker | 563066 | Measure | No scan is possible without confirmation
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563064 | Hazard | Additional strain on patient | 563065 | Cause | Incorrect registration of the marker | 563067 | Measure | Navigation Interface - Scan Confirmation
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563064 | Hazard | Additional strain on patient | 563068 | Cause | Rescan due to user error | 563069 | Measure | No additional measures required
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563064 | Hazard | Additional strain on patient | 563070 | Cause | Rescan due to communication issue | 563071 | Measure | No additional measures required
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563072 | Hazard | Short-term delay | 563073 | Cause | Communication issues | 563074 | Measure | No additional measures required
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563072 | Hazard | Short-term delay | 563075 | Cause | User error | 563076 | Measure | Reference - No scan is possible without confirmation
562334 | Chapter | Perform Clinical Procedure | 563043 | Function | 3D Navigation | 563072 | Hazard | Short-term delay | 563075 | Cause | User error | 563077 | Measure | Reference - Navigation Interface - Scan Confirmation
562334 | Chapter | Perform Clinical Procedure | 563078 | Function | Application Hosting | 563079 | Hazard | Influence on performance or functionality of hosted application | 563080 | Cause | Unauthorized update or execution of unauthorized software | 563081 | Measure | Intrusion detection solution
562334 | Chapter | Perform Clinical Procedure | 563078 | Function | Application Hosting | 563082 | Hazard | Application Hosting PC does not start | 563083 | Cause | Application PC does not start due to HW problem, boot problem | 563084 | Measure | None
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563086 | Hazard | Unwanted radiation | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 563088 | Measure | Label the wireless footswitch with location
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563086 | Hazard | Unwanted radiation | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 563089 | Measure | User Manual - Label the wireless footswitch with location
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563086 | Hazard | Unwanted radiation | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 563090 | Measure | Wireless footswitch and system are paired
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563086 | Hazard | Unwanted radiation | 563087 | Cause | User mistakenly uses a wireless footswitch from another system | 563091 | Measure | Label - wireless footswitch with location
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563086 | Hazard | Unwanted radiation | 563092 | Cause | Send/Receive modules of wireless footswitch are not reliable | 563093 | Measure | Safe hardware modules and signal transmission line
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563094 | Hazard | Unnecessary radiation | 563095 | Cause | x-ray indicators and patient not visible for the user | 563096 | Measure | User Manual -Do not release x-ray out of vieweing distance
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563097 | Hazard | Delayed or no treatment | 563098 | Cause | Interference with other wireless Footswitch | 563099 | Measure | User Manual - Interference wireless footswitch
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563097 | Hazard | Delayed or no treatment | 563100 | Cause | Battery empty | 563101 | Measure | User Manual -Do not release x-ray out of vieweing distance
562334 | Chapter | Perform Clinical Procedure | 563085 | Function | Wireless footswitch | 563097 | Hazard | Delayed or no treatment | 563100 | Cause | Battery empty | 563102 | Measure | High capacity battery
563103 | Chapter | Service System | 563104 | Function | Remote service | 563105 | Hazard | Malfunction of the examination routine | 563106 | Cause | Access via remote service | 563107 | Measure | Reference - Remote access
563103 | Chapter | Service System | 563104 | Function | Remote service | 563105 | Hazard | Malfunction of the examination routine | 563108 | Cause | Install Software on wrong system | 563109 | Measure | Check system requirements before installation starts
563103 | Chapter | Service System | 563110 | Function | Software modification incl. Bugfix | 563111 | Hazard | Delayed or no interventional decision | 563112 | Cause | Impermissible or faulty manipulation of SW from network. | 563113 | Measure | Reference - User Manual - Impermissible or faulty manipulation from network
563103 | Chapter | Service System | 563110 | Function | Software modification incl. Bugfix | 563111 | Hazard | Delayed or no interventional decision | 563120 | Cause | Non approved SW/HW | 563121 | Measure | Reference - User Manual - Use of non-approved parts
563103 | Chapter | Service System | 563110 | Function | Software modification incl. Bugfix | 563111 | Hazard | Delayed or no interventional decision | 563120 | Cause | Non approved SW/HW | 563122 | Measure | Use of unauthorized software
563103 | Chapter | Service System | 563123 | Function | Adjustment (e.g. dose rate adjustment) | 563124 | Hazard | Unwanted radiation released only by SW (closed HW loop in Service Mode, e.g. during calibration) | 563125 | Cause | SW or communication error | 563128 | Measure | Open HW-loop after Service
563103 | Chapter | Service System | 563123 | Function | Adjustment (e.g. dose rate adjustment) | 563129 | Hazard | Increased radiation exposure during service activities | 563130 | Cause | Radiation safety guidelines are not followed | 563131 | Measure | Service Manual - Radiation safety
563103 | Chapter | Service System | 563132 | Function | Replace electrical parts / Replace mechanical parts | 563133 | Hazard | Environmental hazard | 563134 | Cause | Improper disposal of hazardous materials | 563135 | Measure | Dismantling
563103 | Chapter | Service System | 563132 | Function | Replace electrical parts / Replace mechanical parts | 563139 | Hazard | Thermal energy and high temperature | 563140 | Cause | Parts too warm | 563141 | Measure | Service Instruction - Hot Surface of TCP backside
563103 | Chapter | Service System | 563132 | Function | Replace electrical parts / Replace mechanical parts | 563142 | Hazard | ESD | 563143 | Cause | ESD related damages of electronic components | 563144 | Measure | Service Instruction - ESD
563103 | Chapter | Service System | 563132 | Function | Replace electrical parts / Replace mechanical parts | 563145 | Hazard | Spring jumps out when replacing the standard monitor holder on systems with a standard monitor column (non-motorized) | 563146 | Cause | Removing the snap ring without holding the spring | 563147 | Measure | Service Instruction - Hold spring of monitor holder
563103 | Chapter | Service System | 563132 | Function | Replace electrical parts / Replace mechanical parts | 563148 | Hazard | Inaccurate navigation reesult | 563149 | Cause | Removing and remounting the FD cover and optional mechanical interface of the navigation system | 563150 | Measure | Service Instruction - Inform customer to check navigation accuracy
563103 | Chapter | Service System | 563151 | Function | UPS operation | 563152 | Hazard | Electrical shock | 563153 | Cause | Power is not switched off | 563154 | Measure | Service Instruction - Power Presence in System
563103 | Chapter | Service System | 563151 | Function | UPS operation | 563152 | Hazard | Electrical shock | 563153 | Cause | Power is not switched off | 563155 | Measure | Warning labeling
563103 | Chapter | Service System | 563151 | Function | UPS operation | 563152 | Hazard | Electrical shock | 563153 | Cause | Power is not switched off | 563156 | Measure | Label - Unplug the power input connector
563103 | Chapter | Service System | 563151 | Function | UPS operation | 563152 | Hazard | Electrical shock | 563157 | Cause | Inappropiate devices connected to UPS | 563158 | Measure | User Manual - UPS
563103 | Chapter | Service System | 563151 | Function | UPS operation | 563152 | Hazard | Electrical shock | 563157 | Cause | Inappropiate devices connected to UPS | 563159 | Measure | Service Instruction UPS
563103 | Chapter | Service System | 563160 | Function | Installation General | 563161 | Hazard | No hazard on system level | 563162 | Cause | None | 563163 | Measure | None
563103 | Chapter | Service System | 563164 | Function | Software installation | 563165 | Hazard | No hazard on system level | 563166 | Cause | None | 563167 | Measure | None
563103 | Chapter | Service System | 563168 | Function | Service General | 563169 | Hazard | No hazard on system level | 563170 | Cause | None | 563171 | Measure | None
563103 | Chapter | Service System | 563172 | Function | Report unexpected event | 563173 | Hazard | No hazard on system level | 563174 | Cause | None | 563175 | Measure | None
563103 | Chapter | Service System | 563176 | Function | Remote service | 563177 | Hazard | Malfunction of the examination routine | 563178 | Cause | Access via remote service | 563179 | Measure | Remote access
563103 | Chapter | Service System | 563176 | Function | Remote service | 563180 | Hazard | Remote software update | 563181 | Cause | Installing the hotfix or software update | 563182 | Measure | User Manual - Remote Software Update
563103 | Chapter | Service System | 563176 | Function | Remote service | 563180 | Hazard | Remote software update | 563183 | Cause | Failed installation of software | 563184 | Measure | User Manual - Remote Update Failure
563103 | Chapter | Service System | 563185 | Function | Cables | 563186 | Hazard | Electric Shock | 563187 | Cause | Moving Components | 563188 | Measure | Reference - Cable conduit for movable cables
563103 | Chapter | Service System | 563189 | Function | Service documentation | 563190 | Hazard | No hazard on system level | 563191 | Cause | None | 563192 | Measure | None
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563194 | Hazard | Impaired vision through laser-beam | 563195 | Cause | Service staff looking into laser-beam | 563196 | Measure | Reference - Usage of laser class 1
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563197 | Hazard | Excessive temperature, fire, explosion | 563198 | Cause | Improper exchange of batteries | 563199 | Measure | Warning Label - Exchange Battery UPS
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563197 | Hazard | Excessive temperature, fire, explosion | 563198 | Cause | Improper exchange of batteries | 563200 | Measure | Warning Label - Exchange Battery HICSwitch
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563197 | Hazard | Excessive temperature, fire, explosion | 563201 | Cause | Short-circuit of batteries during exchange | 563202 | Measure | Safe battery cases TrolleyU
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563197 | Hazard | Excessive temperature, fire, explosion | 563201 | Cause | Short-circuit of batteries during exchange | 563203 | Measure | Safe battery cases HICSwitch
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563197 | Hazard | Excessive temperature, fire, explosion | 563201 | Cause | Short-circuit of batteries during exchange | 563204 | Measure | Service Caution - Servicing battery components
563103 | Chapter | Service System | 563193 | Function | Maintenance, service / trouble shooting | 563205 | Hazard | Risk of Infection | 563206 | Cause | Opening Chassis / Trolley | 563207 | Measure | Service Instruction - Measures to prevent infection due to contaminated system parts - service
563103 | Chapter | Service System | 563208 | Function | Imaging System, Field service | 563209 | Hazard | Wrong interventional decision | 563210 | Cause | Service parameters changed by unauthorized user | 563211 | Measure | Reference - User Manual - Trained Users
563103 | Chapter | Service System | 563208 | Function | Imaging System, Field service | 563209 | Hazard | Wrong interventional decision | 563210 | Cause | Service parameters changed by unauthorized user | 563212 | Measure | Reference - Service training
563103 | Chapter | Service System | 563213 | Function | 3D Service functions | 563214 | Hazard | Wrong interventional decision or treatment | 563215 | Cause | Unprecise 3D calibration due to a badly positioned calibration phantom | 563216 | Measure | Correct Position of 3D-Calibration Phantom
563103 | Chapter | Service System | 563213 | Function | 3D Service functions | 563214 | Hazard | Wrong interventional decision or treatment | 563215 | Cause | Unprecise 3D calibration due to a badly positioned calibration phantom | 580671 | Measure | Correct Position of 3D-Calibration Phantom
563103 | Chapter | Service System | 563213 | Function | 3D Service functions | 563214 | Hazard | Wrong interventional decision or treatment | 563217 | Cause | Wrong 3D-calibration due to wrong transformation matrices | 563218 | Measure | measuring instrument boards are positioned correctly
563103 | Chapter | Service System | 563213 | Function | 3D Service functions | 563219 | Hazard | Risk of injury due to structure of surface | 563220 | Cause | Sharp edges | 563221 | Measure | Note in interface description
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563224 | Hazard | Electrical safety, Electrical shock | 563225 | Cause | Electrical energy is present after shutdown. | 563226 | Measure | Warning label
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563224 | Hazard | Electrical safety, Electrical shock | 563225 | Cause | Electrical energy is present after shutdown. | 563227 | Measure | Reference - Service Documentation - Voltage Presence
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563228 | Hazard | Loss of data | 563229 | Cause | System shutdown because of a software failure | 563230 | Measure | Reference - Protect shutdown
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563228 | Hazard | Loss of data | 563231 | Cause | User accidentally shuts off system during examination | 563232 | Measure | Protect shutdown
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563228 | Hazard | Loss of data | 563233 | Cause | System restart after X-ray exposure | 563234 | Measure | No measure required
563222 | Chapter | Shutdown System | 563223 | Function | Shut down system & restart | 563235 | Hazard | Delayed or no interventional decision | 563236 | Cause | System automatically switches on/off | 563237 | Measure | Reference - Display system status
563222 | Chapter | Shutdown System | 563238 | Function | Clean System | 563239 | Hazard | Risk of infection | 563240 | Cause | Improper cleaning | 563241 | Measure | User Manual - Cleaning of Wheels
563222 | Chapter | Shutdown System | 563238 | Function | Clean System | 563239 | Hazard | Risk of infection | 563242 | Cause | Contaminated device wheels | 563243 | Measure | Reference - User Manual - Cleaning of Wheels
563222 | Chapter | Shutdown System | 563238 | Function | Clean System | 563239 | Hazard | Risk of infection | 563244 | Cause | Insufficient cleaning | 563245 | Measure | Compliant with IEC60601-1
563222 | Chapter | Shutdown System | 563246 | Function | Image data storage | 563247 | Hazard | Unnecessary radiation | 563248 | Cause | Image data loss | 563249 | Measure | The imaging system performs a controlled shutdown
563250 | Chapter | Additional hazards from subsystems | 563251 | Function | Movements during patient positioning | 563252 | Hazard | Crushing of patient | 563253 | Cause | Contact of Patient with System | 563254 | Measure | User Manual - Crushing zones
563250 | Chapter | Additional hazards from subsystems | 563263 | Function | Virus Protection | 563264 | Hazard | Wrong or no interventional decision | 563265 | Cause | Degraded system performance | 563266 | Measure | Verify access to internet
563250 | Chapter | Additional hazards from subsystems | 563263 | Function | Virus Protection | 563264 | Hazard | Wrong or no interventional decision | 563265 | Cause | Degraded system performance | 563267 | Measure | Verify access to BIOS
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563269 | Hazard | Hazards from fire or smoke | 563270 | Cause | Overheating | 563271 | Measure | Service Documentation usage of wrong fuses
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563269 | Hazard | Hazards from fire or smoke | 563270 | Cause | Overheating | 563272 | Measure | Reference to warning label next to fuses to use correct fuses
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563269 | Hazard | Hazards from fire or smoke | 563270 | Cause | Overheating | 563273 | Measure | Reference to temperature monitoring of the ESU discharge resistors and inspection
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563274 | Hazard | Environmental hazard | 563275 | Cause | Improper disposal of hazardous materials | 563276 | Measure | Improper disposal of ESU (ultra capacitor cells with harmful and flammable liquid)
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563274 | Hazard | Environmental hazard | 563275 | Cause | Improper disposal of hazardous materials | 563277 | Measure | Reference to material request form and label on ESU
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563279 | Cause | Voltage still present after switching-off | 563280 | Measure | Service Documentation - Voltage Presence
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563279 | Cause | Voltage still present after switching-off | 596998 | Measure | Factory Documentation - Voltage Presence
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563279 | Cause | Voltage still present after switching-off | 563282 | Measure | Reference to discharge circuits
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563283 | Cause | ESU requires discharge after switch-off | 563284 | Measure | Service Documentation ESU discharge
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563283 | Cause | ESU requires discharge after switch-off | 563285 | Measure | Supply Chain Documentation ESU discharge
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 563283 | Cause | ESU requires discharge after switch-off | 563286 | Measure | Reference to ESU discharge circuit, display of voltage, warning labels
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 580657 | Cause | Discharge unit defective, intermediate circuit not discharged completely | 580658 | Measure | Warning - Service Documentation - Check discharge of intermediate circuit
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563278 | Hazard | Electric shock | 580657 | Cause | Discharge unit defective, intermediate circuit not discharged completely | 580659 | Measure | Warning - Factory Documentation - Check discharge of intermediate circuit
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563287 | Hazard | Oil leaks out of the monoblock | 563288 | Cause | Component error | 563289 | Measure | Check leakage of monobloc
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563287 | Hazard | Oil leaks out of the monoblock | 563288 | Cause | Component error | 563290 | Measure | Reference to measures against leakage of oil
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563287 | Hazard | Oil leaks out of the monoblock | 563288 | Cause | Component error | 596995 | Measure | Check monobloc leakage
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563291 | Hazard | Leaking cooling liquid | 563292 | Cause | Component error or service error (Refill of Cooling Unit with connected Compensation Bag) | 563293 | Measure | Service Documentation - Disconnect Compensation Bag
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563291 | Hazard | Leaking cooling liquid | 563292 | Cause | Component error or service error (Refill of Cooling Unit with connected Compensation Bag) | 563294 | Measure | Reference to measures against leakage of cooling liquid
563250 | Chapter | Additional hazards from subsystems | 563268 | Function | Documentation Safety Requirements for X-ray Generator | 563295 | Hazard | Unnecessary radiation | 563296 | Cause | Missing error signal about image loss | 563297 | Measure | Missing HV_ON signal
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563299 | Hazard | Risk of burnings | 563300 | Cause | Too high temperature of the detector | 563301 | Measure | User Manual - Detector Temperature
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563299 | Hazard | Risk of burnings | 563300 | Cause | Too high temperature of the detector | 563302 | Measure | Reference to System Warning - Temperature of the Detector
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563303 | Hazard | Malfunction of imaging system during operation due to defective HW | 563304 | Cause | Defective FLC-Hardware | 563305 | Measure | User Manual - Defective FLC HW
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563306 | Hazard | Risk of incorrect interventional decision caused by wrong image processing | 563307 | Cause | Wrong LUT | 563308 | Measure | User Manual - Wrong LUT
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563310 | Cause | Plugging/Unplugging USB devices interfere with performance of imaging system | 563311 | Measure | User Manual - Plug/Unplug USB Devices
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563312 | Cause | Unauthorized installation of SW/HW | 563313 | Measure | Reference to User Manual - Unauthorized Software
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563314 | Cause | Network failure | 563315 | Measure | User Manual - Network Failure
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563316 | Cause | Software lockup | 563317 | Measure | User Manual - Software Lockup
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563318 | Cause | Restart missing after Service session | 563319 | Measure | User Manual - Reboot after Service
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563309 | Hazard | Delayed or no interventional decision | 563318 | Cause | Restart missing after Service session | 563320 | Measure | Message to request reboot after service
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563321 | Hazard | Danger of incorrect interventional decision | 563322 | Cause | Assignment of series/ scenes/ images to the wrong patient | 563323 | Measure | User Manual - Mix-up of Patients
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563321 | Hazard | Danger of incorrect interventional decision | 563324 | Cause | Scale displayed for non-calibrated images | 563325 | Measure | Reference to Caution for 2D measurements
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563326 | Hazard | loss of images or patient data | 563327 | Cause | Postponed loss of recorded patient data | 563329 | Measure | User Manual - Loss of Data USB
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563326 | Hazard | loss of images or patient data | 563331 | Cause | Harddisk or file system defect | 563332 | Measure | User Manual - Hard Disk Failure
563250 | Chapter | Additional hazards from subsystems | 563298 | Function | Documentation user manual for Imaging System | 563333 | Hazard | loss of image quality | 563334 | Cause | Incorrect calibration data used | 563335 | Measure | Note in service documentation
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563337 | Hazard | Electric Shock | 563338 | Cause | Intrusion of Liquids and Instruments | 563339 | Measure | Reference - User Manual - Insufficient cleaning
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563337 | Hazard | Electric Shock | 563338 | Cause | Intrusion of Liquids and Instruments | 563340 | Measure | Reference - Spillage of Liquids on System
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563337 | Hazard | Electric Shock | 563338 | Cause | Intrusion of Liquids and Instruments | 563341 | Measure | User Manual - Intrusion of Instruments xCP
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563342 | Hazard | Overheating | 563343 | Cause | Inappropriate Draping | 563344 | Measure | User Manual - Use Specified Drapes Only
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563342 | Hazard | Overheating | 563345 | Cause | Hot Surface of RCP | 563346 | Measure | User Manual - Avoid Patient Contact with RCP
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563342 | Hazard | Overheating | 563345 | Cause | Hot Surface of RCP | 563347 | Measure | User Manual - Positioning of RCP
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563342 | Hazard | Overheating | 563345 | Cause | Hot Surface of RCP | 563348 | Measure | Reference - Compliance with IEC 60601-1
563250 | Chapter | Additional hazards from subsystems | 563336 | Function | User Manual for Touch Interfaces (xCP) | 563349 | Hazard | Risk of injury | 563350 | Cause | Falling Down of RCP | 563351 | Measure | User Manual - Falling Down
563353 | Chapter | Processes employed | 563354 | Function | Refurbishment | 563355 | Hazard | Biological contamination or infection during refurbishment activities | 563356 | Cause | Contact with contaminated system parts during deinstallation or cleaning for refurbishment | 563357 | Measure | Measures to prevent infection due to contaminated system parts - refurbishment
563353 | Chapter | Processes employed | 563358 | Function | Service / maintenance | 563359 | Hazard | Biological contamination or infection during service activities | 563360 | Cause | Contact with contaminated system parts during service activities | 563361 | Measure | Reference to measures to prevent infection due to contaminated system parts - service
563353 | Chapter | Processes employed | 563358 | Function | Service / maintenance | 563362 | Hazard | Undetected damage of the cooling unit | 563363 | Cause | Overpressure test of cooling unit assemly not performed correctly | 563364 | Measure | Service Documentation - Overpressure test
563353 | Chapter | Processes employed | 563365 | Function | Manufacturing | 578701 | Hazard | None | 578702 | Cause | None | 578703 | Measure | None
563353 | Chapter | Processes employed | 563366 | Function | Supply chain management process | 578704 | Hazard | None | 578705 | Cause | None | 578706 | Measure | None
563353 | Chapter | Processes employed | 563367 | Function | Packaging and transport | 578707 | Hazard | None | 578708 | Cause | None | 578709 | Measure | None
563353 | Chapter | Processes employed | 563368 | Function | Disposal | 578710 | Hazard | None | 578711 | Cause | None | 578712 | Measure | None